Measurements on stator with adhered SAW generating unit
Measurement of frequency dependent impedances adhered variant

Gauge: ZV-Z132 / R&S ZVL
Settings:
- Power: 0 dBm
- ZF-Bandwidth: 100 Hz
- B1 and B2 damping: 0 dB
- Averaging: 10
Data are stored in the file: Network_analyzer_measurements_adhered.ods

Optical deflection measurements for frequencies with increased power drain adhered variant

We determined the frequencies via impedance measurements, where SAW have been generated. For the relevant frequency (f = 3.8 MHz) we performed single frequency measurements. Results are listed in file: vibrometer_measurements_adhered.ods

Measurement Data:
Meas. Nr.	Actuator	Localization	Frequency in Hz	Internal resistance of function generator in Ohm	Applied voltage (peak-peak, including internal resistance) in V	Averaging	Table in vibrometer_measurements.ods	Comment
1	Meas. unit 2	Line; Close to Actuator normal to traveling direction	3.65e6	50	20	10	Meas.Nr.1
2	Meas. unit 2	Line; In traveling direction 26.5 mm behind Meas. position 1; Normal to traveling direction	3.65e6	50	20	10	Meas.Nr.2

Computation of standardized deflection, wave and input power and efficiency for adhered actuator at f1 = 3.8 MHz

dzSmax(u0p=1 V)=3.13e-10 m, PP=9.92e-3 W, Pel=58.577e-3 W, p=0.169 according to Maxima file: Computation_of_results_measurement_adhered_3p8e6Hz_position_1.wxmx

Determination of SAW attenuation on steel substrate

Measurement 2 presents according to Maxima file Computation_of_results_measurement_adhered_3p8e6Hz_position_2.wxmx an average displacement of dzSmax(u0p=1 V)=237.3e-12 m at a distance of 26.5 mm regarding to vibrometer_measurements.ods. According to Maxima file SAW_damping_on_steel_in_dB_per_mm.wxmx there is a damping of 0.09 dB/mm

Comparison of adhered actuator to numerical simulation

For comparison we built up 2D FEM models for the simulation software ANSYS. The APDL file Normal_deflection_series_adhered.txt gives frequency dependent normal deflection of the excited steel substrate surface with average value and standard derivation and the impedance of the SAW generating unit.

Computation of standardized deflection, wave and input power and efficiency for normal simulation model for adhered actuator at f1 = 3.65 MHz

dzSmax(u0p=1 V)=4360e-12 m, PP=1.76e-3 W, Pel=5.79e-3 W, p=0.30 according to Maxima file: Computation_of_results_simulation_adhered_3p65e6Hz.wxmx
Measurements on SAW motor
The SAW motor settings for the documented results are:
Stator: „Measurement unit 1“
Slider: Silicon slider with 16,560 projections with a high of 2 µm and a diameter of 22.5 µm on an area of 16 mm2
Input voltage (amplitude): 50 V
Drive modus: Burst
Burst periods: 50,000
Measurement of blocking force for SAW motor at f = 3.85 MHz

blocking force measurements were made with the force gauge Imada DPX. The results are listed in blocking_force_measurements_SAW_motor_3p85e6Hz.ods

Measurement of time dependent displacements for SAW motor at f = 3.85 MHz

displacement measurements were made with the laser triangulation sensor Keyence LK-G82. The results are listed in time_dependent_diplacement_measurements_SAW_motor_3p85Hz.ods
Measurements on stator with thick film SAW generating unit
Measurement of frequency dependent impedances thick film variant

Gauge: ZV-Z132 / R&S ZVL
Settings:
- Power: 0 dBm
- ZF-Bandwidth: 100 Hz
- B1 and B2 damping: 0 dB
- Averaging: 10
Data are stored in the file: Network_analyzer_measurements_thick_film.ods

Optical deflection measurements for frequencies with increased power drain thick film variant

We determined the frequencies via impedance measurements, where SAW have been generated. For these frequencies (f1 = 4.15 MHz and f2 = 5.45 MHz) we performed optical measurements (vibrometer UHF-120). Results are listed in file: vibrometer_measurements_thick_film.ods

Measurement Data:
Meas. Nr.	Actuator	Localization	Frequency in Hz	Internal resistance of function generator in Ohm	Applied voltage (peak-peak, including internal resistance) in V	Averaging	Table in vibrometer_measurements.ods	Comment
1	Meas. unit 1	Line; Close to Actuator normal to traveling direction	4.15e6	50	20	10	Meas.Nr.1
2	Meas. unit 1	Line; Close to Actuator normal to traveling direction	5.45e6	50	20	10	Meas.Nr.2
3	Meas. unit 1	Line; In traveling direction 16.8 mm behind Meas. position 3; Normal to traveling direction	5.45e6	50	20	10	Meas.Nr.3
4	Meas. unit 1	Line; In traveling direction; Centered to actuator	5.45e6	50	20	10	Meas.Nr.4

Computation of standardized deflection, wave and input power and efficiency for thick film actuator at f1 = 4.15 MHz

dzSmax(u0p=1 V)=48.6e-12 m, PP=0.191e-3 W, Pel=2.318e-3 W, p=0.083 according to Maxima file: Computation_of_results_measurement_thick_film_4p15e6Hz.wxmx

Computation of standardized deflection, wave and input power and efficiency for thick film actuator at f2 = 5.45 MHz

dzSmax(u0p=1 V)=33.8e-12 m, PP=0.122e-3 W, Pel=3.696e-3 W, p=0.033 according to Maxima file: Computation_of_results_measurement_thick_film_5p45e6Hz.wxmx

Comparison of thick film actuator to numerical simulation

For comparison we built up 2D FEM models for the simulation software ANSYS. The APDL file Normal_deflection_series_thick_film.txt gives frequency dependent normal deflection of the excited steel substrate surface with average value and standard derivation and the impedance of the SAW generating unit. The APDL file Normal_deflection_series_thick_film_smooth_junction.txt is modified with a smooth junction from PZT layer to bare steel substrate, realized by incremental reduction of PZT density and stiffness.

Computation of standardized deflection, wave and input power and efficiency for normal simulation model for thick film actuator at f1 = 4.00 MHz

dzSmax(u0p=1 V)=81.6e-12 m, PP=0.620e-3 W, Pel=3.597e-3 W, p=0.145 according to Maxima file: Computation_of_results_simulation_thick_film_4p00e6Hz.wxmx

Computation of standardized deflection, wave and input power and efficiency for normal simulation model for thick film actuator at f1 = 5.35 MHz

dzSmax(u0p=1 V)=182.4e-12 m, PP=3.474e-3 W, Pel=8.344e-3 W, p=0.416 according to Maxima file: Computation_of_results_simulation_thick_film_5p35e6Hz.wxmx

Computation of standardized deflection, wave and input power and efficiency for simulation model with smooth junction for thick film actuator at f1 = 4.00 MHz

dzSmax(u0p=1 V)=160.8e-12 m, PP=2.018e-3 W, Pel=4.922e-3 W, p=0.410 according to Maxima file: Computation_of_results_simulation_thick_film_4p00e6Hz_smooth_junction.wxmx

Computation of standardized deflection, wave and input power and efficiency for simulation model with smooth junction for thick film actuator at f1 = 5.35 MHz

dzSmax(u0p=1 V)=85.0e-12 m, PP=0.755e-3 W, Pel=5.604e-3 W, p=0.135 according to Maxima file: Computation_of_results_simulation_thick_film_5p35e6Hz_smooth_junction.wxmx