{"text_format":"1","workbook":"xlsx","n_sheets":11,"sheets":["FQ","SQ","ConcTable","QCReview","Container1","Summary1","Instructions","Container2","Summary2","Container3","Summary3"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"FQ","index":0,"n_rows":8242,"n_cols":31,"rows_kept":2250,"cut":true,"hidden":false,"merged":[],"bounds":{"rows_first":1000,"rows_last":250,"rows_between":1000,"halvings":1}}
["Operator","Sample Name","Batch Name","Date and Time Acquired","Sample Type","Dilution Factor","Data File Name","Data Path","Instrument Name","Sample Weight or Volume","Final Weight or Volume","Dilution Multiplier","Misc. Info or Comment","Analyte","Element Full Name","Mass","Concentration","Conc RSD","CPS Mean","CPS RSD","CPS Rep1","CPS Rep2","CPS Rep3","Units","ISTD %Rec","ISTD Ref Mass","Tune Step","R Value","Slope","Y intercept","Helper"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,60,null,40,90,50,"ug/l",null,103,1,null,null,null,{"$f":"N2&\"_\"&P2&\"_\"&G2","v":"Li_7_001SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,{"$f":"N3&\"_\"&P3&\"_\"&G3","v":"Be_9_001SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,473.34,null,560.01,490.01,370,"ug/l",null,103,1,null,null,null,"B_10_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,2463.5,null,2140.13,2540.1799999999998,2710.2,"ug/l",null,103,1,null,null,null,"B_11_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,684633.22,null,688994.13,688325.46,676580.07,"ug/l",null,103,1,null,null,null,"Al_27_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25231.06,null,25838.75,25678.25,24176.19,null,1.30524215597224,null,1,null,null,null,"Sc_45_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,2400.17,null,2600.21,2190.14,2410.17,"ug/l",null,103,1,null,null,null,"Ge_74_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,226.67,null,230,260,190,null,2.7296935129026399E-2,null,1,null,null,null,"Rh_103_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,616.67999999999995,null,640.01,620.01,590.01,"ug/l",null,103,1,null,null,null,"Sn_118_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,446.67,null,560.01,290,490.01,"ug/l",null,103,1,null,null,null,"Sb_121_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,376.67,null,400.01,310,420.01,"ug/l",null,103,1,null,null,null,"Sb_123_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,56.67,null,20,30,120,"ug/l",null,103,1,null,null,null,"Te_125_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,36.67,null,30,50,30,"ug/l",null,103,1,null,null,null,"Hg_202_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,490.01,null,410,600.01,460.01,"ug/l",null,103,2,null,null,null,"Al_27_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58362.99,null,59937.919999999998,58704.01,56447.040000000001,"ug/l",null,103,2,null,null,null,"Si_28_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,10,null,0,10,20,"ug/l",null,103,2,null,null,null,"P_31_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,433.34,null,450.01,450.01,400,"ug/l",null,103,2,null,null,null,"S_34_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,20728.63,null,20211.5,21262.5,20711.89,"ug/l",null,103,2,null,null,null,"K_39_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,296.67,null,390,240,260,"ug/l",null,103,2,null,null,null,"Ca_44_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,380,null,620.01,340,180,null,1.44823053188722,null,2,null,null,null,"Sc_45_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,6.67,null,6,6,8,"ug/l",null,103,2,null,null,null,"Ti_49_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,53.33,null,70,30,60,"ug/l",null,103,2,null,null,null,"V_51_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1226.71,null,1180.04,1310.05,1190.04,"ug/l",null,103,2,null,null,null,"Cr_52_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,380,null,390,390.01,360,"ug/l",null,103,2,null,null,null,"Mn_55_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7134.94,null,8202.17,7101.47,6101.17,"ug/l",null,103,2,null,null,null,"Fe_56_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,165,null,145,190,160,"ug/l",null,103,2,null,null,null,"Co_59_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,353.34,null,480.01,340,240,"ug/l",null,103,2,null,null,null,"Ni_60_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,525.01,null,350.01,750.02,475.01,"ug/l",null,103,2,null,null,null,"Cu_65_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,11.67,null,20,5,10,"ug/l",null,103,2,null,null,null,"Ga_71_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,40,null,35,35,50,"ug/l",null,103,2,null,null,null,"Ge_74_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Se_78_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,80,null,130,70,40,"ug/l",null,103,2,null,null,null,"Br_79_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,166.67,null,100,100,300,"ug/l",null,103,2,null,null,null,"Rb_85_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"Y_89_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Zr_90_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mo_98_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,6.67,null,0,10,10,null,2.4565663712031501E-3,null,2,null,null,null,"Rh_103_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.33,null,-2.73,0.36,1.59,"ug/l",null,103,2,null,null,null,"Cd_114_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,0,null,-60,-36,-9,"ug/l",null,103,2,null,null,null,"In_115_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,58.33,null,100,60,15,"ug/l",null,103,2,null,null,null,"Sn_118_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,40,null,85,35,0,"ug/l",null,103,2,null,null,null,"Sb_121_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,100,null,100,200,0,"ug/l",null,103,2,null,null,null,"I_127_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,5,null,0,10,5,"ug/l",null,103,2,null,null,null,"Cs_133_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Ce_140_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,66.67,null,60,70,70,"ug/l",null,103,2,null,null,null,"Tb_159_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"W_184_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Pt_195_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,366.68,null,300.01,400.01,400.01,"ug/l",null,103,2,null,null,null,"Au_197_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Hg_202_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Tl_205_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,76.67,null,110,40,80,"ug/l",null,103,2,null,null,null,"Pb_208_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,10,null,0,0,30,"ug/l",null,103,2,null,null,null,"Bi_209_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,{"$f":"N67&\"_\"&P67&\"_\"&G67","v":"Th_232_001SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:22:58","Sample",1,"001SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_001SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,66.67,null,30,80,90,"ug/l",null,103,1,null,null,null,"Li_7_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,530.01,null,630.01,520.01,440.01,"ug/l",null,103,1,null,null,null,"B_10_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,2110.12,null,2010.11,2070.12,2250.14,"ug/l",null,103,1,null,null,null,"B_11_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,633172.75,null,625637.32999999996,630027.26,643853.66,"ug/l",null,103,1,null,null,null,"Al_27_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,23081.46,null,22474.04,23705.59,23064.74,null,1.19403965915216,null,1,null,null,null,"Sc_45_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,1936.78,null,2090.13,1790.09,1930.11,"ug/l",null,103,1,null,null,null,"Ge_74_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,203.33,null,160,260,190,null,2.4486956512803099E-2,null,1,null,null,null,"Rh_103_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,573.35,null,540.01,770.02,410.01,"ug/l",null,103,1,null,null,null,"Sn_118_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,436.67,null,440.01,360,510.01,"ug/l",null,103,1,null,null,null,"Sb_121_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,373.34,null,340.01,270,510.01,"ug/l",null,103,1,null,null,null,"Sb_123_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,43.33,null,70,20,40,"ug/l",null,103,1,null,null,null,"Te_125_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,26.67,null,30,10,40,"ug/l",null,103,1,null,null,null,"Hg_202_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,66.67,null,0,200.01,0,"ug/l",null,103,2,null,null,null,"Mg_24_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,433.34,null,500.01,410.01,390,"ug/l",null,103,2,null,null,null,"Al_27_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,59600.26,null,60008.25,59456.52,59336,"ug/l",null,103,2,null,null,null,"Si_28_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,33.33,null,50,30,20,"ug/l",null,103,2,null,null,null,"P_31_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,340.01,null,410.01,210,400.01,"ug/l",null,103,2,null,null,null,"S_34_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,21446.3,null,19911.189999999999,21913.67,22514.04,"ug/l",null,103,2,null,null,null,"K_39_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,260,null,150,360,270,"ug/l",null,103,2,null,null,null,"Ca_44_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,373.34,null,360,480.01,280,null,1.4228232015675299,null,2,null,null,null,"Sc_45_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,10,null,8,12,10,"ug/l",null,103,2,null,null,null,"Ti_49_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,66.67,null,100,50,50,"ug/l",null,103,2,null,null,null,"V_51_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1290.05,null,1230.04,1420.06,1220.04,"ug/l",null,103,2,null,null,null,"Cr_52_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,366.67,null,440.01,330,330,"ug/l",null,103,2,null,null,null,"Mn_55_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7034.83,null,5801.03,8602.1,6701.35,"ug/l",null,103,2,null,null,null,"Fe_56_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,113.33,null,135,75,130,"ug/l",null,103,2,null,null,null,"Co_59_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,300,null,210,350,340,"ug/l",null,103,2,null,null,null,"Ni_60_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,575.01,null,500.01,575.01,650.01,"ug/l",null,103,2,null,null,null,"Cu_65_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,8.33,null,10,10,5,"ug/l",null,103,2,null,null,null,"Ga_71_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,31.67,null,50,30,15,"ug/l",null,103,2,null,null,null,"Ge_74_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"As_75_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,5,0,5,"ug/l",null,103,2,null,null,null,"Se_78_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,60,null,30,110,40,"ug/l",null,103,2,null,null,null,"Br_79_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,83.33,null,0,150,100,"ug/l",null,103,2,null,null,null,"Rb_85_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Y_89_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Zr_90_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Mo_98_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,23.33,null,0,50,20,null,8.5979822992110107E-3,null,2,null,null,null,"Rh_103_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.33,null,1.18,1.73,-0.55000000000000004,"ug/l",null,103,2,null,null,null,"Cd_114_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,0,null,-18,-6,-12,"ug/l",null,103,2,null,null,null,"In_115_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,20,null,30,10,20,"ug/l",null,103,2,null,null,null,"Sn_118_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,36.67,null,30,60,20,"ug/l",null,103,2,null,null,null,"Sb_121_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,166.67,null,0,200.01,300.01,"ug/l",null,103,2,null,null,null,"I_127_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,13.33,null,10,25,5,"ug/l",null,103,2,null,null,null,"Cs_133_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"Ce_140_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,43.33,null,30,50,50,"ug/l",null,103,2,null,null,null,"Tb_159_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"W_184_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pt_195_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,416.68,null,700.02,50,500.01,"ug/l",null,103,2,null,null,null,"Au_197_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,10,null,0,0,30,"ug/l",null,103,2,null,null,null,"Hg_202_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,36.67,null,20,10,80,"ug/l",null,103,2,null,null,null,{"$f":"N131&\"_\"&P131&\"_\"&G131","v":"Tl_205_002SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,80,null,90,110,40,"ug/l",null,103,2,null,null,null,"Pb_208_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,10,null,0,10,20,"ug/l",null,103,2,null,null,null,"Bi_209_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:24:26","Sample",1,"002SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_002SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,43.33,null,50,80,0,"ug/l",null,103,1,null,null,null,"Li_7_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,523.34,null,550.01,570.01,450.01,"ug/l",null,103,1,null,null,null,"B_10_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,2456.83,null,2440.16,2800.21,2130.12,"ug/l",null,103,1,null,null,null,"B_11_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,654362.59,null,650950.30000000005,664107.02,648030.46,"ug/l",null,103,1,null,null,null,"Al_27_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,23205,null,23605.5,23525.32,22484.19,null,1.2004309204094801,null,1,null,null,null,"Sc_45_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,1846.77,null,1960.11,1580.08,2000.12,"ug/l",null,103,1,null,null,null,"Ge_74_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,240,null,210,250,260,null,2.8902637195439701E-2,null,1,null,null,null,"Rh_103_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,386.67,null,300,450.01,410.01,"ug/l",null,103,1,null,null,null,"Sn_118_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,503.34,null,600.01,590.01,320,"ug/l",null,103,1,null,null,null,"Sb_121_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,396.67,null,460.01,390.01,340,"ug/l",null,103,1,null,null,null,"Sb_123_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,33.33,null,50,20,30,"ug/l",null,103,1,null,null,null,"Te_125_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,23.33,null,40,10,20,"ug/l",null,103,1,null,null,null,"Hg_202_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,400,null,360,460.01,380,"ug/l",null,103,2,null,null,null,"Al_27_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58436.5,null,58704.04,58252.57,58352.88,"ug/l",null,103,2,null,null,null,"Si_28_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,26.67,null,0,20,60,"ug/l",null,103,2,null,null,null,"P_31_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,363.34,null,520.01,260,310,"ug/l",null,103,2,null,null,null,"S_34_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,21396.1,null,22364,20962.310000000001,20862,"ug/l",null,103,2,null,null,null,"K_39_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,246.67,null,240,220,280,"ug/l",null,103,2,null,null,null,"Ca_44_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,406.67,null,400.01,320,500.01,null,1.54987255683112,null,2,null,null,null,"Sc_45_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,6,null,16,2,0,"ug/l",null,103,2,null,null,null,"Ti_49_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,56.67,null,30,60,80,"ug/l",null,103,2,null,null,null,"V_51_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1466.73,null,1570.07,1550.07,1280.05,"ug/l",null,103,2,null,null,null,"Cr_52_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,283.33,null,260,220,370,"ug/l",null,103,2,null,null,null,"Mn_55_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7368.34,null,6501.24,6801.47,8802.32,"ug/l",null,103,2,null,null,null,"Fe_56_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,123.33,null,110,85,175,"ug/l",null,103,2,null,null,null,"Co_59_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,250,null,280,170,300,"ug/l",null,103,2,null,null,null,"Ni_60_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,358.34,null,250,550.01,275,"ug/l",null,103,2,null,null,null,"Cu_65_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,33.33,null,100,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"Ga_71_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,16.670000000000002,null,15,25,10,"ug/l",null,103,2,null,null,null,"Ge_74_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Se_78_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,60,null,50,80,50,"ug/l",null,103,2,null,null,null,"Br_79_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,100,null,0,0,300.01,"ug/l",null,103,2,null,null,null,"Rb_85_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,13.33,null,10,30,0,"ug/l",null,103,2,null,null,null,"Y_89_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,13.33,null,40,0,0,"ug/l",null,103,2,null,null,null,"Zr_90_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,8.33,null,0,15,8.9,"ug/l",null,103,2,null,null,null,"Mo_98_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Ru_101_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,10,null,0,30,0,null,3.68484955680472E-3,null,2,null,null,null,"Rh_103_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0.67,null,1.45,-0.27,-1.77,"ug/l",null,103,2,null,null,null,"Cd_114_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,0,null,-12,-6,-39,"ug/l",null,103,2,null,null,null,"In_115_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,31.67,null,20,10,65,"ug/l",null,103,2,null,null,null,"Sn_118_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,35,null,55,30,20,"ug/l",null,103,2,null,null,null,"Sb_121_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,166.67,null,200.01,0,300.01,"ug/l",null,103,2,null,null,null,"I_127_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,13.33,null,0,35,5,"ug/l",null,103,2,null,null,null,"Cs_133_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Ba_137_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"W_184_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,{"$f":"N195&\"_\"&P195&\"_\"&G195","v":"Pt_195_003SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,283.33999999999997,null,350.01,350.01,150,"ug/l",null,103,2,null,null,null,"Au_197_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,10,null,10,20,0,"ug/l",null,103,2,null,null,null,"Hg_202_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,16.670000000000002,null,20,20,10,"ug/l",null,103,2,null,null,null,"Tl_205_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,53.33,null,20,90,50,"ug/l",null,103,2,null,null,null,"Pb_208_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,23.33,null,20,30,20,"ug/l",null,103,2,null,null,null,"Bi_209_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:25:54","Sample",1,"003SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_003SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,50,null,60,40,50,"ug/l",null,103,1,null,null,null,"Li_7_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,503.34,null,470.01,500.01,540.01,"ug/l",null,103,1,null,null,null,"B_10_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,2256.81,null,2120.13,2310.15,2340.15,"ug/l",null,103,1,null,null,null,"B_11_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,662151.19999999995,null,649261.98,664761.87,672429.76,"ug/l",null,103,1,null,null,null,"Al_27_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,22951.32,null,23405.200000000001,21843.21,23605.55,null,1.18730748694336,null,1,null,null,null,"Sc_45_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,1776.76,null,1770.09,1640.08,1920.11,"ug/l",null,103,1,null,null,null,"Ge_74_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,203.33,null,240,140,230,null,2.4486956512803099E-2,null,1,null,null,null,"Rh_103_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,516.67999999999995,null,510.01,610.01,430.01,"ug/l",null,103,1,null,null,null,"Sn_118_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,466.68,null,480.01,500.01,420.01,"ug/l",null,103,1,null,null,null,"Sb_121_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,346.67,null,360,350,330,"ug/l",null,103,1,null,null,null,"Sb_123_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,56.67,null,30,90,50,"ug/l",null,103,1,null,null,null,"Te_125_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,13.33,null,30,0,10,"ug/l",null,103,1,null,null,null,"Hg_202_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,450.01,null,400.01,530.01,420.01,"ug/l",null,103,2,null,null,null,"Al_27_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,60650.32,null,59265.760000000002,62105.11,60580.09,"ug/l",null,103,2,null,null,null,"Si_28_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,23.33,null,30,40,0,"ug/l",null,103,2,null,null,null,"P_31_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,370,null,330,390,390.01,"ug/l",null,103,2,null,null,null,"S_34_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19660.87,null,19460.68,20762.189999999999,18759.73,"ug/l",null,103,2,null,null,null,"K_39_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,216.67,null,200,140,310,"ug/l",null,103,2,null,null,null,"Ca_44_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,286.67,null,260,280,320,null,1.0925152037464601,null,2,null,null,null,"Sc_45_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,11.33,null,8,14,12,"ug/l",null,103,2,null,null,null,"Ti_49_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,83.33,null,130,100,20,"ug/l",null,103,2,null,null,null,"V_51_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1373.39,null,1200.04,1550.07,1370.06,"ug/l",null,103,2,null,null,null,"Cr_52_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,463.34,null,460.01,400.01,530.01,"ug/l",null,103,2,null,null,null,"Mn_55_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7435.03,null,6901.41,9102.43,6301.24,"ug/l",null,103,2,null,null,null,"Fe_56_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,98.33,null,95,110,90,"ug/l",null,103,2,null,null,null,"Co_59_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,343.33,null,310,320,400,"ug/l",null,103,2,null,null,null,"Ni_60_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,533.34,null,600.01,475.01,525.01,"ug/l",null,103,2,null,null,null,"Cu_65_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,18.329999999999998,null,25,20,10,"ug/l",null,103,2,null,null,null,"Ga_71_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,16.670000000000002,null,10,25,15,"ug/l",null,103,2,null,null,null,"Ge_74_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"As_75_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,5,null,0,10,5,"ug/l",null,103,2,null,null,null,"Se_78_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,76.67,null,120,90,20,"ug/l",null,103,2,null,null,null,"Br_79_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,116.67,null,150,200.01,0,"ug/l",null,103,2,null,null,null,"Rb_85_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Y_89_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,23.33,null,20,10,40,"ug/l",null,103,2,null,null,null,"Zr_90_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Mo_98_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,0,null,0,0,0,null,0,null,2,null,null,null,"Rh_103_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.33,null,-0.68,-1.0900000000000001,3.18,"ug/l",null,103,2,null,null,null,"Cd_114_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,0,null,-15,-24,-18,"ug/l",null,103,2,null,null,null,"In_115_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,31.67,null,25,40,30,"ug/l",null,103,2,null,null,null,"Sn_118_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,61.67,null,55,60,70,"ug/l",null,103,2,null,null,null,"Sb_121_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,233.34,null,200.01,200,300.01,"ug/l",null,103,2,null,null,null,"I_127_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,11.67,null,5,20,10,"ug/l",null,103,2,null,null,null,"Cs_133_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,{"$f":"N259&\"_\"&P259&\"_\"&G259","v":"Er_166_004SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"W_184_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pt_195_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,383.34,null,100,500.01,550.01,"ug/l",null,103,2,null,null,null,"Au_197_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,23.33,null,60,0,10,"ug/l",null,103,2,null,null,null,"Hg_202_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"Tl_205_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,66.67,null,60,40,100,"ug/l",null,103,2,null,null,null,"Pb_208_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,10,null,10,0,20,"ug/l",null,103,2,null,null,null,"Bi_209_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:27:22","Sample",1,"004SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_004SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,50,null,70,60,20,"ug/l",null,103,1,null,null,null,"Li_7_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,616.67999999999995,null,510.01,610.01,730.02,"ug/l",null,103,1,null,null,null,"B_10_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,2446.84,null,2380.16,2810.22,2150.13,"ug/l",null,103,1,null,null,null,"B_11_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,654334.09,null,650360.42000000004,651772.88,660868.97,"ug/l",null,103,1,null,null,null,"Al_27_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,23518.66,null,22413.91,24005.95,24136.13,null,1.21665704014266,null,1,null,null,null,"Sc_45_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,1946.78,null,1830.1,2170.14,1840.1,"ug/l",null,103,1,null,null,null,"Ge_74_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,170,null,210,110,190,null,2.04727013467698E-2,null,1,null,null,null,"Rh_103_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,406.67,null,370,430.01,420.01,"ug/l",null,103,1,null,null,null,"Sn_118_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,390,null,500.01,300,370,"ug/l",null,103,1,null,null,null,"Sb_121_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,340,null,340,380,300,"ug/l",null,103,1,null,null,null,"Sb_123_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,63.33,null,60,50,80,"ug/l",null,103,1,null,null,null,"Te_125_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,33.33,null,60,40,0,"ug/l",null,103,1,null,null,null,"Hg_202_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,33.33,null,100,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,416.67,null,400,410.01,440.01,"ug/l",null,103,2,null,null,null,"Al_27_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,60894.45,null,60008.11,61262.27,61412.959999999999,"ug/l",null,103,2,null,null,null,"Si_28_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"P_31_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,470.01,null,510.01,520.01,380.01,"ug/l",null,103,2,null,null,null,"S_34_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,20645.400000000001,null,20411.5,19661.310000000001,21863.38,"ug/l",null,103,2,null,null,null,"K_39_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,153.33000000000001,null,150,150,160,"ug/l",null,103,2,null,null,null,"Ca_44_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,333.34,null,440.01,200,360,null,1.2703792196494199,null,2,null,null,null,"Sc_45_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,4.67,null,12,2,0,"ug/l",null,103,2,null,null,null,"Ti_49_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,26.67,null,30,50,0,"ug/l",null,103,2,null,null,null,"V_51_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1420.06,null,1350.05,1650.08,1260.05,"ug/l",null,103,2,null,null,null,"Cr_52_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,376.67,null,410.01,390,330,"ug/l",null,103,2,null,null,null,"Mn_55_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6801.39,null,6201.24,6101.06,8101.88,"ug/l",null,103,2,null,null,null,"Fe_56_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,105,null,90,95,130,"ug/l",null,103,2,null,null,null,"Co_59_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,316.67,null,280,390.01,280,"ug/l",null,103,2,null,null,null,"Ni_60_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,375.01,null,300,425.01,400.01,"ug/l",null,103,2,null,null,null,"Cu_65_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,33.33,null,0,100,0,"ug/l",null,103,2,null,null,null,"Zn_66_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,8.33,null,10,15,0,"ug/l",null,103,2,null,null,null,"Ga_71_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,6.67,null,5,10,5,"ug/l",null,103,2,null,null,null,"Ge_74_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,53.33,null,40,60,60,"ug/l",null,103,2,null,null,null,"Br_79_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,150,null,200,200,50,"ug/l",null,103,2,null,null,null,"Rb_85_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Zr_90_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,10,null,15,5,10,"ug/l",null,103,2,null,null,null,"Mo_98_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,33.33,null,10,20,70,null,1.22828318560157E-2,null,2,null,null,null,"Rh_103_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.1399999999999999,null,-1.23,3.32,1.32,"ug/l",null,103,2,null,null,null,"Cd_114_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-15.67,null,-27,-15,-5,"ug/l",null,103,2,null,null,null,"In_115_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,31.67,null,45,25,25,"ug/l",null,103,2,null,null,null,"Sn_118_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,26.67,null,45,30,5,"ug/l",null,103,2,null,null,null,"Sb_121_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,66.67,null,200.01,0,0,"ug/l",null,103,2,null,null,null,"I_127_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,5,null,0,15,0,"ug/l",null,103,2,null,null,null,"Cs_133_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,{"$f":"N323&\"_\"&P323&\"_\"&G323","v":"Nd_146_005SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pt_195_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,133.33000000000001,null,50,150,200,"ug/l",null,103,2,null,null,null,"Au_197_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Hg_202_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,13.33,null,10,20,10,"ug/l",null,103,2,null,null,null,"Tl_205_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,66.67,null,50,60,90,"ug/l",null,103,2,null,null,null,"Pb_208_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,16.670000000000002,null,0,30,20,"ug/l",null,103,2,null,null,null,"Bi_209_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_005SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T11:28:55","Sample",1,"005SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_005SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,50,null,40,80,30,"ug/l",null,103,1,null,null,null,"Li_7_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,203.33,null,110,340,160,"ug/l",null,103,1,null,null,null,"B_10_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,776.68,null,850.02,790.02,690.01,"ug/l",null,103,1,null,null,null,"B_11_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,16660.939999999999,null,16677.61,17278.16,16027.05,"ug/l",null,103,1,null,null,null,"Al_27_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,21793.17,null,21673.040000000001,21973.39,21733.08,null,1.1273945858115999,null,1,null,null,null,"Sc_45_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,2036.79,null,2180.14,2070.13,1860.1,"ug/l",null,103,1,null,null,null,"Ge_74_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,153.33000000000001,null,120,150,190,null,1.8465573763753101E-2,null,1,null,null,null,"Rh_103_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,266.67,null,270,300,230,"ug/l",null,103,1,null,null,null,"Sn_118_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,33.33,null,0,70,30,"ug/l",null,103,1,null,null,null,"Sb_121_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,23.33,null,20,30,20,"ug/l",null,103,1,null,null,null,"Sb_123_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,63.33,null,70,50,70,"ug/l",null,103,1,null,null,null,"Te_125_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,53.33,null,50,100,10,"ug/l",null,103,1,null,null,null,"Hg_202_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,66.67,null,100,100,0,"ug/l",null,103,2,null,null,null,"Mg_24_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,33.33,null,40,0,60,"ug/l",null,103,2,null,null,null,"Al_27_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,60018.25,null,59566.77,59837.599999999999,60650.39,"ug/l",null,103,2,null,null,null,"Si_28_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"P_31_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,390.01,null,390.01,470.01,310,"ug/l",null,103,2,null,null,null,"S_34_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,21563.19,null,19911.53,22163.8,22614.240000000002,"ug/l",null,103,2,null,null,null,"K_39_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,206.67,null,180,240,200,"ug/l",null,103,2,null,null,null,"Ca_44_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,353.34,null,300,540.01,220,null,1.3466012106084799,null,2,null,null,null,"Sc_45_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,6,null,4,8,6,"ug/l",null,103,2,null,null,null,"Ti_49_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,46.67,null,80,30,30,"ug/l",null,103,2,null,null,null,"V_51_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1226.71,null,1170.04,1100.04,1410.06,"ug/l",null,103,2,null,null,null,"Cr_52_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,380,null,360,420.01,360,"ug/l",null,103,2,null,null,null,"Mn_55_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6134.44,null,6701.26,4500.59,7201.48,"ug/l",null,103,2,null,null,null,"Fe_56_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,96.67,null,105,90,95,"ug/l",null,103,2,null,null,null,"Co_59_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,323.33,null,390,240,340,"ug/l",null,103,2,null,null,null,"Ni_60_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,233.33,null,325,325,50,"ug/l",null,103,2,null,null,null,"Cu_65_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,66.67,null,0,0,200.01,"ug/l",null,103,2,null,null,null,"Zn_66_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Ga_71_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,41.67,null,15,50,60,"ug/l",null,103,2,null,null,null,"Ge_74_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"As_75_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,10,null,20,10,0,"ug/l",null,103,2,null,null,null,"Br_79_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,183.34,null,0,300.01,250,"ug/l",null,103,2,null,null,null,"Rb_85_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zr_90_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mo_98_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,6.67,null,10,0,10,null,2.4565663712031501E-3,null,2,null,null,null,"Rh_103_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0,null,0,-0.82,-0.27,"ug/l",null,103,2,null,null,null,"Cd_114_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,1.67,null,5,-18,-6,"ug/l",null,103,2,null,null,null,"In_115_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,13.33,null,0,30,10,"ug/l",null,103,2,null,null,null,"Sn_118_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Sb_121_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"I_127_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,{"$f":"N387&\"_\"&P387&\"_\"&G387","v":"Cs_133_006SMPL.d"}]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"Tb_159_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pt_195_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,83.33,null,200,0,50,"ug/l",null,103,2,null,null,null,"Au_197_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Hg_202_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"Tl_205_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,36.67,null,40,50,20,"ug/l",null,103,2,null,null,null,"Pb_208_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Bi_209_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:30:24","Sample",1,"006SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_006SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,50,null,30,70,50,"ug/l",null,103,1,null,null,null,"Li_7_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,230,null,180,270,240,"ug/l",null,103,1,null,null,null,"B_10_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,876.69,null,970.03,790.02,870.02,"ug/l",null,103,1,null,null,null,"B_11_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,16260.55,null,15966.95,16107.08,16707.62,"ug/l",null,103,1,null,null,null,"Al_27_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,22343.85,null,22273.79,23204.89,21552.86,null,1.15588194641737,null,1,null,null,null,"Sc_45_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,1620.08,null,1750.09,1540.07,1570.07,"ug/l",null,103,1,null,null,null,"Ge_74_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,180,null,250,200,90,null,2.1676977896579699E-2,null,1,null,null,null,"Rh_103_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,240,null,200,270,250,"ug/l",null,103,1,null,null,null,"Sn_118_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,6.67,null,0,20,0,"ug/l",null,103,1,null,null,null,"Sb_121_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,26.67,null,20,30,30,"ug/l",null,103,1,null,null,null,"Sb_123_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,53.33,null,100,30,30,"ug/l",null,103,1,null,null,null,"Te_125_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,13.33,null,30,0,10,"ug/l",null,103,1,null,null,null,"Hg_202_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,33.33,null,100,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,33.33,null,40,20,40,"ug/l",null,103,2,null,null,null,"Al_27_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,61673.75,null,63379.51,61633.54,60008.2,"ug/l",null,103,2,null,null,null,"Si_28_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"P_31_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,403.34,null,470.01,380.01,360,"ug/l",null,103,2,null,null,null,"S_34_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19493.990000000002,null,20361.61,20211.400000000001,17908.97,"ug/l",null,103,2,null,null,null,"K_39_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,186.67,null,290,130,140,"ug/l",null,103,2,null,null,null,"Ca_44_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,346.67,null,300,320,420.01,null,1.3211938802887899,null,2,null,null,null,"Sc_45_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,11.33,null,10,18,6,"ug/l",null,103,2,null,null,null,"Ti_49_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,60,null,70,80,30,"ug/l",null,103,2,null,null,null,"V_51_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1283.3800000000001,null,1180.04,1380.05,1290.05,"ug/l",null,103,2,null,null,null,"Cr_52_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,406.67,null,410.01,470.01,340,"ug/l",null,103,2,null,null,null,"Mn_55_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6768.12,null,8202.0499999999993,5100.78,7001.53,"ug/l",null,103,2,null,null,null,"Fe_56_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,91.67,null,100,95,80,"ug/l",null,103,2,null,null,null,"Co_59_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,300,null,300,300,300,"ug/l",null,103,2,null,null,null,"Ni_60_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,158.33000000000001,null,250,100,125,"ug/l",null,103,2,null,null,null,"Cu_65_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,8.33,null,0,20,5,"ug/l",null,103,2,null,null,null,"Ga_71_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,15,null,5,10,30,"ug/l",null,103,2,null,null,null,"Ge_74_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,5,5,0,"ug/l",null,103,2,null,null,null,"Se_78_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,16.670000000000002,null,10,40,0,"ug/l",null,103,2,null,null,null,"Br_79_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,133.33000000000001,null,200,50,150,"ug/l",null,103,2,null,null,null,"Rb_85_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zr_90_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Mo_98_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,0,null,0,0,0,null,0,null,2,null,null,null,"Rh_103_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0,null,0,-0.82,-0.55000000000000004,"ug/l",null,103,2,null,null,null,"Cd_114_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,0,null,0,-18,-12,"ug/l",null,103,2,null,null,null,"In_115_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,16.670000000000002,null,0,30,20,"ug/l",null,103,2,null,null,null,"Sn_118_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,{"$f":"N451&\"_\"&P451&\"_\"&G451","v":"Sb_121_007SMPL.d"}]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,166.67,null,100,200,200.01,"ug/l",null,103,2,null,null,null,"I_127_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,8.33,null,10,10,5,"ug/l",null,103,2,null,null,null,"Cs_133_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"Tb_159_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Ir_193_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Pt_195_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,183.34,null,100,200,250.01,"ug/l",null,103,2,null,null,null,"Au_197_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,16.670000000000002,null,10,40,0,"ug/l",null,103,2,null,null,null,"Hg_202_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,10,null,0,20,10,"ug/l",null,103,2,null,null,null,"Tl_205_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,60,null,40,50,90,"ug/l",null,103,2,null,null,null,"Pb_208_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,16.670000000000002,null,0,20,30,"ug/l",null,103,2,null,null,null,"Bi_209_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_007SMPL.d"]
["labman","deionized (DI) Water rinse bottle","CS_2019_10_22_JLS.b","2019-10-22T11:31:53","Sample",1,"007SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_007SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,126.67,null,90,70,220,"ug/l",null,103,1,null,null,null,"Li_7_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,3.33,null,0,10,0,"ug/l",null,103,1,null,null,null,"Be_9_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,573.34,null,600.01,510.01,610.01,"ug/l",null,103,1,null,null,null,"B_10_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,2013.44,null,1990.11,2200.13,1850.09,"ug/l",null,103,1,null,null,null,"B_11_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,54851.96,null,55503.88,55564.09,53487.92,"ug/l",null,103,1,null,null,null,"Al_27_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1795202.37,null,1820626.85,1859106.53,1705873.72,null,92.868611065584901,null,1,null,null,null,"Sc_45_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,3483.69,null,3350.34,3650.39,3450.35,"ug/l",null,103,1,null,null,null,"Ge_74_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,797661.73,null,776996,828321.71,787667.49,null,96.060532013412001,null,1,null,null,null,"Rh_103_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,2540.1999999999998,null,2610.1999999999998,2390.1799999999998,2620.21,"ug/l",null,103,1,null,null,null,"Sn_118_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,356.67,null,250,460.01,360,"ug/l",null,103,1,null,null,null,"Sb_121_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,296.67,null,350,380,160,"ug/l",null,103,1,null,null,null,"Sb_123_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,46.67,null,130,0,10,"ug/l",null,103,1,null,null,null,"Te_125_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,16.670000000000002,null,20,10,20,"ug/l",null,103,1,null,null,null,"Hg_202_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,100,null,100,0,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,66.67,null,40,100,60,"ug/l",null,103,2,null,null,null,"Al_27_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,55855.07,null,57570.36,54721.59,55273.26,"ug/l",null,103,2,null,null,null,"Si_28_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,46.67,null,20,80,40,"ug/l",null,103,2,null,null,null,"P_31_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,413.34,null,350,430.01,460.01,"ug/l",null,103,2,null,null,null,"S_34_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,20595.23,null,19961.05,19861.189999999999,21963.45,"ug/l",null,103,2,null,null,null,"K_39_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,2356.83,null,2220.14,2220.14,2630.2,"ug/l",null,103,2,null,null,null,"Ca_44_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,22560.83,null,22253.72,21753.15,23675.63,null,85.981581718104707,null,2,null,null,null,"Sc_45_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,9.33,null,6,14,8,"ug/l",null,103,2,null,null,null,"Ti_49_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,490.01,null,510.01,610.01,350,"ug/l",null,103,2,null,null,null,"V_51_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1476.73,null,1590.07,1580.08,1260.05,"ug/l",null,103,2,null,null,null,"Cr_52_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,570.01,null,530.01,560.01,620.01,"ug/l",null,103,2,null,null,null,"Mn_55_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,8402.09,null,9602.59,7801.86,7801.81,"ug/l",null,103,2,null,null,null,"Fe_56_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,110,null,190,65,75,"ug/l",null,103,2,null,null,null,"Co_59_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,510.01,null,510.01,520.01,500.01,"ug/l",null,103,2,null,null,null,"Ni_60_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,2033.46,null,2275.16,1725.09,2100.13,"ug/l",null,103,2,null,null,null,"Cu_65_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,4133.93,null,3800.43,3800.59,4800.76,"ug/l",null,103,2,null,null,null,"Zn_66_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,10,null,5,20,5,"ug/l",null,103,2,null,null,null,"Ga_71_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,15,null,15,5,25,"ug/l",null,103,2,null,null,null,"Ge_74_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"As_75_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Se_78_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,516.67999999999995,null,580.01,440.01,530.01,"ug/l",null,103,2,null,null,null,"Br_79_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,50,null,50,0,100,"ug/l",null,103,2,null,null,null,"Rb_85_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,703.35,null,670.01,760.02,680.02,"ug/l",null,103,2,null,null,null,"Sr_88_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,20,null,20,10,30,"ug/l",null,103,2,null,null,null,"Y_89_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,60,null,30,80,70,"ug/l",null,103,2,null,null,null,"Zr_90_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,420,null,525.01,330,405,"ug/l",null,103,2,null,null,null,"Mo_98_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,247178.75,null,239757.01,252264.7,249514.55,null,91.081651967187597,null,2,null,null,null,"Rh_103_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,66.67,null,66.67,0,133.33000000000001,"ug/l",null,103,2,null,null,null,"Ag_109_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,9.4,null,13.55,12.59,2.0499999999999998,"ug/l",null,103,2,null,null,null,{"$f":"N515&\"_\"&P515&\"_\"&G515","v":"Cd_114_008SMPL.d"}]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-65.33,null,-54,-75,-67,"ug/l",null,103,2,null,null,null,"In_115_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,120,null,90,125,145,"ug/l",null,103,2,null,null,null,"Sn_118_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,63.33,null,85,30,75,"ug/l",null,103,2,null,null,null,"Sb_121_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,400.02,null,0,500.02,700.03,"ug/l",null,103,2,null,null,null,"I_127_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,8.33,null,5,15,5,"ug/l",null,103,2,null,null,null,"Cs_133_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,153.33000000000001,null,180,145,135,"ug/l",null,103,2,null,null,null,"Ba_137_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,63.33,null,30,80,80,"ug/l",null,103,2,null,null,null,"Ce_140_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Nd_146_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Tb_159_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,60,null,30,60,90,"ug/l",null,103,2,null,null,null,"Ir_193_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"Pt_195_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,2516.88,null,2850.27,2150.15,2550.2199999999998,"ug/l",null,103,2,null,null,null,"Au_197_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Hg_202_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"Tl_205_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,510.01,null,580.01,530.01,420.01,"ug/l",null,103,2,null,null,null,"Pb_208_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,70,null,70,50,90,"ug/l",null,103,2,null,null,null,"Bi_209_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Th_232_008SMPL.d"]
["labman","Reagent blank1","CS_2019_10_22_JLS.b","2019-10-22T11:33:21","Sample",1,"008SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,26.67,null,20,30,30,"ug/l",null,103,2,null,null,null,"U_238_008SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,73.33,null,70,60,90,"ug/l",null,103,1,null,null,null,"Li_7_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,400.01,null,300,450.01,450.01,"ug/l",null,103,1,null,null,null,"B_10_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,1770.09,null,1840.1,1810.09,1660.08,"ug/l",null,103,1,null,null,null,"B_11_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,34335.550000000003,null,33670.99,33380.379999999997,35955.29,"ug/l",null,103,1,null,null,null,"Al_27_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1794832.37,null,1729744.19,1816276.54,1838476.38,null,92.849470562447095,null,1,null,null,null,"Sc_45_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,3093.61,null,3040.26,3450.34,2790.22,"ug/l",null,103,1,null,null,null,"Ge_74_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,811800.04,null,810195.07,809522.8,815682.25,null,97.763175130680906,null,1,null,null,null,"Rh_103_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,2246.8200000000002,null,2360.17,2460.1799999999998,1920.11,"ug/l",null,103,1,null,null,null,"Sn_118_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,206.67,null,310,170,140,"ug/l",null,103,1,null,null,null,"Sb_121_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,176.67,null,160,150,220,"ug/l",null,103,1,null,null,null,"Sb_123_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,76.67,null,50,100,80,"ug/l",null,103,1,null,null,null,"Te_125_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,23.33,null,10,50,10,"ug/l",null,103,1,null,null,null,"Hg_202_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,133.33000000000001,null,100,0,300,"ug/l",null,103,2,null,null,null,"Mg_24_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,43.33,null,90,40,0,"ug/l",null,103,2,null,null,null,"Al_27_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58540.14,null,58021.89,58994.89,58603.65,"ug/l",null,103,2,null,null,null,"Si_28_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,10,null,30,0,0,"ug/l",null,103,2,null,null,null,"P_31_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,423.34,null,420.01,450.01,400,"ug/l",null,103,2,null,null,null,"S_34_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19193.759999999998,null,19160.240000000002,17658.810000000001,20762.23,"ug/l",null,103,2,null,null,null,"K_39_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,290,null,330,230,310,"ug/l",null,103,2,null,null,null,"Ca_44_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,24409.99,null,25237.82,24256.33,23735.82,null,93.028901854531895,null,2,null,null,null,"Sc_45_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,11.33,null,18,6,10,"ug/l",null,103,2,null,null,null,"Ti_49_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,373.34,null,420.01,360,340,"ug/l",null,103,2,null,null,null,"V_51_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1466.73,null,1250.04,1700.09,1450.06,"ug/l",null,103,2,null,null,null,"Cr_52_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,463.34,null,540.01,470.01,380,"ug/l",null,103,2,null,null,null,"Mn_55_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7768.55,null,7001.56,6301.16,10002.93,"ug/l",null,103,2,null,null,null,"Fe_56_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,88.33,null,65,85,115,"ug/l",null,103,2,null,null,null,"Co_59_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,410.01,null,470.01,420.01,340,"ug/l",null,103,2,null,null,null,"Ni_60_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,600.01,null,750.02,400.01,650.01,"ug/l",null,103,2,null,null,null,"Cu_65_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,300.01,null,500.02,0,400.01,"ug/l",null,103,2,null,null,null,"Zn_66_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,5,null,0,5,10,"ug/l",null,103,2,null,null,null,"Ga_71_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,46.67,null,40,55,45,"ug/l",null,103,2,null,null,null,"Ge_74_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"As_75_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,500.01,null,560.01,510.01,430.01,"ug/l",null,103,2,null,null,null,"Br_79_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,50,null,0,50,100,"ug/l",null,103,2,null,null,null,"Rb_85_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,13.33,null,30,0,10,"ug/l",null,103,2,null,null,null,"Sr_88_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"Zr_90_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,146.66999999999999,null,150,140,150,"ug/l",null,103,2,null,null,null,"Mo_98_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,257207.43,null,253889.61,259876.97,257855.72,null,94.777069672521307,null,2,null,null,null,{"$f":"N579&\"_\"&P579&\"_\"&G579","v":"Rh_103_009SMPL.d"}]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,11.11,null,33.33,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0.43,null,-3.13,0.59,3.82,"ug/l",null,103,2,null,null,null,"Cd_114_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-60.67,null,-69,-75,-38,"ug/l",null,103,2,null,null,null,"In_115_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,106.67,null,115,125,80,"ug/l",null,103,2,null,null,null,"Sn_118_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sb_121_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,366.68,null,300.01,700.03,100,"ug/l",null,103,2,null,null,null,"I_127_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"Cs_133_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,38.33,null,40,25,50,"ug/l",null,103,2,null,null,null,"Ba_137_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"W_184_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,43.33,null,80,40,10,"ug/l",null,103,2,null,null,null,"Ir_193_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,13.33,null,30,10,0,"ug/l",null,103,2,null,null,null,"Pt_195_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1483.42,null,1700.1,1600.09,1150.06,"ug/l",null,103,2,null,null,null,"Au_197_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,26.67,null,0,70,10,"ug/l",null,103,2,null,null,null,"Hg_202_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,10,null,0,10,20,"ug/l",null,103,2,null,null,null,"Tl_205_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,173.33,null,210,180,130,"ug/l",null,103,2,null,null,null,"Pb_208_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,10,null,30,0,0,"ug/l",null,103,2,null,null,null,"Bi_209_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Th_232_009SMPL.d"]
["labman","Reagent blank2","CS_2019_10_22_JLS.b","2019-10-22T11:34:53","Sample",1,"009SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_009SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,63.33,null,50,50,90,"ug/l",null,103,1,null,null,null,"Li_7_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,6.67,null,20,0,0,"ug/l",null,103,1,null,null,null,"Be_9_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,453.34,null,440.01,460.01,460.01,"ug/l",null,103,1,null,null,null,"B_10_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,1683.41,null,1690.08,1560.07,1800.09,"ug/l",null,103,1,null,null,null,"B_11_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,29333.48,null,29293.46,29093.08,29613.9,"ug/l",null,103,1,null,null,null,"Al_27_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1912102.73,null,1981705.28,1958339.5,1796263.41,null,98.916048712398606,null,1,null,null,null,"Sc_45_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,3073.61,null,3000.26,3250.31,2970.26,"ug/l",null,103,1,null,null,null,"Ge_74_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,828868.3,null,810491.16,836202.1,839911.63,null,99.818665255661301,null,1,null,null,null,"Rh_103_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1880.11,null,2010.12,1970.12,1660.08,"ug/l",null,103,1,null,null,null,"Sn_118_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,170,null,160,160,190,"ug/l",null,103,1,null,null,null,"Sb_121_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,126.67,null,140,80,160,"ug/l",null,103,1,null,null,null,"Sb_123_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,23.33,null,30,10,30,"ug/l",null,103,1,null,null,null,"Te_125_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,23.33,null,20,10,40,"ug/l",null,103,1,null,null,null,"Hg_202_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,33.33,null,100,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,30,null,40,20,30,"ug/l",null,103,2,null,null,null,"Al_27_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,57560.480000000003,null,57339.74,57600.63,57741.07,"ug/l",null,103,2,null,null,null,"Si_28_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,20,null,30,20,10,"ug/l",null,103,2,null,null,null,"P_31_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,353.34,null,370,290,400.01,"ug/l",null,103,2,null,null,null,"S_34_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19794.45,null,21512.99,17758.79,20111.580000000002,"ug/l",null,103,2,null,null,null,"K_39_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,330,null,240,410.01,340,"ug/l",null,103,2,null,null,null,"Ca_44_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,24122.84,null,24196.29,24336.47,23835.75,null,91.934531915672096,null,2,null,null,null,"Sc_45_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,14,null,18,16,8,"ug/l",null,103,2,null,null,null,"Ti_49_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,346.67,null,460.01,320,260,"ug/l",null,103,2,null,null,null,"V_51_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1460.06,null,1460.06,1530.07,1390.06,"ug/l",null,103,2,null,null,null,"Cr_52_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,400.01,null,420.01,370,410.01,"ug/l",null,103,2,null,null,null,"Mn_55_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6668.04,null,5400.86,8402.0300000000007,6201.23,"ug/l",null,103,2,null,null,null,"Fe_56_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,103.33,null,125,100,85,"ug/l",null,103,2,null,null,null,"Co_59_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,336.67,null,190,430.01,390,"ug/l",null,103,2,null,null,null,"Ni_60_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,508.34,null,450.01,325,750.02,"ug/l",null,103,2,null,null,null,"Cu_65_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,100,null,0,100,200,"ug/l",null,103,2,null,null,null,"Zn_66_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Ga_71_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,28.33,null,25,35,25,"ug/l",null,103,2,null,null,null,"Ge_74_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"As_75_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,490.01,null,510.01,500.01,460.01,"ug/l",null,103,2,null,null,null,"Br_79_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,50,null,100,0,50,"ug/l",null,103,2,null,null,null,"Rb_85_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,20,null,30,20,10,"ug/l",null,103,2,null,null,null,"Sr_88_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,13.33,null,10,20,10,"ug/l",null,103,2,null,null,null,{"$f":"N643&\"_\"&P643&\"_\"&G643","v":"Zr_90_010SMPL.d"}]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,48.33,null,55,70,20,"ug/l",null,103,2,null,null,null,"Mo_98_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,259497.36,null,261703.04000000001,258800.39,257988.65,null,95.620873198799501,null,2,null,null,null,"Rh_103_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,-3.18,null,-2.86,-2.3199999999999998,-4.3600000000000003,"ug/l",null,103,2,null,null,null,"Cd_114_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-65,null,-63,-41,-91,"ug/l",null,103,2,null,null,null,"In_115_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,116.67,null,105,85,160,"ug/l",null,103,2,null,null,null,"Sn_118_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,10,null,5,5,20,"ug/l",null,103,2,null,null,null,"Sb_121_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,133.34,null,100,200.01,100,"ug/l",null,103,2,null,null,null,"I_127_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Cs_133_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Ba_137_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,16.670000000000002,null,40,0,10,"ug/l",null,103,2,null,null,null,"W_184_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,56.67,null,30,50,90,"ug/l",null,103,2,null,null,null,"Ir_193_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,16.670000000000002,null,0,40,10,"ug/l",null,103,2,null,null,null,"Pt_195_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1150.06,null,1000.04,1100.06,1350.08,"ug/l",null,103,2,null,null,null,"Au_197_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,20,null,40,10,10,"ug/l",null,103,2,null,null,null,"Hg_202_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Tl_205_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,46.67,null,40,80,20,"ug/l",null,103,2,null,null,null,"Pb_208_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,13.33,null,20,0,20,"ug/l",null,103,2,null,null,null,"Bi_209_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Th_232_010SMPL.d"]
["labman","Reagent blank3","CS_2019_10_22_JLS.b","2019-10-22T11:36:21","Sample",1,"010SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_010SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,null,null,66.67,null,120,30,50,"ug/l",null,103,1,null,null,null,"Li_7_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,null,null,0,null,0,0,0,"ug/l",null,103,1,null,null,null,"Be_9_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,null,null,363.34,null,450.01,380,260,"ug/l",null,103,1,null,null,null,"B_10_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,null,null,1290.04,null,1410.05,1220.04,1240.04,"ug/l",null,103,1,null,null,null,"B_11_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,30962.95,null,30465.38,30585.56,31837.9,"ug/l",null,103,1,null,null,null,"Al_27_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1947013.56,null,1986144.97,1954193.56,1900702.16,null,100.72204042843499,null,1,null,null,null,"Sc_45_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,3033.6,null,2740.21,3110.28,3250.3,"ug/l",null,103,1,null,null,null,"Ge_74_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,813665.4,null,820446.47,812021.86,808527.88,null,97.9878164626017,null,1,null,null,null,"Rh_103_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1453.4,null,1490.07,1460.06,1410.06,"ug/l",null,103,1,null,null,null,"Sn_118_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,146.66999999999999,null,160,160,120,"ug/l",null,103,1,null,null,null,"Sb_121_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,93.33,null,90,110,80,"ug/l",null,103,1,null,null,null,"Sb_123_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,30,null,30,20,40,"ug/l",null,103,1,null,null,null,"Te_125_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,30,null,20,70,0,"ug/l",null,103,1,null,null,null,"Hg_202_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,66.67,null,0,0,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,null,null,20,null,30,20,10,"ug/l",null,103,2,null,null,null,"Al_27_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,57256.06,null,56306.46,57861.25,57600.46,"ug/l",null,103,2,null,null,null,"Si_28_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,66.67,null,50,80,70,"ug/l",null,103,2,null,null,null,"P_31_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,476.68,null,520.01,450.01,460.01,"ug/l",null,103,2,null,null,null,"S_34_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18693.03,null,19010.060000000001,18259.240000000002,18809.79,"ug/l",null,103,2,null,null,null,"K_39_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,266.67,null,190,370.01,240,"ug/l",null,103,2,null,null,null,"Ca_44_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,24643.599999999999,null,23996.04,24997.57,24937.18,null,93.919200116263994,null,2,null,null,null,"Sc_45_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,12,null,12,18,6,"ug/l",null,103,2,null,null,null,"Ti_49_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,496.68,null,410.01,590.01,490.01,"ug/l",null,103,2,null,null,null,"V_51_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1643.42,null,1730.09,1370.06,1830.1,"ug/l",null,103,2,null,null,null,"Cr_52_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,336.67,null,330,410.01,270,"ug/l",null,103,2,null,null,null,"Mn_55_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6301.23,null,7601.79,5600.96,5700.95,"ug/l",null,103,2,null,null,null,"Fe_56_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,93.33,null,140,70,70,"ug/l",null,103,2,null,null,null,"Co_59_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,826.69,null,940.03,690.02,850.02,"ug/l",null,103,2,null,null,null,"Ni_60_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,591.67999999999995,null,475.01,600.01,700.02,"ug/l",null,103,2,null,null,null,"Cu_65_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,100,null,300.01,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,6.67,null,0,15,5,"ug/l",null,103,2,null,null,null,"Ga_71_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,30,null,30,40,20,"ug/l",null,103,2,null,null,null,"Ge_74_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"As_75_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,536.67999999999995,null,410.01,610.01,590.01,"ug/l",null,103,2,null,null,null,"Br_79_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,33.33,null,50,50,0,"ug/l",null,103,2,null,null,null,{"$f":"N707&\"_\"&P707&\"_\"&G707","v":"Rb_85_011SMPL.d"}]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,16.670000000000002,null,30,20,0,"ug/l",null,103,2,null,null,null,"Sr_88_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,13.33,null,0,20,20,"ug/l",null,103,2,null,null,null,"Zr_90_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,25,null,30,15,30,"ug/l",null,103,2,null,null,null,"Mo_98_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,259701.02,null,256852.22,265967.81,256283.04,null,95.695920073156501,null,2,null,null,null,"Rh_103_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,5,null,5,0,10,"ug/l",null,103,2,null,null,null,"Pd_105_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,-0.39,null,-3,-1.36,3.18,"ug/l",null,103,2,null,null,null,"Cd_114_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-38,null,-66,-30,-18,"ug/l",null,103,2,null,null,null,"In_115_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,63.33,null,110,50,30,"ug/l",null,103,2,null,null,null,"Sn_118_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Sb_121_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,700.03,null,600.01,900.03,600.04,"ug/l",null,103,2,null,null,null,"I_127_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,8.33,null,0,25,0,"ug/l",null,103,2,null,null,null,"Cs_133_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,16.670000000000002,null,10,25,15,"ug/l",null,103,2,null,null,null,"Ba_137_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"W_184_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,46.67,null,40,70,30,"ug/l",null,103,2,null,null,null,"Ir_193_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,null,null,13.33,null,20,0,20,"ug/l",null,103,2,null,null,null,"Pt_195_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1033.3800000000001,null,750.02,1250.07,1100.04,"ug/l",null,103,2,null,null,null,"Au_197_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Hg_202_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Tl_205_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,86.67,null,90,90,80,"ug/l",null,103,2,null,null,null,"Pb_208_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,33.33,null,50,40,10,"ug/l",null,103,2,null,null,null,"Bi_209_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_011SMPL.d"]
["labman","Reagent blank4","CS_2019_10_22_JLS.b","2019-10-22T11:37:49","Sample",1,"011SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_011SMPL.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-2.5000000000000001E-2,null,63.33,null,90,70,30,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.4339999999999999,null,356.67,null,260,460.01,350,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.5349999999999999,null,1580.07,null,1340.05,1750.09,1650.08,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,5.0999999999999997E-2,24.05,34519.18,24.05,33650.879999999997,35123.65,34783.01,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1933056.12,null,1984277,1905764.82,1909126.53,null,100,null,1,null,null,null,"Sc_45_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,1.6E-2,146.43,3296.98,146.43,3060.27,3560.36,3270.31,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,830374.05,null,821147.96,823703.74,846270.46,null,100,null,1,null,null,null,"Rh_103_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1176.71,null,1020.03,1340.05,1170.04,"ug/l",null,103,1,null,null,null,"Sn_118_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,150,null,170,120,160,"ug/l",null,103,1,null,null,null,"Sb_121_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,86.67,null,110,80,70,"ug/l",null,103,1,null,null,null,"Sb_123_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-3.286,null,26.67,null,40,20,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,9.843,0.73,10,0.73,10,10,10,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,66.67,null,0,200.01,0,"ug/l",null,103,2,null,null,null,"Mg_24_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,-9.9000000000000005E-2,null,36.67,null,10,50,50,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58593.63,null,59205.5,59155.46,57419.93,"ug/l",null,103,2,null,null,null,"Si_28_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,53.33,null,40,60,60,"ug/l",null,103,2,null,null,null,"P_31_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,450.01,null,410,540.01,400.01,"ug/l",null,103,2,null,null,null,"S_34_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18976.82,null,19410.689999999999,17608.669999999998,19911.11,"ug/l",null,103,2,null,null,null,"K_39_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,153.33000000000001,null,120,240,100,"ug/l",null,103,2,null,null,null,"Ca_44_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,26239.15,null,26379.25,25137.56,27200.63,null,100,null,2,null,null,null,"Sc_45_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,9.33,null,10,8,10,"ug/l",null,103,2,null,null,null,"Ti_49_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,426.67,null,360,530.01,390.01,"ug/l",null,103,2,null,null,null,"V_51_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1376.72,null,1410.06,1230.04,1490.07,"ug/l",null,103,2,null,null,null,"Cr_52_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,310,null,250,430.01,250,"ug/l",null,103,2,null,null,null,"Mn_55_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6034.51,null,5200.8599999999997,4500.6000000000004,8402.08,"ug/l",null,103,2,null,null,null,"Fe_56_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,73.33,null,75,75,70,"ug/l",null,103,2,null,null,null,"Co_59_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,880.03,null,790.02,1100.04,750.02,"ug/l",null,103,2,null,null,null,"Ni_60_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,775.02,null,675.01,750.02,900.03,"ug/l",null,103,2,null,null,null,"Cu_65_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,100,null,0,0,300.01,"ug/l",null,103,2,null,null,null,"Zn_66_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,18.329999999999998,null,10,20,25,"ug/l",null,103,2,null,null,null,"Ga_71_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,28.33,null,35,30,20,"ug/l",null,103,2,null,null,null,"Ge_74_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,10,null,10,0,20,"ug/l",null,103,2,null,null,null,{"$f":"N771&\"_\"&P771&\"_\"&G771","v":"As_75_012CALS.d"}]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Se_78_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,490.01,null,420.01,590.01,460.01,"ug/l",null,103,2,null,null,null,"Br_79_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,116.67,null,200,100,50,"ug/l",null,103,2,null,null,null,"Rb_85_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"Sr_88_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,20,null,20,10,30,"ug/l",null,103,2,null,null,null,"Zr_90_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,26.67,null,40,20,20,"ug/l",null,103,2,null,null,null,"Mo_98_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,271381.5,null,268242.18,274470.65999999997,271431.65999999997,null,100,null,2,null,null,null,"Rh_103_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,33.33,null,100,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0.85,null,-0.04,4.2300000000000004,-1.64,"ug/l",null,103,2,null,null,null,"Cd_114_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-36.67,null,-45,-29,-36,"ug/l",null,103,2,null,null,null,"In_115_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,66.67,null,75,65,60,"ug/l",null,103,2,null,null,null,"Sn_118_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Sb_121_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,200,null,100,400.01,100,"ug/l",null,103,2,null,null,null,"I_127_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,16.670000000000002,null,30,5,15,"ug/l",null,103,2,null,null,null,"Cs_133_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,38.33,null,45,20,50,"ug/l",null,103,2,null,null,null,"Ba_137_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Ce_140_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,16.670000000000002,null,10,10,30,"ug/l",null,103,2,null,null,null,"W_184_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,50,null,50,100,0,"ug/l",null,103,2,null,null,null,"Ir_193_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,3.1E-2,5991.73,3.33,5991.73,10,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,966.71,null,1000.05,1000.03,900.04,"ug/l",null,103,2,null,null,null,"Au_197_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,13.33,null,0,40,0,"ug/l",null,103,2,null,null,null,"Hg_202_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tl_205_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,80,null,10,100,130,"ug/l",null,103,2,null,null,null,"Pb_208_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Bi_209_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Th_232_012CALS.d"]
["labman","Reagent blank5","CS_2019_10_22_JLS.b","2019-10-22T11:39:17","CalStd",1,"012CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_012CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,50,7.08,1599347.79,7.08,1534850.45,1552296.54,1710896.38,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,1.25,3.46,13007.89,3.46,12894.47,12944.52,13184.68,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,125,2.96,137617.37,2.96,138675.41,138020,136156.71,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,125,2.58,631654.61,2.58,656820.54,625691.28,612452.02,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,50,6.76,3248556.72,6.76,3566844.01,3161283.08,3017543.08,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1943295.75,null,1982380.6,1895262.63,1952244.03,null,100.529712333666,null,1,null,null,null,"Sc_45_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,1,10.59,14866.3,10.59,13885.49,15076.47,15636.94,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,853204.18,null,880074.05,828143.27,851395.22,null,102.749378617386,null,1,null,null,null,"Rh_103_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1243.3800000000001,null,1440.06,1300.05,990.03,"ug/l",null,103,1,null,null,null,"Sn_118_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,216.67,null,330,150,170,"ug/l",null,103,1,null,null,null,"Sb_121_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,136.66999999999999,null,100,90,220,"ug/l",null,103,1,null,null,null,"Sb_123_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,1,240.37,66.67,240.37,80,40,80,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,3,281.61,26.67,281.61,50,20,10,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,219068.39,null,219001.23,208173.66,230030.28,"ug/l",null,103,2,null,null,null,"Mg_24_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,50,3.66,1820.09,3.66,1690.08,1920.1,1850.1,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,56948.44,null,56908.28,56948.49,56988.54,"ug/l",null,103,2,null,null,null,"Si_28_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,13605.11,null,13144.8,13805.27,13865.27,"ug/l",null,103,2,null,null,null,"P_31_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,483.34,null,490.01,630.01,330,"ug/l",null,103,2,null,null,null,"S_34_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19544.150000000001,null,19811.240000000002,18309.53,20511.689999999999,"ug/l",null,103,2,null,null,null,"K_39_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,4530.58,null,4780.6400000000003,4150.4799999999996,4660.62,"ug/l",null,103,2,null,null,null,"Ca_44_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,24530.13,null,23355.41,25317.77,24917.21,null,93.486767354222906,null,2,null,null,null,"Sc_45_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,19.329999999999998,null,20,16,22,"ug/l",null,103,2,null,null,null,"Ti_49_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,14929.57,null,15556.77,14395.79,14836.16,"ug/l",null,103,2,null,null,null,"V_51_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,24116.28,null,22764.49,24907.32,24677.03,"ug/l",null,103,2,null,null,null,"Cr_52_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,6377.82,null,6051.03,6461.2,6621.23,"ug/l",null,103,2,null,null,null,"Mn_55_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,6467.9,null,6501.18,7601.68,5300.83,"ug/l",null,103,2,null,null,null,"Fe_56_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,22771.21,null,22173.81,23395.32,22744.5,"ug/l",null,103,2,null,null,null,"Co_59_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,433.34,null,420.01,480.01,400,"ug/l",null,103,2,null,null,null,"Ni_60_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,216656.23,null,214892.17,215829.55,219246.98,"ug/l",null,103,2,null,null,null,"Cu_65_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,533.35,null,600.01,500.02,500.01,"ug/l",null,103,2,null,null,null,{"$f":"N835&\"_\"&P835&\"_\"&G835","v":"Zn_66_013CALS.d"}]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,3228.64,null,3245.3,2945.25,3495.36,"ug/l",null,103,2,null,null,null,"Ga_71_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,238.33,null,325,220,170,"ug/l",null,103,2,null,null,null,"Ge_74_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,22557.81,null,22234.04,22834.87,22604.53,"ug/l",null,103,2,null,null,null,"As_75_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3847.09,null,3840.43,3835.42,3865.43,"ug/l",null,103,2,null,null,null,"Se_78_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,570.01,null,570.01,490.01,650.01,"ug/l",null,103,2,null,null,null,"Br_79_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,261902.69,null,250736.68,265553.77,269417.61,"ug/l",null,103,2,null,null,null,"Rb_85_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,58411.16,null,57605.04,58648.71,58979.73,"ug/l",null,103,2,null,null,null,"Sr_88_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Y_89_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,26.67,null,50,20,10,"ug/l",null,103,2,null,null,null,"Zr_90_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,2281.8200000000002,null,2260.15,2295.15,2290.15,"ug/l",null,103,2,null,null,null,"Mo_98_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,272969.55,null,261047.38,276980.21000000002,280881.07,null,100.585173762152,null,2,null,null,null,"Rh_103_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,138381.79999999999,null,135634.56,136075.85,143435,"ug/l",null,103,2,null,null,null,"Ag_109_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,37.68,null,35.82,31.14,46.09,"ug/l",null,103,2,null,null,null,"Cd_114_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,15434.37,null,15138.76,15489.1,15675.25,"ug/l",null,103,2,null,null,null,"In_115_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,85,null,80,105,70,"ug/l",null,103,2,null,null,null,"Sn_118_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,10,null,0,5,25,"ug/l",null,103,2,null,null,null,"Sb_121_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,1033.43,null,1600.17,1300.1099999999999,200,"ug/l",null,103,2,null,null,null,"I_127_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,786.69,null,780.02,790.02,790.02,"ug/l",null,103,2,null,null,null,"Cs_133_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,22156.32,null,20822.93,22670.32,22975.72,"ug/l",null,103,2,null,null,null,"Ba_137_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,16.670000000000002,null,40,0,10,"ug/l",null,103,2,null,null,null,"Ce_140_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,63.33,null,60,70,60,"ug/l",null,103,2,null,null,null,"Gd_156_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,40,null,40,40,40,"ug/l",null,103,2,null,null,null,"W_184_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,60,null,50,70,60,"ug/l",null,103,2,null,null,null,"Ir_193_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,1,352.35,6.67,352.35,0,0,20,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1566.76,null,1500.09,1450.08,1750.12,"ug/l",null,103,2,null,null,null,"Au_197_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Hg_202_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,70482.5,null,70003.73,70586.259999999995,70857.52,"ug/l",null,103,2,null,null,null,"Tl_205_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,479627.78,null,471988.24,480371.24,486523.86,"ug/l",null,103,2,null,null,null,"Pb_208_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,65555.17,null,63737.599999999999,64761.919999999998,68166,"ug/l",null,103,2,null,null,null,"Bi_209_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,75110.42,null,74286.460000000006,73965.039999999994,77079.77,"ug/l",null,103,2,null,null,null,"Th_232_013CALS.d"]
["labman","Standard NMS-59","CS_2019_10_22_JLS.b","2019-10-22T11:41:11","CalStd",1,"013CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,21211.18,null,21845.33,20143.12,21645.08,"ug/l",null,103,2,null,null,null,"U_238_013CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,2.1999999999999999E-2,34.68,1560.07,34.68,1540.06,1790.09,1350.05,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,173.21,3.33,173.21,10,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,3.2879999999999998,14.83,5507.49,14.83,5870.94,5660.86,4990.68,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,3.2440000000000002,6.36,25494.21,6.36,26318.69,25337.33,24826.61,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,2.8000000000000001E-2,22.16,34091.730000000003,22.16,34352.199999999997,33360.36,34562.629999999997,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,2010744.19,null,1991575.28,1935590.44,2105066.84,null,104.018924713576,null,1,null,null,null,"Sc_45_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-6.0000000000000001E-3,null,3140.28,null,2960.25,3330.32,3130.28,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,855917.7,null,851401.08,845851.32,870500.69,null,103.076161066304,null,1,null,null,null,"Rh_103_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1006.7,null,910.03,1030.03,1080.03,"ug/l",null,103,1,null,null,null,"Sn_118_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,123.33,null,80,110,180,"ug/l",null,103,1,null,null,null,"Sb_121_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,123.33,null,170,130,70,"ug/l",null,103,1,null,null,null,"Sb_123_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-3.758,null,23.33,null,20,20,30,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,8.5129999999999999,58.01,13.33,58.01,20,20,0,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,66.67,null,0,0,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,-0.193,null,33.33,null,30,50,20,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58048.67,null,58583.62,56396.94,59165.45,"ug/l",null,103,2,null,null,null,"Si_28_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,33.33,null,10,20,70,"ug/l",null,103,2,null,null,null,"P_31_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,380.01,null,280,430.01,430.01,"ug/l",null,103,2,null,null,null,"S_34_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18676.47,null,19260.39,17558.54,19210.48,"ug/l",null,103,2,null,null,null,"K_39_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,286.67,null,310,290,260,"ug/l",null,103,2,null,null,null,"Ca_44_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25952.04,null,25978.720000000001,25918.68,25958.71,null,98.905795208787296,null,2,null,null,null,"Sc_45_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,10.67,null,4,16,12,"ug/l",null,103,2,null,null,null,"Ti_49_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,413.34,null,430.01,330,480.01,"ug/l",null,103,2,null,null,null,"V_51_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1656.74,null,1650.08,1690.08,1630.07,"ug/l",null,103,2,null,null,null,"Cr_52_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,363.34,null,390.01,410.01,290,"ug/l",null,103,2,null,null,null,"Mn_55_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,5967.76,null,6501.23,5400.92,6001.14,"ug/l",null,103,2,null,null,null,"Fe_56_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,78.33,null,80,55,100,"ug/l",null,103,2,null,null,null,{"$f":"N899&\"_\"&P899&\"_\"&G899","v":"Co_59_014CALS.d"}]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,666.68,null,790.02,560.01,650.01,"ug/l",null,103,2,null,null,null,"Ni_60_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,875.03,null,1050.04,750.02,825.03,"ug/l",null,103,2,null,null,null,"Cu_65_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,200,null,300.01,100,200,"ug/l",null,103,2,null,null,null,"Zn_66_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,18.329999999999998,null,5,30,20,"ug/l",null,103,2,null,null,null,"Ga_71_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,41.67,null,40,35,50,"ug/l",null,103,2,null,null,null,"Ge_74_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,16.670000000000002,null,20,10,20,"ug/l",null,103,2,null,null,null,"As_75_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Se_78_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,430.01,null,500.01,350,440.01,"ug/l",null,103,2,null,null,null,"Br_79_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,416.68,null,100,550.03,600.02,"ug/l",null,103,2,null,null,null,"Rb_85_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,60,null,100,50,30,"ug/l",null,103,2,null,null,null,"Sr_88_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,10,null,0,30,0,"ug/l",null,103,2,null,null,null,"Zr_90_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,58.15,null,35,44.45,95,"ug/l",null,103,2,null,null,null,"Mo_98_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Ru_101_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,273678.46000000002,null,269958.98,270863.59000000003,280212.81,null,100.8463952038,null,2,null,null,null,"Rh_103_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,1211.1600000000001,null,1533.41,1000.04,1100.04,"ug/l",null,103,2,null,null,null,"Ag_109_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,-1.73,null,-1.5,-0.82,-2.86,"ug/l",null,103,2,null,null,null,"Cd_114_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-19.670000000000002,null,-23,22,-58,"ug/l",null,103,2,null,null,null,"In_115_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,63.33,null,55,30,105,"ug/l",null,103,2,null,null,null,"Sn_118_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Sb_121_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,633.36,null,600.01,200.01,1100.05,"ug/l",null,103,2,null,null,null,"I_127_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,10,null,20,0,10,"ug/l",null,103,2,null,null,null,"Cs_133_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"Ba_137_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,10,null,0,30,0,"ug/l",null,103,2,null,null,null,"W_184_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,53.33,null,60,80,20,"ug/l",null,103,2,null,null,null,"Ir_193_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,-1.034,null,0,null,0,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,766.69,null,750.02,550.02,1000.04,"ug/l",null,103,2,null,null,null,"Au_197_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Hg_202_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,93.33,null,90,130,60,"ug/l",null,103,2,null,null,null,"Tl_205_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,396.67,null,340,460.01,390.01,"ug/l",null,103,2,null,null,null,"Pb_208_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,1910.12,null,1930.12,1930.13,1870.11,"ug/l",null,103,2,null,null,null,"Bi_209_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,3943.84,null,3710.45,4230.58,3890.49,"ug/l",null,103,2,null,null,null,"Th_232_014CALS.d"]
["labman","Reagent blank6","CS_2019_10_22_JLS.b","2019-10-22T11:42:59","CalStd",1,"014CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,30,null,30,10,50,"ug/l",null,103,2,null,null,null,"U_238_014CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,0,42543351613016.102,853.36,42543351613016.102,930.03,740.02,890.02,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,0,2481590589483380,1906.77,2481590589483380,2050.11,1650.07,2020.12,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,0,null,9188.9500000000007,null,9322.35,9242.2999999999993,9002.2000000000007,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0,null,31921.17,null,33410.43,31406.959999999999,30946.11,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1942873.93,null,1965393.1,2003219.35,1860009.35,null,100.507890928878,null,1,null,null,null,"Sc_45_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,0,2.51937564989757E+16,3180.3,2.51937564989757E+16,3260.31,3220.32,3060.27,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,846069.39,null,853507.88,852306.32,832393.97,null,101.890152588904,null,1,null,null,null,"Rh_103_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,880.03,null,960.03,780.02,900.03,"ug/l",null,103,1,null,null,null,"Sn_118_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,100,null,90,140,70,"ug/l",null,103,1,null,null,null,"Sb_121_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,133.33000000000001,null,90,130,180,"ug/l",null,103,1,null,null,null,"Sb_123_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,0,null,56.67,null,70,30,70,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,0,1.79842004768259E+18,33.33,1.79842004768259E+18,60,30,10,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,300.01,null,200,300.01,400.01,"ug/l",null,103,2,null,null,null,"Mg_24_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0,7338114408805570,40,7338114408805570,20,70,30,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58222.51,null,58784.24,57861.35,58021.95,"ug/l",null,103,2,null,null,null,"Si_28_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,43.33,null,50,0,80,"ug/l",null,103,2,null,null,null,"P_31_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,396.67,null,370,490.01,330,"ug/l",null,103,2,null,null,null,"S_34_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18943.38,null,18559.650000000001,18809.78,19460.72,"ug/l",null,103,2,null,null,null,"K_39_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,250,null,250,210,290,"ug/l",null,103,2,null,null,null,"Ca_44_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25985.49,null,24777.11,26239.23,26940.13,null,99.033289192331495,null,2,null,null,null,"Sc_45_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,10.67,null,12,10,10,"ug/l",null,103,2,null,null,null,"Ti_49_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,370,null,400,300,410.01,"ug/l",null,103,2,null,null,null,"V_51_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1416.73,null,1390.06,1310.05,1550.07,"ug/l",null,103,2,null,null,null,{"$f":"N963&\"_\"&P963&\"_\"&G963","v":"Cr_52_015CALS.d"}]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,420,null,580.01,310,370,"ug/l",null,103,2,null,null,null,"Mn_55_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7301.56,null,8301.99,6901.41,6701.28,"ug/l",null,103,2,null,null,null,"Fe_56_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,110,null,95,165,70,"ug/l",null,103,2,null,null,null,"Co_59_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,896.69,null,930.03,840.02,920.03,"ug/l",null,103,2,null,null,null,"Ni_60_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,491.67,null,375,575.01,525.01,"ug/l",null,103,2,null,null,null,"Cu_65_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Zn_66_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,8.33,null,5,10,10,"ug/l",null,103,2,null,null,null,"Ga_71_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,45,null,40,70,25,"ug/l",null,103,2,null,null,null,"Ge_74_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"As_75_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Se_78_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,576.67999999999995,null,650.01,440.01,640.01,"ug/l",null,103,2,null,null,null,"Br_79_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,183.33,null,250,100,200,"ug/l",null,103,2,null,null,null,"Rb_85_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"Sr_88_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,10,null,10,20,0,"ug/l",null,103,2,null,null,null,"Zr_90_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,11.67,null,20,10,5,"ug/l",null,103,2,null,null,null,"Mo_98_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,271179.06,null,267075.84000000003,270234.15999999997,276227.18,null,99.925403905571997,null,2,null,null,null,"Rh_103_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,255.56,null,200,300.01,266.67,"ug/l",null,103,2,null,null,null,"Ag_109_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0.99,null,3.96,-1.64,0.64,"ug/l",null,103,2,null,null,null,"Cd_114_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-33.67,null,-45,-26,-30,"ug/l",null,103,2,null,null,null,"In_115_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,61.67,null,75,60,50,"ug/l",null,103,2,null,null,null,"Sn_118_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,3.33,null,5,0,5,"ug/l",null,103,2,null,null,null,"Sb_121_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,166.67,null,100,300.01,100,"ug/l",null,103,2,null,null,null,"I_127_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Cs_133_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,5,null,5,10,0,"ug/l",null,103,2,null,null,null,"Ba_137_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,10,null,0,30,0,"ug/l",null,103,2,null,null,null,"W_184_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,43.33,null,60,50,20,"ug/l",null,103,2,null,null,null,"Ir_193_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,0,1.21017193828113E+18,3.33,1.21017193828113E+18,0,0,10,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1150.05,null,950.03,1100.04,1400.07,"ug/l",null,103,2,null,null,null,"Au_197_015CALS.d"]
["labman","Reagent blank7","CS_2019_10_22_JLS.b","2019-10-22T11:44:47","CalStd",1,"015CALS.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,393.34,null,320,490.01,370,"ug/l",null,103,2,null,null,null,"Bi_209_015CALS.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-0.44800000000000001,null,1433.39,null,1370.05,1460.06,1470.06,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,130,null,110,100,180,"ug/l",null,103,1,null,null,null,"Sb_123_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,57717.52,null,58533.35,58984.86,55634.35,"ug/l",null,103,2,null,null,null,"Si_28_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,1623.41,null,1620.08,1530.07,1720.08,"ug/l",null,103,2,null,null,null,"Cr_52_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,43.33,null,40,30,60,"ug/l",null,103,2,null,null,null,"Ge_74_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,26.67,null,20,30,30,"ug/l",null,103,2,null,null,null,"Sr_88_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,-1.47,null,-0.82,-1.91,-1.68,"ug/l",null,103,2,null,null,null,"Cd_114_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,333.34,null,400.01,200,400.01,"ug/l",null,103,2,null,null,null,"I_127_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_016SMPL.d"]
["labman","Reagent blank8","CS_2019_10_22_JLS.b","2019-10-22T11:46:41","Sample",1,"016SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,913.36,null,1040.04,900.03,800.02,"ug/l",null,103,2,null,null,null,"Th_232_016SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,22.556999999999999,1.7,121399.39,1.7,122442.51,117139.04,124616.62,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-1.1930000000000001,null,46.67,null,70,10,60,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,4227.17,null,4250.51,4290.51,4140.49,"ug/l",null,103,2,null,null,null,"Ca_44_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,1523.4,null,1540.07,1500.06,1530.07,"ug/l",null,103,2,null,null,null,"Mn_55_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,4357.2299999999996,null,4370.55,4510.62,4190.51,"ug/l",null,103,2,null,null,null,"As_75_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1818.11,null,1807.15,1850.46,1796.73,"ug/l",null,103,2,null,null,null,"Cd_114_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,2902.92,null,2869.25,2864.25,2975.27,"ug/l",null,103,2,null,null,null,"In_115_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,14376.64,null,14907.14,14346.6,13876.18,"ug/l",null,103,2,null,null,null,"Th_232_017SMPL.d"]
["labman","QC1-1","CS_2019_10_22_JLS.b","2019-10-22T11:48:10","Sample",1,"017SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,4007.19,null,3850.48,3810.47,4360.63,"ug/l",null,103,2,null,null,null,"U_238_017SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,303.22899999999998,9.56,2763.57,9.56,2550.1999999999998,2700.22,3040.28,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,1063.3699999999999,null,1040.03,1230.04,920.03,"ug/l",null,103,2,null,null,null,"Ca_44_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25504.82,null,27881.54,23775.71,24857.22,null,97.201420676282197,null,2,null,null,null,"Sc_45_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"As_75_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,261823.01,null,255757.01,267408.90000000002,262303.12,null,96.477840235977794,null,2,null,null,null,"Rh_103_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,3367.03,null,3340.35,3580.42,3180.32,"ug/l",null,103,2,null,null,null,"Hg_202_018SMPL.d"]
["labman","QC2-1","CS_2019_10_22_JLS.b","2019-10-22T11:49:40","Sample",1,"018SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,13.33,null,40,0,0,"ug/l",null,103,2,null,null,null,"U_238_018SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,841037.12,null,834776,841928.58,846406.78,null,101.284128113573,null,1,null,null,null,"Rh_103_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,57463.49,null,58603.72,57279.62,56507.14,"ug/l",null,103,2,null,null,null,"Si_28_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25124.31,null,26259.13,25097.56,24016.25,null,95.7512591872906,null,2,null,null,null,"Sc_45_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,791.69,null,825.02,675.02,875.03,"ug/l",null,103,2,null,null,null,"Cu_65_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,66.67,null,100,40,60,"ug/l",null,103,2,null,null,null,"Sr_88_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,265577.44,null,260556.83,271891.68,264283.82,null,97.861292436416406,null,2,null,null,null,"Rh_103_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,55158.67,null,50276.12,53384.6,61815.3,"ug/l",null,103,2,null,null,null,"I_127_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_019SMPL.d"]
["labman","QC3-1","CS_2019_10_22_JLS.b","2019-10-22T11:51:09","Sample",1,"019SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,36.67,null,60,30,20,"ug/l",null,103,2,null,null,null,"Tl_205_019SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,422.48599999999999,5.83,5420.8,5.83,5680.87,5290.76,5290.76,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,14779.77,null,15236.82,15256.85,13845.64,"ug/l",null,103,1,null,null,null,{"$f":"N1283&\"_\"&P1283&\"_\"&G1283","v":"Sn_118_020SMPL.d"}]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,123.33,null,130,210,30,"ug/l",null,103,2,null,null,null,"P_31_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,13.33,null,18,14,8,"ug/l",null,103,2,null,null,null,"Ti_49_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,29658.62,null,27722.17,32229.51,29024.18,"ug/l",null,103,2,null,null,null,"Zn_66_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,20,null,10,20,30,"ug/l",null,103,2,null,null,null,"Y_89_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,11.67,null,5,30,0,"ug/l",null,103,2,null,null,null,"Pd_105_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Cs_133_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_020SMPL.d"]
["labman","acid rinse","CS_2019_10_22_JLS.b","2019-10-22T11:52:43","Sample",1,"020SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,216.67,null,140,280,230,"ug/l",null,103,2,null,null,null,"Pb_208_020SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,12.146000000000001,9.0500000000000007,76065.899999999994,9.0500000000000007,74611.570000000007,71178.2,82407.929999999993,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,390,null,350,360,460.01,"ug/l",null,103,1,null,null,null,"Sb_121_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,390,null,350,530.01,290,"ug/l",null,103,2,null,null,null,"S_34_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,10336.459999999999,null,9202.51,11403.8,10403.08,"ug/l",null,103,2,null,null,null,"Fe_56_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,28.33,null,35,30,20,"ug/l",null,103,2,null,null,null,"Ga_71_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,443.34,null,570.01,340,420.01,"ug/l",null,103,2,null,null,null,"Zr_90_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,22.22,null,0,33.33,33.33,"ug/l",null,103,2,null,null,null,"Ag_109_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,16.670000000000002,null,25,10,15,"ug/l",null,103,2,null,null,null,"Ba_137_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,55.713000000000001,88.23,13.33,88.23,20,20,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_021SMPL.d"]
["labman","base rinse","CS_2019_10_22_JLS.b","2019-10-22T11:54:13","Sample",1,"021SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,220,null,260,270,130,"ug/l",null,103,2,null,null,null,"Bi_209_021SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,36867.730000000003,null,36927.769999999997,37098.32,36577.089999999997,null,1.9072248523359401,null,1,null,null,null,"Sc_45_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,20695.39,null,19760.91,20061.490000000002,22263.77,"ug/l",null,103,2,null,null,null,"K_39_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,88.33,null,115,65,85,"ug/l",null,103,2,null,null,null,"Co_59_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,40,null,20,60,40,"ug/l",null,103,2,null,null,null,"Zr_90_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,15,null,25,5,15,"ug/l",null,103,2,null,null,null,"Mo_98_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,43.33,null,40,20,70,"ug/l",null,103,2,null,null,null,"Sb_121_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,11902.902,91.66,20,91.66,20,40,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_022SMPL.d"]
["labman","nitric","CS_2019_10_22_JLS.b","2019-10-22T11:55:41","Sample",1,"022SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1550.08,null,1750.11,1250.05,1650.09,"ug/l",null,103,2,null,null,null,"Au_197_022SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,33.33,null,100,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,-8.2000000000000003E-2,null,36.67,null,30,60,20,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,81.67,null,95,80,70,"ug/l",null,103,2,null,null,null,"Co_59_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,540.01,null,550.01,560.01,510.01,"ug/l",null,103,2,null,null,null,"Br_79_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,66.67,null,0,50,150,"ug/l",null,103,2,null,null,null,"Rb_85_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,8.33,null,15,0,10,"ug/l",null,103,2,null,null,null,"Sb_121_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_023SMPL.d"]
["labman","Reagent blank9","CS_2019_10_22_JLS.b","2019-10-22T11:57:10","Sample",1,"023SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,11304.17,null,11504.25,11454.28,10953.98,"ug/l",null,103,2,null,null,null,"Au_197_023SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,206.67,null,220,280,120,"ug/l",null,103,1,null,null,null,"Sb_121_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,-0.23200000000000001,null,30,null,30,20,40,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,350,null,340,380,330,"ug/l",null,103,2,null,null,null,"V_51_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,5,null,0,5,10,"ug/l",null,103,2,null,null,null,"Ga_71_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,100,null,100,100,100,"ug/l",null,103,2,null,null,null,"Rb_85_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,11.11,null,0,0,33.33,"ug/l",null,103,2,null,null,null,"Ag_109_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_024SMPL.d"]
["labman","Blank control1","CS_2019_10_22_JLS.b","2019-10-22T11:58:42","CalStd",1,"024SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Bi_209_024SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","B","boron",10,-0.56799999999999995,null,1243.3800000000001,null,1240.04,1320.05,1170.04,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Sb","antimony",123,null,null,5404.21,null,5500.9,5640.96,5070.76,"ug/l",null,103,1,null,null,null,"Sb_123_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","K","potassium",39,null,null,70286.53,null,69633.600000000006,68629.97,72596.02,"ug/l",null,103,2,null,null,null,"K_39_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Cr","chromium",52,null,null,2290.15,null,2320.15,2330.16,2220.14,"ug/l",null,103,2,null,null,null,"Cr_52_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Ge","germanium",74,null,null,61.67,null,45,90,50,"ug/l",null,103,2,null,null,null,"Ge_74_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Sr","strontium",88,null,null,3673.73,null,3870.44,3250.31,3900.45,"ug/l",null,103,2,null,null,null,"Sr_88_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Cd","cadmium",114,null,null,5.44,null,9.0500000000000007,2.1,5.18,"ug/l",null,103,2,null,null,null,"Cd_114_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Ce","cerium",140,null,null,407497.09,null,413568.94,410794.21,398128.12,"ug/l",null,103,2,null,null,null,"Ce_140_025SMPL.d"]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","W","tungsten",184,null,null,420.01,null,520.01,310,430.01,"ug/l",null,103,2,null,null,null,{"$f":"N1667&\"_\"&P1667&\"_\"&G1667","v":"W_184_025SMPL.d"}]
["labman","19286427-001","CS_2019_10_22_JLS.b","2019-10-22T12:00:10","Sample",1,"025SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9065763","Th","thorium",232,null,null,50,null,50,50,50,"ug/l",null,103,2,null,null,null,"Th_232_025SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","B","boron",11,-0.77900000000000003,null,5297.43,null,5430.81,5010.68,5450.81,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","Te","tellurium",125,0.03,7626.49,56.67,7626.49,40,80,50,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","Ca","calcium",44,null,null,2436.84,null,2460.17,2440.1799999999998,2410.16,"ug/l",null,103,2,null,null,null,"Ca_44_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","Mn","manganese",55,null,null,866.69,null,810.02,780.02,1010.03,"ug/l",null,103,2,null,null,null,"Mn_55_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","As","arsenic",75,null,null,16.670000000000002,null,0,0,50,"ug/l",null,103,2,null,null,null,"As_75_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","In","indium",115,null,null,-28.33,null,-18,-36,-31,"ug/l",null,103,2,null,null,null,"In_115_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","Nd","neodymium",146,null,null,40,null,50,10,60,"ug/l",null,103,2,null,null,null,"Nd_146_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","Ir","iridium",193,null,null,16.670000000000002,null,0,20,30,"ug/l",null,103,2,null,null,null,"Ir_193_026SMPL.d"]
["labman","19286427-002","CS_2019_10_22_JLS.b","2019-10-22T12:01:38","Sample",1,"026SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"green top tube","U","uranium",238,null,null,453.34,null,470.01,450.01,440.01,"ug/l",null,103,2,null,null,null,"U_238_026SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,839205.93,null,848725.46,831656.08,837236.24,null,101.063601794623,null,1,null,null,null,"Rh_103_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-344.24200000000002,null,830.03,null,1050.04,770.02,670.02,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,26780.01,null,26559.68,28522.6,25257.74,null,102.061271301506,null,2,null,null,null,"Sc_45_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,20,null,10,30,20,"ug/l",null,103,2,null,null,null,"As_75_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Se_78_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,263657.13,null,261418.77,263509.84000000003,266042.77,null,97.153684634607202,null,2,null,null,null,"Rh_103_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,40,null,50,50,20,"ug/l",null,103,2,null,null,null,"Nd_146_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,433.34,null,560.01,420.01,320,"ug/l",null,103,2,null,null,null,"Gd_156_027SMPL.d"]
["labman","19286427-003","CS_2019_10_22_JLS.b","2019-10-22T12:03:06","Sample",1,"027SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,106.67,null,50,190,80,"ug/l",null,103,2,null,null,null,"Tl_205_027SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,817419.13,null,812138.58,818281.55,821837.25,null,98.439868561082505,null,1,null,null,null,"Rh_103_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,910.03,null,970.03,870.03,890.02,"ug/l",null,103,1,null,null,null,"Sn_118_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25718.41,null,24997.39,26239.1,25918.74,null,98.015420725064203,null,2,null,null,null,"Sc_45_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,3592.05,null,3525.37,3375.33,3875.44,"ug/l",null,103,2,null,null,null,"Cu_65_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,475230.92,null,488033.21,466773.94,470885.61,"ug/l",null,103,2,null,null,null,"Zn_66_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,255764.21,null,256660.54,249648.71,260983.38,null,94.245263586500897,null,2,null,null,null,"Rh_103_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,566.69000000000005,null,500.02,400.01,800.03,"ug/l",null,103,2,null,null,null,"I_127_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,11.67,null,0,30,5,"ug/l",null,103,2,null,null,null,"Cs_133_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,136.66999999999999,null,110,130,170,"ug/l",null,103,2,null,null,null,"Tl_205_028SMPL.d"]
["labman","19286427-004","CS_2019_10_22_JLS.b","2019-10-22T12:04:34","Sample",1,"028SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,30,null,40,40,10,"ug/l",null,103,2,null,null,null,"Bi_209_028SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,840809.6,null,848432.25,816602.02,857394.52,null,101.256728012086,null,1,null,null,null,"Rh_103_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,58225.83,null,58563.48,57410.03,58703.98,"ug/l",null,103,2,null,null,null,"Si_28_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2636.87,null,2650.2,2790.23,2470.1799999999998,"ug/l",null,103,2,null,null,null,"Cr_52_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,3483.72,null,2450.1799999999998,3750.45,4250.5200000000004,"ug/l",null,103,2,null,null,null,"Cu_65_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,3527.03,null,3640.38,3400.34,3540.36,"ug/l",null,103,2,null,null,null,"Sr_88_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,261981.32,null,263566.28999999998,252741.71,269635.95,null,96.536173861028303,null,2,null,null,null,"Rh_103_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,700.06,null,900.06,1000.1,200.01,"ug/l",null,103,2,null,null,null,"I_127_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,356.67,null,350,360.01,360,"ug/l",null,103,2,null,null,null,"W_184_029SMPL.d"]
["labman","19286427-005","CS_2019_10_22_JLS.b","2019-10-22T12:06:03","Sample",1,"029SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,130,null,150,120,120,"ug/l",null,103,2,null,null,null,"Tl_205_029SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-0.79900000000000004,null,5080.71,null,5160.72,5270.76,4810.6400000000003,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-2.4910000000000001,null,33.33,null,30,50,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,90,null,60,100,110,"ug/l",null,103,2,null,null,null,"P_31_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,546.67999999999995,null,580.01,480.01,580.01,"ug/l",null,103,2,null,null,null,"Mn_55_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,806841.59,null,778571.7,826112.03,815841.03,"ug/l",null,103,2,null,null,null,"Zn_66_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,216.67,null,140,210,300,"ug/l",null,103,2,null,null,null,"Y_89_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-31.67,null,-37,-31,-27,"ug/l",null,103,2,null,null,null,"In_115_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,23.33,null,20,40,10,"ug/l",null,103,2,null,null,null,"Cs_133_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,36.67,null,90,10,10,"ug/l",null,103,2,null,null,null,"Ir_193_030SMPL.d"]
["labman","19286427-006","CS_2019_10_22_JLS.b","2019-10-22T12:07:32","Sample",1,"030SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,1933.45,null,2020.13,1790.1,1990.13,"ug/l",null,103,2,null,null,null,"Pb_208_030SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,21.131,4.84,1340226.18,4.84,1365185.92,1385975.92,1269516.7,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-281.50900000000001,null,670.02,null,670.02,720.02,620.01,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,2593.52,null,2510.1799999999998,2560.1799999999998,2710.21,"ug/l",null,103,2,null,null,null,"S_34_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,30827.01,null,29224.31,30726.45,32530.28,"ug/l",null,103,2,null,null,null,"Fe_56_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Se_78_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,953.36,null,930.03,840.02,1090.04,"ug/l",null,103,2,null,null,null,"Zr_90_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,58.33,null,60,45,70,"ug/l",null,103,2,null,null,null,"Sn_118_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,166160.82,null,164351.13,167942.13,166189.20000000001,"ug/l",null,103,2,null,null,null,"Ba_137_031SMPL.d"]
["labman","19286427-007","CS_2019_10_22_JLS.b","2019-10-22T12:09:02","Sample",1,"031SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,5.78,119.32,20,119.32,20,40,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_031SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,3.2170000000000001,0.64,101096.94,0.64,101596.14,100299.32,101395.35,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1925268.72,null,2014200.75,1807389.97,1954215.44,null,99.597145856267502,null,1,null,null,null,"Sc_45_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,12504.41,null,13405.01,12604.58,11503.64,"ug/l",null,103,2,null,null,null,"Mg_24_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,24,null,38,12,22,"ug/l",null,103,2,null,null,null,"Ti_49_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,148.33000000000001,null,140,145,160,"ug/l",null,103,2,null,null,null,"Co_59_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,22791.46,null,22454.35,22514.43,23405.61,"ug/l",null,103,2,null,null,null,"Br_79_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,40,null,70,10,40,"ug/l",null,103,2,null,null,null,"Sn_118_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,566.67999999999995,null,480.01,635.01,585.01,"ug/l",null,103,2,null,null,null,"Sb_121_032SMPL.d"]
["labman","19286427-008","CS_2019_10_22_JLS.b","2019-10-22T12:10:36","Sample",1,"032SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,63.33,null,50,50,90,"ug/l",null,103,2,null,null,null,"Tb_159_032SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,3.1859999999999999,1.55,100024.29,1.55,97604.7,100670.88,101797.3,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,1E-3,114.44,13.33,114.44,0,10,30,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,13872.12,null,11903.94,14405.91,15306.5,"ug/l",null,103,2,null,null,null,"Mg_24_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,23.33,null,16,20,34,"ug/l",null,103,2,null,null,null,"Ti_49_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,700.02,null,750.02,630.01,720.02,"ug/l",null,103,2,null,null,null,"V_51_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,25678.86,null,25358.49,26459.97,25218.12,"ug/l",null,103,2,null,null,null,"Br_79_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,13.33,null,35,0,5,"ug/l",null,103,2,null,null,null,"Pd_105_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,22.22,null,66.67,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,26.67,null,20,40,20,"ug/l",null,103,2,null,null,null,"Tb_159_033SMPL.d"]
["labman","19286427-009","CS_2019_10_22_JLS.b","2019-10-22T12:12:06","Sample",1,"033SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,2263.5,null,2110.14,2240.17,2440.19,"ug/l",null,103,2,null,null,null,"Pb_208_033SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,2E-3,137.21,23.33,137.21,60,0,10,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,6948.11,null,6771.36,6951.45,7121.51,"ug/l",null,103,1,null,null,null,"Sb_121_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,2536.84,null,2440.16,2710.2,2460.17,"ug/l",null,103,2,null,null,null,"S_34_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,746.68,null,720.02,890.02,630.01,"ug/l",null,103,2,null,null,null,"V_51_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,675.01,null,615.01,790.02,620.01,"ug/l",null,103,2,null,null,null,{"$f":"N2243&\"_\"&P2243&\"_\"&G2243","v":"Ga_71_034SMPL.d"}]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,783.35,null,870.02,770.02,710.02,"ug/l",null,103,2,null,null,null,"Zr_90_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,164334.82,null,162119.6,164826.25,166058.60999999999,"ug/l",null,103,2,null,null,null,"Ba_137_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,46.67,null,50,30,60,"ug/l",null,103,2,null,null,null,"Er_166_034SMPL.d"]
["labman","19286427-010","CS_2019_10_22_JLS.b","2019-10-22T12:13:36","Sample",1,"034SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,30,null,60,20,10,"ug/l",null,103,2,null,null,null,"Bi_209_034SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,2086336.74,null,1936390.75,2309437,2013182.47,null,107.929445090174,null,1,null,null,null,"Sc_45_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,5337.52,null,5480.9,5400.88,5130.78,"ug/l",null,103,1,null,null,null,"Sb_123_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,74118.42,null,71139.56,77164.3,74051.41,"ug/l",null,103,2,null,null,null,"K_39_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2800.23,null,2760.22,2810.23,2830.23,"ug/l",null,103,2,null,null,null,"Cr_52_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,35,null,35,55,15,"ug/l",null,103,2,null,null,null,"Ge_74_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,15505.3,null,15737.16,14741.29,16037.45,"ug/l",null,103,2,null,null,null,"Mo_98_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,6.59,null,2.23,2.5,15.05,"ug/l",null,103,2,null,null,null,"Cd_114_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,404992.14,null,404907.06,409575.38,400493.98,"ug/l",null,103,2,null,null,null,"Ce_140_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,3483.74,null,2850.26,3700.46,3900.51,"ug/l",null,103,2,null,null,null,"Au_197_035SMPL.d"]
["labman","19286427-011","CS_2019_10_22_JLS.b","2019-10-22T12:15:14","Sample",1,"035SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,33.33,null,60,20,20,"ug/l",null,103,2,null,null,null,"Th_232_035SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,2.5999999999999999E-2,125.64,3467.01,125.64,3080.27,3840.42,3480.34,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-1.841,null,40,null,60,30,30,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,2136.8000000000002,null,2120.13,2320.15,1970.11,"ug/l",null,103,2,null,null,null,"Ca_44_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,630.01,null,610.01,600.01,680.01,"ug/l",null,103,2,null,null,null,"Ni_60_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,30,null,60,30,0,"ug/l",null,103,2,null,null,null,"As_75_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-21.67,null,-26,-24,-15,"ug/l",null,103,2,null,null,null,"In_115_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,23.33,null,50,20,0,"ug/l",null,103,2,null,null,null,"Nd_146_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,313.33,null,300,280,360,"ug/l",null,103,2,null,null,null,"Hg_202_036SMPL.d"]
["labman","19286427-012","CS_2019_10_22_JLS.b","2019-10-22T12:16:47","Sample",1,"036SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,413.34,null,380.01,590.01,270,"ug/l",null,103,2,null,null,null,"U_238_036SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,821420.8,null,830650.77,823834.99,809776.63,null,98.921780295190402,null,1,null,null,null,"Rh_103_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,56814.8,null,57590.44,55524.15,57329.81,"ug/l",null,103,2,null,null,null,"Si_28_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25731.75,null,26279.13,25357.93,25558.2,null,98.066273496699097,null,2,null,null,null,"Sc_45_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,3983.8,null,3475.35,4350.5600000000004,4125.49,"ug/l",null,103,2,null,null,null,"Cu_65_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,258908.34,null,257419.98,263072.98,256232.05,null,95.403826961921396,null,2,null,null,null,"Rh_103_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,666.69,null,700.03,400.01,900.04,"ug/l",null,103,2,null,null,null,"I_127_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,606.67999999999995,null,560.01,660.01,600.01,"ug/l",null,103,2,null,null,null,"Hg_202_037SMPL.d"]
["labman","19286427-013","CS_2019_10_22_JLS.b","2019-10-22T12:18:17","Sample",1,"037SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,120,null,100,180,80,"ug/l",null,103,2,null,null,null,"Tl_205_037SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-0.79200000000000004,null,5207.42,null,4680.6000000000004,5440.83,5500.82,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,55199.83,null,54571.22,55795.06,55233.21,"ug/l",null,103,2,null,null,null,"Si_28_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,56.67,null,110,10,50,"ug/l",null,103,2,null,null,null,"P_31_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,3400.34,null,3025.26,3325.32,3850.45,"ug/l",null,103,2,null,null,null,"Cu_65_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,3667.06,null,3510.36,3950.45,3540.37,"ug/l",null,103,2,null,null,null,"Sr_88_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,206.67,null,230,190,200,"ug/l",null,103,2,null,null,null,"Y_89_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,533.36,null,800.05,300.01,500.02,"ug/l",null,103,2,null,null,null,"I_127_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,396.67,null,460.01,400.01,330,"ug/l",null,103,2,null,null,null,"W_184_038SMPL.d"]
["labman","19286427-014","CS_2019_10_22_JLS.b","2019-10-22T12:19:45","Sample",1,"038SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,96.67,null,160,40,90,"ug/l",null,103,2,null,null,null,"Tl_205_038SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-0.86199999999999999,null,4797.3,null,4950.67,4670.59,4770.63,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-0.59699999999999998,null,50,null,50,20,80,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,63.33,null,20,100,70,"ug/l",null,103,2,null,null,null,"P_31_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,613.35,null,610.01,480.01,750.02,"ug/l",null,103,2,null,null,null,"Mn_55_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,26.67,null,0,60,20,"ug/l",null,103,2,null,null,null,"As_75_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,213.33,null,240,230,170,"ug/l",null,103,2,null,null,null,"Y_89_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-25.67,null,-39,-2,-36,"ug/l",null,103,2,null,null,null,"In_115_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,18.329999999999998,null,10,15,30,"ug/l",null,103,2,null,null,null,"Cs_133_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,43.33,null,20,50,60,"ug/l",null,103,2,null,null,null,"Ir_193_039SMPL.d"]
["labman","19286427-015","CS_2019_10_22_JLS.b","2019-10-22T12:21:14","Sample",1,"039SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,403.34,null,440.01,350,420.01,"ug/l",null,103,2,null,null,null,"U_238_039SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,22.962,6.1,1472541.28,6.1,1438710.13,1598306.23,1380607.48,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-222.58799999999999,null,543.34,null,460.01,550.01,620.01,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,{"$f":"N2627&\"_\"&P2627&\"_\"&G2627","v":"Hg_202_040SMPL.d"}]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25765.15,null,25898.68,24777.14,26619.62,null,98.193538814270397,null,2,null,null,null,"Sc_45_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,33064.43,null,33932.36,32730.15,32530.78,"ug/l",null,103,2,null,null,null,"Fe_56_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,6.67,null,10,5,5,"ug/l",null,103,2,null,null,null,"Se_78_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,933.36,null,810.02,920.03,1070.04,"ug/l",null,103,2,null,null,null,"Zr_90_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,50,null,45,60,45,"ug/l",null,103,2,null,null,null,"Sn_118_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,456.68,null,460.01,420.01,490.01,"ug/l",null,103,2,null,null,null,"Gd_156_040SMPL.d"]
["labman","19286427-016","CS_2019_10_22_JLS.b","2019-10-22T12:22:47","Sample",1,"040SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,2.2719999999999998,144.84,10,144.84,10,20,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_040SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,3.1629999999999998,1.91,101348.18,1.91,102289.78,100148.42,101606.33,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1972877.11,null,2301026.06,1859197.47,1758407.79,null,102.0600017587,null,1,null,null,null,"Sc_45_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,14205.92,null,15606.68,14205.99,12805.1,"ug/l",null,103,2,null,null,null,"Mg_24_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,30,null,38,22,30,"ug/l",null,103,2,null,null,null,"Ti_49_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,113.33,null,95,185,60,"ug/l",null,103,2,null,null,null,"Co_59_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,33862.720000000001,null,34363.800000000003,33251.5,33972.870000000003,"ug/l",null,103,2,null,null,null,"Br_79_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,8.33,null,10,10,5,"ug/l",null,103,2,null,null,null,"Pd_105_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,595.01,null,655.01,555.01,575.01,"ug/l",null,103,2,null,null,null,"Sb_121_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,43.33,null,50,50,30,"ug/l",null,103,2,null,null,null,"Tb_159_041SMPL.d"]
["labman","19286427-017","CS_2019_10_22_JLS.b","2019-10-22T12:24:17","Sample",1,"041SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,2883.62,null,3050.32,2500.1999999999998,3100.34,"ug/l",null,103,2,null,null,null,"Au_197_041SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-0.89,null,4740.6099999999997,null,4610.58,4810.63,4800.63,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,590.01,null,580.01,590.01,600.01,"ug/l",null,103,1,null,null,null,"Sn_118_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,13572.19,null,12104.56,14406.03,14205.97,"ug/l",null,103,2,null,null,null,"Mg_24_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,37.33,null,38,28,46,"ug/l",null,103,2,null,null,null,"Ti_49_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,36367.69,null,35906.71,37239.5,35956.85,"ug/l",null,103,2,null,null,null,"Br_79_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Pd_105_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,456.67,null,530.01,490.01,350,"ug/l",null,103,2,null,null,null,"Gd_156_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,43.33,null,0,70,60,"ug/l",null,103,2,null,null,null,"Tb_159_042SMPL.d"]
["labman","19286427-018","CS_2019_10_22_JLS.b","2019-10-22T12:25:45","Sample",1,"042SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,2543.5500000000002,null,2750.25,2440.1999999999998,2440.19,"ug/l",null,103,2,null,null,null,"Pb_208_042SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,686.68,null,670.01,690.01,700.02,"ug/l",null,103,1,null,null,null,"Sn_118_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,7108.17,null,6731.36,7411.62,7181.53,"ug/l",null,103,1,null,null,null,"Sb_121_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,31.33,null,32,28,34,"ug/l",null,103,2,null,null,null,"Ti_49_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,1049910.54,null,1022834.1,1078816.82,1048080.69,"ug/l",null,103,2,null,null,null,"Zn_66_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,810.02,null,755.02,910.02,765.02,"ug/l",null,103,2,null,null,null,"Ga_71_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,21.67,null,10,50,5,"ug/l",null,103,2,null,null,null,"Pd_105_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,8.33,null,0,0,25,"ug/l",null,103,2,null,null,null,"Cs_133_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,162409.22,null,161897.60999999999,162074.47,163255.57999999999,"ug/l",null,103,2,null,null,null,"Ba_137_043SMPL.d"]
["labman","19286427-019","CS_2019_10_22_JLS.b","2019-10-22T12:27:15","Sample",1,"043SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,2383.52,null,2090.14,2700.24,2360.1799999999998,"ug/l",null,103,2,null,null,null,"Pb_208_043SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,22.765999999999998,0.84,1453580.24,0.84,1440904.82,1458562.17,1461273.73,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,6754.69,null,6811.37,6771.37,6681.33,"ug/l",null,103,1,null,null,null,"Sb_121_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,2530.1799999999998,null,2360.16,2730.21,2500.17,"ug/l",null,103,2,null,null,null,"S_34_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,30259.38,null,31428.46,29224.33,30125.360000000001,"ug/l",null,103,2,null,null,null,"Fe_56_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,806.69,null,690.02,935.03,795.02,"ug/l",null,103,2,null,null,null,"Ga_71_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,1120.04,null,1120.04,1100.04,1140.04,"ug/l",null,103,2,null,null,null,"Zr_90_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,33.33,null,0,33.33,66.67,"ug/l",null,103,2,null,null,null,"Ag_109_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,161722.96,null,161665.94,161882.71,161620.23000000001,"ug/l",null,103,2,null,null,null,"Ba_137_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,6.7439999999999998,158.13999999999999,23.33,158.13999999999999,0,10,60,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_044SMPL.d"]
["labman","19286427-020","CS_2019_10_22_JLS.b","2019-10-22T12:28:50","Sample",1,"044SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,20,null,40,10,10,"ug/l",null,103,2,null,null,null,{"$f":"N2947&\"_\"&P2947&\"_\"&G2947","v":"Bi_209_044SMPL.d"}]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,207875.01,null,208487.2,213698.24,201439.59,null,10.753697640110801,null,1,null,null,null,"Sc_45_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,100,null,0,100,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,21613.26,null,22964.76,21162.78,20712.240000000002,"ug/l",null,103,2,null,null,null,"K_39_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,111.67,null,80,150,105,"ug/l",null,103,2,null,null,null,"Co_59_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,116.67,null,85,115,150,"ug/l",null,103,2,null,null,null,"Ge_74_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,376.67,null,415.01,370,345,"ug/l",null,103,2,null,null,null,"Mo_98_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,58.33,null,85,25,65,"ug/l",null,103,2,null,null,null,"Sb_121_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,2540.1999999999998,null,2500.19,2770.23,2350.17,"ug/l",null,103,2,null,null,null,"Ce_140_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,5884.5,null,6151.28,5851.09,5651.12,"ug/l",null,103,2,null,null,null,"Au_197_045SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:30:20","Sample",1,"045SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,23.33,null,0,40,30,"ug/l",null,103,2,null,null,null,"Th_232_045SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,4.6340000000000003,13.56,2943.58,13.56,2750.22,2980.26,3100.27,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,50.033000000000001,44.78,60,44.78,40,60,80,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,170,null,200,150,160,"ug/l",null,103,2,null,null,null,"Ca_44_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,363.34,null,470.01,330,290,"ug/l",null,103,2,null,null,null,"Ni_60_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,150,null,200.01,100,150,"ug/l",null,103,2,null,null,null,"Rb_85_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_046SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T12:31:50","Sample",1,"046SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,26.67,null,20,40,20,"ug/l",null,103,2,null,null,null,"Hg_202_046SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,413.15100000000001,7.45,516.67999999999995,7.45,480.01,570.01,500.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,983.36,null,840.02,1100.04,1010.03,null,0.11842414022763099,null,1,null,null,null,"Rh_103_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,49797.75,null,51532.46,51161.33,46699.47,"ug/l",null,103,2,null,null,null,"Si_28_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2050.12,null,1840.1,2120.13,2190.14,"ug/l",null,103,2,null,null,null,"Cr_52_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,516.67999999999995,null,700.02,450.01,400.01,"ug/l",null,103,2,null,null,null,"Cu_65_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Sr_88_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,300.01,null,400.01,200,300.01,"ug/l",null,103,2,null,null,null,"I_127_047SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T12:33:19","Sample",1,"047SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"W_184_047SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.4139999999999999,null,380,null,410.01,380,350,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.524,null,1643.41,null,1590.07,1800.09,1540.07,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,49530.28,null,50599.8,49456.68,48534.35,"ug/l",null,103,2,null,null,null,"Si_28_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2193.4699999999998,null,1960.11,2340.16,2280.15,"ug/l",null,103,2,null,null,null,"Cr_52_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,466.68,null,440.01,470.01,490.01,"ug/l",null,103,2,null,null,null,"Mn_55_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,16.670000000000002,null,50,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0.65,null,-0.55000000000000004,3.18,-0.68,"ug/l",null,103,2,null,null,null,"Cd_114_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-15,null,-12,-18,-15,"ug/l",null,103,2,null,null,null,"In_115_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_048SMPL.d"]
["labman","Blank control2","CS_2019_10_22_JLS.b","2019-10-22T12:34:53","Sample",1,"048SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,10,null,10,0,20,"ug/l",null,103,2,null,null,null,"Th_232_048SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.5640000000000001,null,1540.07,null,1690.08,1270.05,1660.08,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-1.0189999999999999,null,50,null,40,70,40,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,233.33,null,240,250,210,"ug/l",null,103,2,null,null,null,"Ca_44_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,396.67,null,450.01,400,340,"ug/l",null,103,2,null,null,null,"Mn_55_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"As_75_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-22,null,-33,-12,-21,"ug/l",null,103,2,null,null,null,"In_115_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,83.33,null,100,90,60,"ug/l",null,103,2,null,null,null,"Ir_193_049SMPL.d"]
["labman","Reagent blank10","CS_2019_10_22_JLS.b","2019-10-22T12:36:23","Sample",1,"049SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"U_238_049SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,851746.11,null,868604.13,837941.55,848692.64,null,102.573786265062,null,1,null,null,null,"Rh_103_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-1.4390000000000001,null,36.67,null,30,40,40,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25391.27,null,25097.48,26339.26,24737.07,null,96.768657618946904,null,2,null,null,null,"Sc_45_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,114536.37,null,118392.95,108834.82,116381.33,"ug/l",null,103,2,null,null,null,"Fe_56_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,608.34,null,590.01,655.01,580.01,"ug/l",null,103,2,null,null,null,"Se_78_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,270965.48,null,271675.33,267845.90999999997,273375.19,null,99.846701660454599,null,2,null,null,null,"Rh_103_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,36.67,null,30,40,40,"ug/l",null,103,2,null,null,null,{"$f":"N3331&\"_\"&P3331&\"_\"&G3331","v":"Sn_118_050SMPL.d"}]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_050SMPL.d"]
["labman","QC1-2","CS_2019_10_22_JLS.b","2019-10-22T12:37:52","Sample",1,"050SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,11287.37,null,10403.450000000001,11684.29,11774.37,"ug/l",null,103,2,null,null,null,"Tl_205_050SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-1.6E-2,null,350,null,400,300,350,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,53039.47,null,51287.33,52781.99,55049.1,"ug/l",null,103,1,null,null,null,"Sn_118_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,200.01,null,100,300.01,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,554.01,null,570.01,616.01,476.01,"ug/l",null,103,2,null,null,null,"Ti_49_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,600.02,null,500.01,500.01,800.04,"ug/l",null,103,2,null,null,null,"Zn_66_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,983.36,null,870.02,920.03,1160.04,"ug/l",null,103,2,null,null,null,"Br_79_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,3105.28,null,3250.31,2980.26,3085.28,"ug/l",null,103,2,null,null,null,"Pd_105_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,2920.25,null,2845.24,2810.23,3105.29,"ug/l",null,103,2,null,null,null,"Sb_121_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Tb_159_051SMPL.d"]
["labman","QC2-2","CS_2019_10_22_JLS.b","2019-10-22T12:39:20","Sample",1,"051SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,200,null,130,260,210,"ug/l",null,103,2,null,null,null,"Pb_208_051SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,173.21,3.33,173.21,10,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,183.33,null,270,110,170,"ug/l",null,103,1,null,null,null,"Sb_121_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,406.67,null,450.01,410,360,"ug/l",null,103,2,null,null,null,"S_34_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,403.34,null,460.01,420.01,330,"ug/l",null,103,2,null,null,null,"V_51_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Ga_71_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,15,null,5,10,30,"ug/l",null,103,2,null,null,null,"Pd_105_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,255.56,null,266.67,333.34,166.67,"ug/l",null,103,2,null,null,null,"Ag_109_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,15,null,10,10,25,"ug/l",null,103,2,null,null,null,"Ba_137_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,90,null,110,70,90,"ug/l",null,103,2,null,null,null,"Pb_208_052SMPL.d"]
["labman","QC3-2","CS_2019_10_22_JLS.b","2019-10-22T12:40:53","Sample",1,"052SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,26.67,null,0,20,60,"ug/l",null,103,2,null,null,null,"Bi_209_052SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1935276.06,null,2096708.09,1766454.35,1942665.75,null,100.114841294442,null,1,null,null,null,"Sc_45_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,390.01,null,420.01,280,470.01,"ug/l",null,103,2,null,null,null,"S_34_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19060.22,null,18459.68,19660.71,19060.259999999998,"ug/l",null,103,2,null,null,null,"K_39_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,3373.66,null,3455.34,3420.33,3245.3,"ug/l",null,103,2,null,null,null,"Ga_71_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,20,null,20,30,10,"ug/l",null,103,2,null,null,null,"Zr_90_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,90,null,85,80,105,"ug/l",null,103,2,null,null,null,"Mo_98_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,22082.87,null,22044.5,22379.919999999998,21824.2,"ug/l",null,103,2,null,null,null,"Ba_137_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,-1.034,null,0,null,0,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_053SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T12:42:21","Sample",1,"053SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,67707.600000000006,null,68065.88,67001.13,68055.8,"ug/l",null,103,2,null,null,null,"Bi_209_053SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1853012.52,null,1948985.44,1794672.94,1815379.19,null,95.859220400111298,null,1,null,null,null,"Sc_45_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,166.67,null,100,200.01,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18743.240000000002,null,19410.66,17808.87,19010.2,"ug/l",null,103,2,null,null,null,"K_39_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,70,null,55,85,70,"ug/l",null,103,2,null,null,null,"Co_59_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,843.36,null,970.03,810.02,750.02,"ug/l",null,103,2,null,null,null,{"$f":"N3587&\"_\"&P3587&\"_\"&G3587","v":"Br_79_054SMPL.d"}]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,39.82,null,35,19.45,65,"ug/l",null,103,2,null,null,null,"Mo_98_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Sb_121_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_054SMPL.d"]
["labman","Reagent blank11","CS_2019_10_22_JLS.b","2019-10-22T12:43:48","Sample",1,"054SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,2550.2399999999998,null,2300.19,3050.34,2300.1999999999998,"ug/l",null,103,2,null,null,null,"Au_197_054SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,1E-3,173.21,6.67,173.21,20,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-8.2000000000000003E-2,null,8471.94,null,8581.99,8652.0300000000007,8181.8,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19760.91,null,19560.63,21663,18059.11,"ug/l",null,103,2,null,null,null,"K_39_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,105,null,85,130,100,"ug/l",null,103,2,null,null,null,"Co_59_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,896.69,null,780.02,1070.03,840.02,"ug/l",null,103,2,null,null,null,"Br_79_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,51.67,null,50,45,60,"ug/l",null,103,2,null,null,null,"Mo_98_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,10,null,0,15,15,"ug/l",null,103,2,null,null,null,"Sb_121_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_055SMPL.d"]
["labman","Blank control3","CS_2019_10_22_JLS.b","2019-10-22T12:45:17","CalStd",1,"055SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,2150.17,null,2150.19,1950.13,2350.1999999999998,"ug/l",null,103,2,null,null,null,"Au_197_055SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Be","beryllium",9,1E-3,86.63,6.67,86.63,0,10,10,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Ge","germanium",74,-1.2999999999999999E-2,null,2970.26,null,3010.26,2890.24,3010.27,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Al","aluminium",27,-0.83399999999999996,null,10,null,20,10,0,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","V","vanadium",51,null,null,426.67,null,560.01,320,400.01,"ug/l",null,103,2,null,null,null,"V_51_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Ni","nickel",60,null,null,413.34,null,450.01,350,440.01,"ug/l",null,103,2,null,null,null,"Ni_60_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Rb","rubidium",85,null,null,133.34,null,0,300.01,100,"ug/l",null,103,2,null,null,null,"Rb_85_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Ag","silver",109,null,null,588.91,null,466.68,833.36,466.68,"ug/l",null,103,2,null,null,null,"Ag_109_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_056SMPL.d"]
["labman","19286426-001","CS_2019_10_22_JLS.b","2019-10-22T12:46:51","Sample",1,"056SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 9-254-001","Hg","mercury",202,null,null,130,null,90,100,200,"ug/l",null,103,2,null,null,null,"Hg_202_056SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","B","boron",10,-0.93400000000000005,null,886.69,null,870.02,830.02,960.03,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Sb","antimony",123,null,null,90,null,70,140,60,"ug/l",null,103,1,null,null,null,"Sb_123_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Si","silicon",28,null,null,46676.12,null,47441.36,46579.27,46007.72,"ug/l",null,103,2,null,null,null,"Si_28_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Cr","chromium",52,null,null,2276.8200000000002,null,2650.2,2060.12,2120.13,"ug/l",null,103,2,null,null,null,"Cr_52_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Rb","rubidium",85,null,null,100,null,100,100,100,"ug/l",null,103,2,null,null,null,"Rb_85_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Sr","strontium",88,null,null,26.67,null,50,0,30,"ug/l",null,103,2,null,null,null,"Sr_88_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Cd","cadmium",114,null,null,-0.24,null,-0.14000000000000001,-1.77,1.18,"ug/l",null,103,2,null,null,null,"Cd_114_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","W","tungsten",184,null,null,33.33,null,0,80,20,"ug/l",null,103,2,null,null,null,"W_184_057SMPL.d"]
["labman","19286426-002","CS_2019_10_22_JLS.b","2019-10-22T12:48:21","Sample",1,"057SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash vial","Th","thorium",232,null,null,593.35,null,640.02,640.01,500.01,"ug/l",null,103,2,null,null,null,"Th_232_057SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,200,null,190,170,240,"ug/l",null,103,1,null,null,null,"Sb_123_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-4.101,null,20,null,40,10,10,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2020.12,null,2040.12,1990.12,2030.12,"ug/l",null,103,2,null,null,null,"Cr_52_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,43.33,null,15,40,75,"ug/l",null,103,2,null,null,null,"Ge_74_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,-0.5,null,4.09,-1.64,-3.95,"ug/l",null,103,2,null,null,null,"Cd_114_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,26.67,null,30,30,20,"ug/l",null,103,2,null,null,null,"Ce_140_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_058SMPL.d"]
["labman","19286426-003","CS_2019_10_22_JLS.b","2019-10-22T12:49:51","Sample",1,"058SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,426.67,null,320,470.01,490.01,"ug/l",null,103,2,null,null,null,"Th_232_058SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-0.05,null,2650.2,null,2760.22,2690.21,2500.1799999999998,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-1.956,null,40,null,20,40,60,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,233.33,null,250,230,220,"ug/l",null,103,2,null,null,null,{"$f":"N3907&\"_\"&P3907&\"_\"&G3907","v":"Ca_44_059SMPL.d"}]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,360,null,360,370.01,350,"ug/l",null,103,2,null,null,null,"Ni_60_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-30,null,-18,-36,-36,"ug/l",null,103,2,null,null,null,"In_115_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,106.67,null,40,150,130,"ug/l",null,103,2,null,null,null,"Hg_202_059SMPL.d"]
["labman","19286426-004","CS_2019_10_22_JLS.b","2019-10-22T12:51:21","Sample",1,"059SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_059SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,826707.02,null,835526.86,834577.8,810016.39,null,99.558387373504004,null,1,null,null,null,"Rh_103_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,45262.559999999998,null,45536.54,46168.13,44083,"ug/l",null,103,2,null,null,null,"Si_28_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,26626.37,null,26239.119999999999,26219.06,27420.94,null,101.475759374289,null,2,null,null,null,"Sc_45_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,641.67999999999995,null,600.02,550.01,775.02,"ug/l",null,103,2,null,null,null,"Cu_65_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Se_78_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,259589.23,null,257855.78,253704.7,267207.21999999997,null,95.654727139960997,null,2,null,null,null,"Rh_103_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,433.35,null,300.01,400.01,600.03,"ug/l",null,103,2,null,null,null,"I_127_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_060SMPL.d"]
["labman","19286426-005","CS_2019_10_22_JLS.b","2019-10-22T12:52:55","Sample",1,"060SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"Tl_205_060SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-1.7999999999999999E-2,null,276.67,null,280,280,270,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,376.67,null,350,330,450.01,"ug/l",null,103,1,null,null,null,"Sn_118_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,40,null,30,30,60,"ug/l",null,103,2,null,null,null,"P_31_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,8.67,null,8,12,6,"ug/l",null,103,2,null,null,null,"Ti_49_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,200.01,null,400.02,0,200,"ug/l",null,103,2,null,null,null,"Zn_66_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Pd_105_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,13.33,null,0,30,10,"ug/l",null,103,2,null,null,null,"Cs_133_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_061SMPL.d"]
["labman","19286426-006","CS_2019_10_22_JLS.b","2019-10-22T12:54:24","Sample",1,"061SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,116.67,null,110,110,130,"ug/l",null,103,2,null,null,null,"Pb_208_061SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,2.7080000000000002,2.4,205162.55,2.4,204414.07999999999,204152.71,206920.86,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,416.67,null,490.01,400,360,"ug/l",null,103,1,null,null,null,"Sb_121_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,343.34,null,310,440.01,280,"ug/l",null,103,2,null,null,null,"S_34_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,16274.56,null,16108.23,14305.79,18409.66,"ug/l",null,103,2,null,null,null,"Fe_56_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,13.33,null,5,10,25,"ug/l",null,103,2,null,null,null,"Ga_71_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,53.33,null,40,30,90,"ug/l",null,103,2,null,null,null,"Zr_90_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,13.33,null,25,15,0,"ug/l",null,103,2,null,null,null,"Cs_133_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,45,null,40,30,65,"ug/l",null,103,2,null,null,null,"Ba_137_062SMPL.d"]
["labman","19286426-007","CS_2019_10_22_JLS.b","2019-10-22T12:55:52","Sample",1,"062SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,-1.034,null,0,null,0,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_062SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,1.139,3.67,103062.33,3.67,105180.96,103078.89,100927.15,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1866243.15,null,1862978.57,1935118.88,1800632,null,96.543661299296502,null,1,null,null,null,"Sc_45_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,523.34,null,560.01,420.01,590.01,"ug/l",null,103,2,null,null,null,"S_34_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,9069.32,null,8002,8902.4599999999991,10303.51,"ug/l",null,103,2,null,null,null,"Fe_56_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,68.33,null,65,55,85,"ug/l",null,103,2,null,null,null,"Co_59_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Zr_90_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,38.33,null,30,60,25,"ug/l",null,103,2,null,null,null,"Sn_118_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,5,null,5,0,10,"ug/l",null,103,2,null,null,null,"Sb_121_063SMPL.d"]
["labman","19286426-008","CS_2019_10_22_JLS.b","2019-10-22T12:57:21","Sample",1,"063SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,8.4000000000000005E-2,2298.8000000000002,3.33,2298.8000000000002,0,0,10,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_063SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-2.1000000000000001E-2,null,183.33,null,200,220,130,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1930515.86,null,2064652.31,1887644.35,1839250.91,null,99.868588398541604,null,1,null,null,null,"Sc_45_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,100,null,100,0,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,13.33,null,10,26,4,"ug/l",null,103,2,null,null,null,"Ti_49_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,78.33,null,115,60,60,"ug/l",null,103,2,null,null,null,"Co_59_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,620.01,null,760.02,570.01,530.01,"ug/l",null,103,2,null,null,null,"Br_79_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,5,null,15,0,0,"ug/l",null,103,2,null,null,null,"Pd_105_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,5,null,0,10,5,"ug/l",null,103,2,null,null,null,"Sb_121_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"Tb_159_064SMPL.d"]
["labman","19286426-009","CS_2019_10_22_JLS.b","2019-10-22T12:58:55","Sample",1,"064SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1016.71,null,1150.05,950.04,950.04,"ug/l",null,103,2,null,null,null,"Au_197_064SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,{"$f":"N4291&\"_\"&P4291&\"_\"&G4291","v":"Be_9_065SMPL.d"}]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,330,null,340,290,360,"ug/l",null,103,1,null,null,null,"Sb_121_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,15.135999999999999,4.78,550.01,4.78,560.01,520.01,570.01,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,383.34,null,360,330,460.01,"ug/l",null,103,2,null,null,null,"V_51_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,506.68,null,540.01,530.01,450.01,"ug/l",null,103,2,null,null,null,"Ni_60_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,83.33,null,0,50,200,"ug/l",null,103,2,null,null,null,"Rb_85_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,300.01,null,266.67,233.34,400.01,"ug/l",null,103,2,null,null,null,"Ag_109_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_065SMPL.d"]
["labman","19286426-010","CS_2019_10_22_JLS.b","2019-10-22T13:00:23","Sample",1,"065SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Bi_209_065SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.2989999999999999,null,503.34,null,580.01,410.01,520.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,303.33,null,280,330,300,"ug/l",null,103,1,null,null,null,"Sb_123_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,44076.4,null,43712.21,43200.98,45316.01,"ug/l",null,103,2,null,null,null,"Si_28_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2363.5,null,2130.14,2600.19,2360.16,"ug/l",null,103,2,null,null,null,"Cr_52_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,26.67,null,25,30,25,"ug/l",null,103,2,null,null,null,"Ge_74_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,30,null,30,10,50,"ug/l",null,103,2,null,null,null,"Sr_88_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,8.4700000000000006,null,4.91,9.32,11.18,"ug/l",null,103,2,null,null,null,"Cd_114_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,266.67,null,300.01,400.01,100,"ug/l",null,103,2,null,null,null,"I_127_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"W_184_066SMPL.d"]
["labman","19286426-011","CS_2019_10_22_JLS.b","2019-10-22T13:01:53","Sample",1,"066SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,140,null,110,120,190,"ug/l",null,103,2,null,null,null,"Th_232_066SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.3939999999999999,null,2273.48,null,2210.14,2310.15,2300.15,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-2.1850000000000001,null,36.67,null,20,50,40,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,256.67,null,260,270,240,"ug/l",null,103,2,null,null,null,"Ca_44_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,466.67,null,390,500.01,510.01,"ug/l",null,103,2,null,null,null,"Mn_55_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,4.12,null,7.05,-0.41,5.73,"ug/l",null,103,2,null,null,null,"Cd_114_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-12,null,-21,-9,-6,"ug/l",null,103,2,null,null,null,"In_115_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,96.67,null,90,60,140,"ug/l",null,103,2,null,null,null,"Th_232_067SMPL.d"]
["labman","19286426-012","CS_2019_10_22_JLS.b","2019-10-22T13:03:23","Sample",1,"067SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_067SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-4.8220000000000001,null,13.33,null,20,0,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,17625.330000000002,null,18359.419999999998,17358.32,17158.259999999998,"ug/l",null,103,2,null,null,null,"K_39_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,203.33,null,180,180,250,"ug/l",null,103,2,null,null,null,"Ca_44_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,26.67,null,20,30,30,"ug/l",null,103,2,null,null,null,"Ge_74_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,5,null,5,5,5,"ug/l",null,103,2,null,null,null,"Mo_98_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,10,null,20,0,10,"ug/l",null,103,2,null,null,null,"Ce_140_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1150.0899999999999,null,950.04,1550.18,950.04,"ug/l",null,103,2,null,null,null,"Au_197_068SMPL.d"]
["labman","19286426-013","CS_2019_10_22_JLS.b","2019-10-22T13:04:56","Sample",1,"068SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,96.67,null,30,130,130,"ug/l",null,103,2,null,null,null,"Th_232_068SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-3.3000000000000002E-2,null,2753.55,null,2520.19,2960.25,2780.22,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.622,91.49,60,91.49,80,40,60,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,243.33,null,310,200,220,"ug/l",null,103,2,null,null,null,"Ca_44_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,650.01,null,560.01,700.01,690.01,"ug/l",null,103,2,null,null,null,"Ni_60_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,50,null,0,150,0,"ug/l",null,103,2,null,null,null,"Rb_85_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,1.67,null,0,5,0,"ug/l",null,103,2,null,null,null,"Ru_101_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_069SMPL.d"]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,{"$f":"N4611&\"_\"&P4611&\"_\"&G4611","v":"Nd_146_069SMPL.d"}]
["labman","19286426-014","CS_2019_10_22_JLS.b","2019-10-22T13:06:26","Sample",1,"069SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,33.33,null,20,40,40,"ug/l",null,103,2,null,null,null,"Hg_202_069SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.399,null,406.67,null,340,380,500.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,858126.29,null,859761.32,860771.86,853845.69,null,103.342136782244,null,1,null,null,null,"Rh_103_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,42282.14,null,42539.47,42318.879999999997,41988.08,"ug/l",null,103,2,null,null,null,"Si_28_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,26546.26,null,25518.09,26679.71,27440.98,null,101.17043948583699,null,2,null,null,null,"Sc_45_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,650.02,null,650.02,550.01,750.02,"ug/l",null,103,2,null,null,null,"Cu_65_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,10,null,0,30,0,"ug/l",null,103,2,null,null,null,"Sr_88_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,268952.18,null,268273.21999999997,273062.71000000002,265520.59999999998,null,99.104830899183099,null,2,null,null,null,"Rh_103_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,233.34,null,200.01,300.01,200.01,"ug/l",null,103,2,null,null,null,"I_127_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,10,null,10,20,0,"ug/l",null,103,2,null,null,null,"W_184_070SMPL.d"]
["labman","19286426-015","CS_2019_10_22_JLS.b","2019-10-22T13:07:55","Sample",1,"070SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tl_205_070SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.508,null,1810.09,null,1950.11,1850.1,1630.07,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,410.01,null,430.01,450.01,350,"ug/l",null,103,1,null,null,null,"Sn_118_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,23.33,null,10,40,20,"ug/l",null,103,2,null,null,null,"P_31_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,393.34,null,430.01,400,350,"ug/l",null,103,2,null,null,null,"Mn_55_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,400.01,null,600.02,200,400.01,"ug/l",null,103,2,null,null,null,"Zn_66_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-16,null,-9,-18,-21,"ug/l",null,103,2,null,null,null,"In_115_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Cs_133_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,36.67,null,10,50,50,"ug/l",null,103,2,null,null,null,"Ir_193_071SMPL.d"]
["labman","19286426-016","CS_2019_10_22_JLS.b","2019-10-22T13:09:24","Sample",1,"071SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,53.33,null,50,20,90,"ug/l",null,103,2,null,null,null,"Pb_208_071SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,4.343,0.77,307088.24,0.77,299414.43,311353.39,310496.90999999997,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-13.061,null,63.33,null,60,80,50,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,406.67,null,390,390,440.01,"ug/l",null,103,2,null,null,null,"S_34_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,11804.05,null,12604.58,10303.09,12504.47,"ug/l",null,103,2,null,null,null,"Fe_56_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,20,null,0,0,60,"ug/l",null,103,2,null,null,null,"Zr_90_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,35,null,80,10,15,"ug/l",null,103,2,null,null,null,"Sn_118_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,33.33,null,30,50,20,"ug/l",null,103,2,null,null,null,"Ir_193_072SMPL.d"]
["labman","19286426-017","CS_2019_10_22_JLS.b","2019-10-22T13:10:59","Sample",1,"072SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,1.1259999999999999,332.34,6.67,332.34,0,20,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_072SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-2.1000000000000001E-2,null,193.33,null,240,80,260,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-19.619,null,80,null,80,100,60,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,666.71,null,300.01,1000.09,700.02,"ug/l",null,103,2,null,null,null,"Mg_24_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,10636.85,null,11503.92,10503.74,9902.8799999999992,"ug/l",null,103,2,null,null,null,"Fe_56_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,686.68,null,720.02,580.01,760.02,"ug/l",null,103,2,null,null,null,"Br_79_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,85,null,195,30,30,"ug/l",null,103,2,null,null,null,"Sn_118_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_073SMPL.d"]
["labman","19286426-018","CS_2019_10_22_JLS.b","2019-10-22T13:12:28","Sample",1,"073SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_073SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-2.1000000000000001E-2,null,186.67,null,250,170,140,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,466.67,null,510.01,380,510.01,"ug/l",null,103,1,null,null,null,"Sn_118_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,10,null,6,20,4,"ug/l",null,103,2,null,null,null,"Ti_49_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,66.67,null,100,0,100,"ug/l",null,103,2,null,null,null,"Zn_66_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,656.68,null,560.01,570.01,840.02,"ug/l",null,103,2,null,null,null,"Br_79_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,3.33,null,5,5,0,"ug/l",null,103,2,null,null,null,"Pd_105_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,8.33,null,15,10,0,"ug/l",null,103,2,null,null,null,"Sb_121_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_074SMPL.d"]
["labman","19286426-019","CS_2019_10_22_JLS.b","2019-10-22T13:13:57","Sample",1,"074SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,73.33,null,90,80,50,"ug/l",null,103,2,null,null,null,"Pb_208_074SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,260,null,140,250,390,"ug/l",null,103,1,null,null,null,"Sb_121_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,446.67,null,240,600.01,500.01,"ug/l",null,103,2,null,null,null,"S_34_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,543.34,null,620.01,520.01,490.01,"ug/l",null,103,2,null,null,null,"V_51_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,13.33,null,20,20,0,"ug/l",null,103,2,null,null,null,"Ga_71_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,200,null,300,100,200,"ug/l",null,103,2,null,null,null,{"$f":"N4995&\"_\"&P4995&\"_\"&G4995","v":"Rb_85_075SMPL.d"}]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,111.11,null,66.67,200,66.67,"ug/l",null,103,2,null,null,null,"Ag_109_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,55,null,70,40,55,"ug/l",null,103,2,null,null,null,"Ba_137_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_075SMPL.d"]
["labman","19286426-020","CS_2019_10_22_JLS.b","2019-10-22T13:15:26","Sample",1,"075SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,26.67,null,40,40,0,"ug/l",null,103,2,null,null,null,"Bi_209_075SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,106.38500000000001,22.97,533.34,22.97,560.01,640.01,400.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,700.02,null,770.02,730.02,600.01,"ug/l",null,103,1,null,null,null,"Sb_123_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19210.400000000001,null,19510.55,18259.57,19861.09,"ug/l",null,103,2,null,null,null,"K_39_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2713.54,null,2650.2,2750.22,2740.21,"ug/l",null,103,2,null,null,null,"Cr_52_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,110,null,125,85,120,"ug/l",null,103,2,null,null,null,"Ge_74_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,25,null,20,20,35,"ug/l",null,103,2,null,null,null,"Mo_98_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,-6.04,null,-6.98,-0.21,-10.93,"ug/l",null,103,2,null,null,null,"Cd_114_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,353.34,null,360.01,300,400.01,"ug/l",null,103,2,null,null,null,"Ce_140_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"W_184_076SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:17:00","Sample",1,"076SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,73.33,null,70,120,30,"ug/l",null,103,2,null,null,null,"Th_232_076SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,0.79700000000000004,20.23,2556.86,20.23,2300.15,2550.1999999999998,2820.23,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,13.542,30.2,36.67,30.2,30,40,40,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,253.33,null,180,290,290,"ug/l",null,103,2,null,null,null,"Ca_44_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,450.01,null,350,440.01,560.01,"ug/l",null,103,2,null,null,null,"Ni_60_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,10,null,0,10,20,"ug/l",null,103,2,null,null,null,"Ce_140_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_077SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T13:18:31","Sample",1,"077SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,70,null,60,40,110,"ug/l",null,103,2,null,null,null,"Hg_202_077SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,101.789,28.33,1466.73,28.33,1420.06,1350.06,1630.08,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,1083.3699999999999,null,1240.05,1130.04,880.02,null,0.13046770857676401,null,1,null,null,null,"Rh_103_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,170,null,190,190,130,"ug/l",null,103,2,null,null,null,"Ca_44_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,426.67,null,380,350,550.01,"ug/l",null,103,2,null,null,null,"Ni_60_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,391.67,null,325,350,500.01,"ug/l",null,103,2,null,null,null,"Cu_65_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,633.35,null,900.04,600.01,400.01,"ug/l",null,103,2,null,null,null,"I_127_078SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T13:20:00","Sample",1,"078SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,16.670000000000002,null,20,20,10,"ug/l",null,103,2,null,null,null,"Hg_202_078SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.4950000000000001,null,293.33,null,210,260,410,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,836945.17,null,845887.8,842528.11,822419.6,null,100.791344170773,null,1,null,null,null,"Rh_103_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,40668.519999999997,null,40925.699999999997,40815.58,40264.269999999997,"ug/l",null,103,2,null,null,null,"Si_28_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2823.56,null,2620.1999999999998,2660.2,3190.29,"ug/l",null,103,2,null,null,null,{"$f":"N5251&\"_\"&P5251&\"_\"&G5251","v":"Cr_52_079SMPL.d"}]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,1100.04,null,1200.04,1000.03,1100.04,"ug/l",null,103,2,null,null,null,"Cu_65_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,20,null,20,10,30,"ug/l",null,103,2,null,null,null,"Sr_88_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.97,null,5.05,2.36,-1.5,"ug/l",null,103,2,null,null,null,"Cd_114_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,700.03,null,1200.06,600.01,300.01,"ug/l",null,103,2,null,null,null,"I_127_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"W_184_079SMPL.d"]
["labman","Blank control4","CS_2019_10_22_JLS.b","2019-10-22T13:21:28","Sample",1,"079SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tl_205_079SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.575,null,1370.05,null,1600.07,1090.03,1420.06,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-3.24,null,26.67,null,50,10,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,13.33,null,20,10,10,"ug/l",null,103,2,null,null,null,"P_31_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,406.67,null,340,360,520.01,"ug/l",null,103,2,null,null,null,"Mn_55_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,1200.07,null,600.02,1300.08,1700.11,"ug/l",null,103,2,null,null,null,"Zn_66_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-23,null,-27,-25,-17,"ug/l",null,103,2,null,null,null,"In_115_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Cs_133_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,43.33,null,40,40,50,"ug/l",null,103,2,null,null,null,"Ir_193_080SMPL.d"]
["labman","Reagent blank12","CS_2019_10_22_JLS.b","2019-10-22T13:23:01","Sample",1,"080SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_080SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,10.724,0.48,718707.46,0.48,703866.32,724654.05,727602.02,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-4.4779999999999998,null,16.670000000000002,null,10,20,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,2480.17,null,2250.14,2540.1799999999998,2650.2,"ug/l",null,103,2,null,null,null,"P_31_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,1436.73,null,1240.05,1580.07,1490.07,"ug/l",null,103,2,null,null,null,"Mn_55_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,4697.3,null,4660.62,4770.66,4660.62,"ug/l",null,103,2,null,null,null,"As_75_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,3059.61,null,2963.26,3119.3,3096.28,"ug/l",null,103,2,null,null,null,"In_115_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,4255.54,null,4145.51,4225.54,4395.58,"ug/l",null,103,2,null,null,null,"Cs_133_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,33.33,null,10,50,40,"ug/l",null,103,2,null,null,null,"Ir_193_081SMPL.d"]
["labman","QC1-3","CS_2019_10_22_JLS.b","2019-10-22T13:24:30","Sample",1,"081SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,3927.17,null,3990.52,4120.5600000000004,3670.43,"ug/l",null,103,2,null,null,null,"U_238_081SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.82499999999999996,1.95,86523.91,1.95,86634.3,85780.43,87156.99,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-2388.8519999999999,null,5734.38,null,5761.07,6011.14,5430.93,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25398.01,null,24997.39,25297.95,25898.68,null,96.794331726235001,null,2,null,null,null,"Sc_45_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,8201.9699999999993,null,6501.19,9502.6,8602.1299999999992,"ug/l",null,103,2,null,null,null,"Fe_56_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-2754.72,null,-2699.38,-2845.41,-2719.38,"ug/l",null,103,2,null,null,null,"In_115_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,4677.32,null,4615.63,4800.6899999999996,4615.63,"ug/l",null,103,2,null,null,null,"Sn_118_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_082SMPL.d"]
["labman","QC2-3","CS_2019_10_22_JLS.b","2019-10-22T13:26:00","Sample",1,"082SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"U_238_082SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-1.7999999999999999E-2,null,276.67,null,290,310,230,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,530.01,null,640.01,400.01,550.01,"ug/l",null,103,1,null,null,null,"Sn_118_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25618.27,null,27020.25,25337.86,24496.69,null,97.633764512667099,null,2,null,null,null,"Sc_45_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,9.33,null,4,10,14,"ug/l",null,103,2,null,null,null,"Ti_49_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,0,5,5,"ug/l",null,103,2,null,null,null,"Se_78_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,263419.59000000003,null,257560.78,263550.45,269147.53999999998,null,97.066155946518094,null,2,null,null,null,"Rh_103_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,36.67,null,40,10,60,"ug/l",null,103,2,null,null,null,"Pd_105_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_083SMPL.d"]
["labman","QC3-3","CS_2019_10_22_JLS.b","2019-10-22T13:27:30","Sample",1,"083SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,13.33,null,0,10,30,"ug/l",null,103,2,null,null,null,"Tl_205_083SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,49.865000000000002,2.14,1630605.18,2.14,1616925.91,1591748.26,1683141.38,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,583.34,null,520.01,610.01,620.01,"ug/l",null,103,1,null,null,null,{"$f":"N5571&\"_\"&P5571&\"_\"&G5571","v":"Sn_118_084SMPL.d"}]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,13481.66,null,13625.1,13905.3,12914.57,"ug/l",null,103,2,null,null,null,"P_31_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,14,null,12,22,8,"ug/l",null,103,2,null,null,null,"Ti_49_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,1166.74,null,600.02,1000.04,1900.17,"ug/l",null,103,2,null,null,null,"Zn_66_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Y_89_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,25,null,20,35,20,"ug/l",null,103,2,null,null,null,"Pd_105_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,730.02,null,730.02,735.02,725.02,"ug/l",null,103,2,null,null,null,"Cs_133_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_084SMPL.d"]
["labman","NMS 59","CS_2019_10_22_JLS.b","2019-10-22T13:29:03","Sample",1,"084SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,474349.94,null,472654.13,466689.21,483706.48,"ug/l",null,103,2,null,null,null,"Pb_208_084SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.21299999999999999,8.64,45533.23,8.64,45426.35,46478.89,44694.46,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,156.66999999999999,null,150,120,200,"ug/l",null,103,1,null,null,null,"Sb_121_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,373.34,null,400.01,360,360,"ug/l",null,103,2,null,null,null,"S_34_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,496.68,null,620.01,440.01,430.01,"ug/l",null,103,2,null,null,null,"V_51_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,5,null,0,0,15,"ug/l",null,103,2,null,null,null,"Ga_71_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,13.33,null,10,0,30,"ug/l",null,103,2,null,null,null,"Zr_90_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,2789.14,null,3133.64,2966.94,2266.83,"ug/l",null,103,2,null,null,null,"Ag_109_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,16.670000000000002,null,10,25,15,"ug/l",null,103,2,null,null,null,"Ba_137_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,2E-3,117668.35,3.33,117668.35,0,0,10,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_085SMPL.d"]
["labman","Reagent blank13","CS_2019_10_22_JLS.b","2019-10-22T13:30:31","Sample",1,"085SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,1253.3900000000001,null,1190.05,1350.06,1220.05,"ug/l",null,103,2,null,null,null,"Bi_209_085SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,2021817.31,null,1921534.03,2153197.62,1990720.28,null,104.59175460909,null,1,null,null,null,"Sc_45_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,126.67,null,200,60,120,"ug/l",null,103,1,null,null,null,"Sb_123_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,17441.88,null,19060.060000000001,17408.59,15856.99,"ug/l",null,103,2,null,null,null,"K_39_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,116.67,null,155,90,105,"ug/l",null,103,2,null,null,null,"Co_59_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,28.33,null,35,40,10,"ug/l",null,103,2,null,null,null,"Ge_74_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,63.33,null,45,85,60,"ug/l",null,103,2,null,null,null,"Mo_98_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,11.67,null,0,5,30,"ug/l",null,103,2,null,null,null,"Sb_121_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,3066.99,null,2850.29,3400.39,2950.29,"ug/l",null,103,2,null,null,null,"Au_197_086SMPL.d"]
["labman","Blank control5","CS_2019_10_22_JLS.b","2019-10-22T13:32:00","CalStd",1,"086SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,1650.09,null,1630.09,1440.07,1880.12,"ug/l",null,103,2,null,null,null,"Th_232_086SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Ge","germanium",74,-4.9000000000000002E-2,null,2570.1999999999998,null,2270.15,2830.23,2610.21,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Al","aluminium",27,-0.13800000000000001,null,33.33,null,10,50,40,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Ca","calcium",44,null,null,230,null,220,310,160,"ug/l",null,103,2,null,null,null,"Ca_44_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Ni","nickel",60,null,null,483.34,null,530.01,450.01,470.01,"ug/l",null,103,2,null,null,null,"Ni_60_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Mo","molybdenum",98,null,null,21.67,null,30,20,15,"ug/l",null,103,2,null,null,null,"Mo_98_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Au","gold",197,null,null,2083.4899999999998,null,2200.17,2000.14,2050.15,"ug/l",null,103,2,null,null,null,"Au_197_087SMPL.d"]
["labman","19286424-001","CS_2019_10_22_JLS.b","2019-10-22T13:33:29","Sample",1,"087SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"lot 7-016-001","Hg","mercury",202,null,null,110,null,70,120,140,"ug/l",null,103,2,null,null,null,"Hg_202_087SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","B","boron",10,-0.48899999999999999,null,1340.05,null,1360.05,1090.03,1570.07,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Al","aluminium",27,-0.72899999999999998,null,13.33,null,20,0,20,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Si","silicon",28,null,null,38290.03,null,38320.03,37528.449999999997,39021.599999999999,"ug/l",null,103,2,null,null,null,"Si_28_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Ni","nickel",60,null,null,400.01,null,410.01,240,550.01,"ug/l",null,103,2,null,null,null,"Ni_60_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Rb","rubidium",85,null,null,283.33999999999997,null,300.01,100,450.01,"ug/l",null,103,2,null,null,null,"Rb_85_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Sr","strontium",88,null,null,193.33,null,180,220,180,"ug/l",null,103,2,null,null,null,"Sr_88_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_088SMPL.d"]
["labman","19286424-002","CS_2019_10_22_JLS.b","2019-10-22T13:35:04","Sample",1,"088SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"acid wash bottle","W","tungsten",184,null,null,13.33,null,10,20,10,"ug/l",null,103,2,null,null,null,"W_184_088SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-0.77,null,1046.7,null,1040.03,1060.03,1040.03,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,203.33,null,210,180,220,"ug/l",null,103,1,null,null,null,"Sb_123_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,38066.35,null,38199.919999999998,37809.25,38189.89,"ug/l",null,103,2,null,null,null,"Si_28_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2550.19,null,2280.15,2530.19,2840.23,"ug/l",null,103,2,null,null,null,"Cr_52_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,13.33,null,15,15,10,"ug/l",null,103,2,null,null,null,"Ge_74_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,13.33,null,0,40,0,"ug/l",null,103,2,null,null,null,"Sr_88_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.29,null,2.77,3.05,-1.95,"ug/l",null,103,2,null,null,null,"Cd_114_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,666.7,null,200.01,900.04,900.05,"ug/l",null,103,2,null,null,null,"I_127_089SMPL.d"]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,13.33,null,40,0,0,"ug/l",null,103,2,null,null,null,{"$f":"N5955&\"_\"&P5955&\"_\"&G5955","v":"W_184_089SMPL.d"}]
["labman","19286424-003","CS_2019_10_22_JLS.b","2019-10-22T13:36:33","Sample",1,"089SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,433.34,null,530.01,340,430.01,"ug/l",null,103,2,null,null,null,"Th_232_089SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-0.85799999999999998,null,4637.25,null,4660.59,4560.5600000000004,4690.6000000000004,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-4.3310000000000004,null,16.670000000000002,null,10,20,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,266.67,null,250,290,260,"ug/l",null,103,2,null,null,null,"Ca_44_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,426.67,null,520.01,500.01,260,"ug/l",null,103,2,null,null,null,"Mn_55_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,13.33,null,20,0,20,"ug/l",null,103,2,null,null,null,"As_75_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-69.34,null,-163.01,-18,-27,"ug/l",null,103,2,null,null,null,"In_115_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,40,null,20,30,70,"ug/l",null,103,2,null,null,null,"Ir_193_090SMPL.d"]
["labman","19286424-004","CS_2019_10_22_JLS.b","2019-10-22T13:38:03","Sample",1,"090SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"U_238_090SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.38200000000000001,3.63,53882.46,3.63,53648.480000000003,53497.93,54500.97,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-37.890999999999998,null,116.67,null,130,80,140,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25057.54,null,25037.39,24777.31,25357.91,null,95.496766663143504,null,2,null,null,null,"Sc_45_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,9936.34,null,12304.63,9202.39,8302,"ug/l",null,103,2,null,null,null,"Fe_56_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,6.67,null,5,15,0,"ug/l",null,103,2,null,null,null,"Se_78_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,250587.16,null,245598.51,253248.75,252914.21,null,92.337597318412193,null,2,null,null,null,"Rh_103_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,45,null,60,50,25,"ug/l",null,103,2,null,null,null,"Sn_118_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_091SMPL.d"]
["labman","19286424-005","CS_2019_10_22_JLS.b","2019-10-22T13:39:32","Sample",1,"091SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,9.4E-2,2082.12,3.33,2082.12,0,10,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_091SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-1.7999999999999999E-2,null,276.67,null,260,270,300,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,373.33,null,370,350,400,"ug/l",null,103,1,null,null,null,"Sn_118_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,200,null,100,300.01,200,"ug/l",null,103,2,null,null,null,"Mg_24_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,10,null,18,0,12,"ug/l",null,103,2,null,null,null,"Ti_49_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,104275.24,null,97568.11,99982.89,115274.72,"ug/l",null,103,2,null,null,null,"Zn_66_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,490.01,null,560.01,400,510.01,"ug/l",null,103,2,null,null,null,"Br_79_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,10,null,0,30,0,"ug/l",null,103,2,null,null,null,"Pd_105_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_092SMPL.d"]
["labman","19286424-006","CS_2019_10_22_JLS.b","2019-10-22T13:41:05","Sample",1,"092SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,116.67,null,110,140,100,"ug/l",null,103,2,null,null,null,"Pb_208_092SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,410.01,null,380.01,450.01,400.01,"ug/l",null,103,1,null,null,null,"Sn_118_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,326.67,null,210,410.01,360,"ug/l",null,103,1,null,null,null,"Sb_121_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,12,null,6,16,14,"ug/l",null,103,2,null,null,null,"Ti_49_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,90967.96,null,93748.79,86512.53,92642.559999999998,"ug/l",null,103,2,null,null,null,"Zn_66_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,15,null,20,15,10,"ug/l",null,103,2,null,null,null,"Ga_71_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,6.67,null,0,5,15,"ug/l",null,103,2,null,null,null,"Pd_105_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,5,null,0,10,5,"ug/l",null,103,2,null,null,null,"Cs_133_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_093SMPL.d"]
["labman","19286424-007","CS_2019_10_22_JLS.b","2019-10-22T13:42:34","Sample",1,"093SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,116.67,null,70,150,130,"ug/l",null,103,2,null,null,null,"Pb_208_093SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.218,61.19,44096.19,61.19,52722.31,40093.78,39472.47,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,313.33,null,280,320,340,"ug/l",null,103,1,null,null,null,"Sb_121_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,423.34,null,520.01,340,410.01,"ug/l",null,103,2,null,null,null,"S_34_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,380.01,null,340,390.01,410.01,"ug/l",null,103,2,null,null,null,"Mn_55_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,86746.79,null,85506.98,87617.75,87115.63,"ug/l",null,103,2,null,null,null,"Zn_66_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,5,null,10,5,0,"ug/l",null,103,2,null,null,null,{"$f":"N6275&\"_\"&P6275&\"_\"&G6275","v":"Pd_105_094SMPL.d"}]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,3.33,null,5,5,0,"ug/l",null,103,2,null,null,null,"Cs_133_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,20,null,20,30,10,"ug/l",null,103,2,null,null,null,"Ir_193_094SMPL.d"]
["labman","19286424-008","CS_2019_10_22_JLS.b","2019-10-22T13:44:03","Sample",1,"094SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,100,null,60,100,140,"ug/l",null,103,2,null,null,null,"Pb_208_094SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.14000000000000001,20.67,39305.49,20.67,38580.589999999997,38791.03,40544.839999999997,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-9.6419999999999995,null,53.33,null,30,70,60,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,376.67,null,360,460.01,310,"ug/l",null,103,2,null,null,null,"S_34_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,8068.64,null,7201.64,8602.14,8402.1299999999992,"ug/l",null,103,2,null,null,null,"Fe_56_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,11.67,null,25,5,5,"ug/l",null,103,2,null,null,null,"Ga_71_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,50,null,60,90,0,"ug/l",null,103,2,null,null,null,"Zr_90_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,23.33,null,5,20,45,"ug/l",null,103,2,null,null,null,"Sn_118_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Ba_137_095SMPL.d"]
["labman","19286424-009","CS_2019_10_22_JLS.b","2019-10-22T13:45:33","Sample",1,"095SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,-1.034,null,0,null,0,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_095SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-1.9E-2,null,240,null,190,250,280,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1815065.65,null,1897696.69,1767729.97,1779770.29,null,93.896169611975495,null,1,null,null,null,"Sc_45_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18142.8,null,19110.560000000001,18109.419999999998,17208.419999999998,"ug/l",null,103,2,null,null,null,"K_39_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,76.67,null,70,75,85,"ug/l",null,103,2,null,null,null,"Co_59_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,560.01,null,460.01,700.02,520.01,"ug/l",null,103,2,null,null,null,"Br_79_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,14.63,null,20,15,8.9,"ug/l",null,103,2,null,null,null,"Mo_98_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,16.670000000000002,null,10,5,35,"ug/l",null,103,2,null,null,null,"Sb_121_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Ce_140_096SMPL.d"]
["labman","19286424-010","CS_2019_10_22_JLS.b","2019-10-22T13:47:06","Sample",1,"096SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,1266.73,null,1650.11,950.03,1200.05,"ug/l",null,103,2,null,null,null,"Au_197_096SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-4.1000000000000002E-2,null,2556.86,null,2400.17,2490.1799999999998,2780.22,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.106,496.85,40,496.85,30,30,60,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,416.68,null,390.01,430.01,430.01,"ug/l",null,103,2,null,null,null,"V_51_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,410,null,390,330,510.01,"ug/l",null,103,2,null,null,null,"Ni_60_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,150,null,150,100,200,"ug/l",null,103,2,null,null,null,"Rb_85_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,20,null,20,20,20,"ug/l",null,103,2,null,null,null,"Sb_121_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_097SMPL.d"]
["labman","19286424-011","CS_2019_10_22_JLS.b","2019-10-22T13:48:36","Sample",1,"097SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_097SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-0.95,null,830.02,null,800.02,810.02,880.02,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,223.33,null,280,250,140,"ug/l",null,103,1,null,null,null,"Sb_123_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,353.33,null,350,310,400,"ug/l",null,103,2,null,null,null,"V_51_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2736.88,null,3040.27,2380.16,2790.22,"ug/l",null,103,2,null,null,null,"Cr_52_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,50,null,0,50,100,"ug/l",null,103,2,null,null,null,"Rb_85_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,133.34,null,233.34,166.67,0,"ug/l",null,103,2,null,null,null,"Ag_109_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,0.47,null,-1.0900000000000001,1.32,1.18,"ug/l",null,103,2,null,null,null,"Cd_114_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_098SMPL.d"]
["labman","19286424-012","CS_2019_10_22_JLS.b","2019-10-22T13:50:05","Sample",1,"098SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,33.33,null,20,50,30,"ug/l",null,103,2,null,null,null,"Bi_209_098SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-0.93100000000000005,null,866.69,null,690.01,870.02,1040.03,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,220,null,140,250,270,"ug/l",null,103,1,null,null,null,"Sb_123_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18092.59,null,18509.66,17709,18059.099999999999,"ug/l",null,103,2,null,null,null,"K_39_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2730.21,null,2840.23,2570.19,2780.22,"ug/l",null,103,2,null,null,null,"Cr_52_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,15,null,15,5,25,"ug/l",null,103,2,null,null,null,"Ge_74_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,38.33,null,30,40,45,"ug/l",null,103,2,null,null,null,"Mo_98_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,2.4700000000000002,null,3.32,-1.0900000000000001,5.18,"ug/l",null,103,2,null,null,null,"Cd_114_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"W_184_099SMPL.d"]
["labman","19286424-013","CS_2019_10_22_JLS.b","2019-10-22T13:51:47","Sample",1,"099SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,80,null,40,110,90,"ug/l",null,103,2,null,null,null,"Th_232_099SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-6.4000000000000001E-2,null,2343.4899999999998,null,2120.13,2500.1799999999998,2410.17,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-4.391,null,16.670000000000002,null,10,20,20,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,216.67,null,140,190,320,"ug/l",null,103,2,null,null,null,"Ca_44_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,383.34,null,410.01,320,420.01,"ug/l",null,103,2,null,null,null,{"$f":"N6659&\"_\"&P6659&\"_\"&G6659","v":"Mn_55_100SMPL.d"}]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,0,0,20,"ug/l",null,103,2,null,null,null,"As_75_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-20,null,-12,-30,-18,"ug/l",null,103,2,null,null,null,"In_115_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,23.33,null,20,20,30,"ug/l",null,103,2,null,null,null,"Hg_202_100SMPL.d"]
["labman","19286424-014","CS_2019_10_22_JLS.b","2019-10-22T13:53:21","Sample",1,"100SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_100SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,839649.05,null,840475.61,843295.53,835176,null,101.116965697098,null,1,null,null,null,"Rh_103_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-7.335,null,50,null,40,100,10,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,25044.13,null,25077.5,25698.33,24356.560000000001,null,95.445672521870605,null,2,null,null,null,"Sc_45_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,641.67999999999995,null,600.01,825.02,500.01,"ug/l",null,103,2,null,null,null,"Cu_65_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Se_78_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,255468.28,null,252406.89,261286.87,252711.09,null,94.136219061849602,null,2,null,null,null,"Rh_103_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,333.35,null,300.01,0,700.03,"ug/l",null,103,2,null,null,null,"I_127_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_101SMPL.d"]
["labman","19286424-015","CS_2019_10_22_JLS.b","2019-10-22T13:54:51","Sample",1,"101SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Tl_205_101SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-0.02,null,196.67,null,180,120,290,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,330,null,440.01,300,250,"ug/l",null,103,1,null,null,null,"Sn_118_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,16.670000000000002,null,0,0,50,"ug/l",null,103,2,null,null,null,"P_31_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,9.33,null,10,14,4,"ug/l",null,103,2,null,null,null,"Ti_49_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,98808.99,null,102697.73,98774.38,94954.85,"ug/l",null,103,2,null,null,null,"Zn_66_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,257795.39,null,255257.01,257754.37,260374.78,null,94.993721630496793,null,2,null,null,null,"Rh_103_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,13.33,null,10,10,20,"ug/l",null,103,2,null,null,null,"Pd_105_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,13.33,null,25,0,15,"ug/l",null,103,2,null,null,null,"Cs_133_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"Tl_205_102SMPL.d"]
["labman","19286424-016","CS_2019_10_22_JLS.b","2019-10-22T13:56:21","Sample",1,"102SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,133.33000000000001,null,100,180,120,"ug/l",null,103,2,null,null,null,"Pb_208_102SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.13,21.47,39976.99,21.47,38129.660000000003,41446.85,40354.449999999997,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,50,null,100,50,0,"ug/l",null,103,2,null,null,null,"P_31_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,386.67,null,440.01,460.01,260,"ug/l",null,103,2,null,null,null,"S_34_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,100287.14,null,101391.87,105416.58,94052.97,"ug/l",null,103,2,null,null,null,"Zn_66_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,13.33,null,0,20,20,"ug/l",null,103,2,null,null,null,"Zr_90_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"Cs_133_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,16.670000000000002,null,0,10,40,"ug/l",null,103,2,null,null,null,"Ir_193_103SMPL.d"]
["labman","19286424-017","CS_2019_10_22_JLS.b","2019-10-22T13:57:50","Sample",1,"103SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,1.0720000000000001,340.38,6.67,340.38,0,20,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_103SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.17599999999999999,3.66,43411.37,3.66,42850.03,43431.37,43952.7,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-2.7890000000000001,null,40,null,10,70,40,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,{"$f":"N6915&\"_\"&P6915&\"_\"&G6915","v":"Hg_202_104SMPL.d"}]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,366.67,null,310,480.01,310,"ug/l",null,103,2,null,null,null,"S_34_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,9602.7199999999993,null,11103.52,8602.24,9102.41,"ug/l",null,103,2,null,null,null,"Fe_56_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Se_78_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,26.67,null,10,20,50,"ug/l",null,103,2,null,null,null,"Zr_90_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,25,null,30,25,20,"ug/l",null,103,2,null,null,null,"Sn_118_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Ba_137_104SMPL.d"]
["labman","19286424-018","CS_2019_10_22_JLS.b","2019-10-22T13:59:23","Sample",1,"104SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,-1.034,null,0,null,0,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_104SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-2.1000000000000001E-2,null,196.67,null,190,260,140,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1806208.67,null,1723510.13,1826370.75,1868745.13,null,93.4379842585532,null,1,null,null,null,"Sc_45_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,133.34,null,200.01,0,200.01,"ug/l",null,103,2,null,null,null,"Mg_24_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,12.67,null,12,16,10,"ug/l",null,103,2,null,null,null,"Ti_49_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,96.67,null,85,80,125,"ug/l",null,103,2,null,null,null,"Co_59_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,520.01,null,440.01,550.01,570.01,"ug/l",null,103,2,null,null,null,"Br_79_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,23.15,null,24.45,20,25,"ug/l",null,103,2,null,null,null,"Mo_98_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,25,null,40,10,25,"ug/l",null,103,2,null,null,null,"Sb_121_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_105SMPL.d"]
["labman","19286424-019","CS_2019_10_22_JLS.b","2019-10-22T14:00:52","Sample",1,"105SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,900.03,null,1050.04,600.01,1050.04,"ug/l",null,103,2,null,null,null,"Au_197_105SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-6.4000000000000001E-2,null,2433.5,null,2470.17,2510.1799999999998,2320.16,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,{"$f":"N7043&\"_\"&P7043&\"_\"&G7043","v":"Ge_74_106SMPL.d"}]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.42899999999999999,457.4,53.33,457.4,130,20,10,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,363.33,null,330,360,400,"ug/l",null,103,2,null,null,null,"V_51_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,716.68,null,800.02,640.01,710.01,"ug/l",null,103,2,null,null,null,"Ni_60_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,50,null,100,50,0,"ug/l",null,103,2,null,null,null,"Rb_85_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,77.78,null,33.33,100,100,"ug/l",null,103,2,null,null,null,"Ag_109_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_106SMPL.d"]
["labman","19286424-020","CS_2019_10_22_JLS.b","2019-10-22T14:02:22","Sample",1,"106SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,26.67,null,20,40,20,"ug/l",null,103,2,null,null,null,"Hg_202_106SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,123.14,14.58,433.34,14.58,340,490.01,470.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,670.01,null,640.01,670.01,700.02,"ug/l",null,103,1,null,null,null,"Sb_123_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,356.67,null,260,370,440.01,"ug/l",null,103,2,null,null,null,"S_34_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,296.67,null,210,280,400.01,"ug/l",null,103,2,null,null,null,"V_51_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,3.33,null,5,5,0,"ug/l",null,103,2,null,null,null,"Ga_71_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,166.67,null,250,50,200,"ug/l",null,103,2,null,null,null,"Rb_85_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,622.23,null,666.68,466.67,733.35,"ug/l",null,103,2,null,null,null,"Ag_109_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Tb_159_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_107SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:03:53","Sample",1,"107SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,36.67,null,10,10,90,"ug/l",null,103,2,null,null,null,"Bi_209_107SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,303.33999999999997,null,260,240,410.01,"ug/l",null,103,1,null,null,null,"Sb_121_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,270,null,240,340,230,"ug/l",null,103,1,null,null,null,"Sb_123_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,103.33,null,70,140,100,"ug/l",null,103,2,null,null,null,"V_51_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,11.67,null,10,0,25,"ug/l",null,103,2,null,null,null,"Ga_71_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,11.67,null,0,15,20,"ug/l",null,103,2,null,null,null,"Ge_74_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,144.44999999999999,null,66.67,366.68,0,"ug/l",null,103,2,null,null,null,"Ag_109_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,11.67,null,15,10,10,"ug/l",null,103,2,null,null,null,"Ba_137_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,23.33,null,20,50,0,"ug/l",null,103,2,null,null,null,"Ce_140_108SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:05:29","Sample",1,"108SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,133.33000000000001,null,110,120,170,"ug/l",null,103,2,null,null,null,{"$f":"N7235&\"_\"&P7235&\"_\"&G7235","v":"Bi_209_108SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,29207.18,null,30425.72,30556.16,26639.67,null,1.5109330288712901,null,1,null,null,null,"Sc_45_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,400,null,540.01,350,310,"ug/l",null,103,1,null,null,null,"Sb_123_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,17992.47,null,19360.46,15856.98,18759.96,"ug/l",null,103,2,null,null,null,"K_39_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,91.67,null,115,80,80,"ug/l",null,103,2,null,null,null,"Co_59_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,15,null,15,5,25,"ug/l",null,103,2,null,null,null,"Ge_74_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,13.33,null,10,10,20,"ug/l",null,103,2,null,null,null,"Mo_98_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.99,null,1.73,0.64,3.59,"ug/l",null,103,2,null,null,null,"Cd_114_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"Ce_140_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,950.04,null,750.03,1000.04,1100.06,"ug/l",null,103,2,null,null,null,"Au_197_109SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:06:57","Sample",1,"109SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Th_232_109SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-7.4999999999999997E-2,null,2323.4899999999998,null,1950.11,2500.1799999999998,2520.1799999999998,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,-0.33900000000000002,null,26.67,null,30,10,40,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,200,null,170,260,170,"ug/l",null,103,2,null,null,null,"Ca_44_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,616.67999999999995,null,660.01,530.01,660.01,"ug/l",null,103,2,null,null,null,"Ni_60_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_110SMPL.d"]
["labman","Blank control6","CS_2019_10_22_JLS.b","2019-10-22T14:08:29","Sample",1,"110SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,10,null,10,10,10,"ug/l",null,103,2,null,null,null,"Hg_202_110SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.4910000000000001,null,283.33999999999997,null,400.01,180,270,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,797861.08,null,780991.94,805712.64,806878.66,null,96.084538865006905,null,1,null,null,null,"Rh_103_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,32258.46,null,32028.01,33049.93,31697.43,"ug/l",null,103,2,null,null,null,"Si_28_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,23241.72,null,23034.73,21893.4,24797.03,null,88.576508585645101,null,2,null,null,null,"Sc_45_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,633.35,null,850.02,650.02,400.01,"ug/l",null,103,2,null,null,null,"Cu_65_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,43.33,null,50,50,30,"ug/l",null,103,2,null,null,null,"Sr_88_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,253281.2,null,248045.66,255003.3,256794.64,null,93.330311756696702,null,2,null,null,null,"Rh_103_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,533.35,null,300.01,300.01,1000.04,"ug/l",null,103,2,null,null,null,"I_127_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_111SMPL.d"]
["labman","Reagent blank14","CS_2019_10_22_JLS.b","2019-10-22T14:09:59","Sample",1,"111SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tl_205_111SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,23.12,1.02,117145.08,1.02,117741.87,115619.11,118074.27,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,473.34,null,370,540.01,510.01,"ug/l",null,103,1,null,null,null,"Sn_118_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,2810.22,null,2580.19,2750.21,3100.26,"ug/l",null,103,2,null,null,null,"P_31_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,1626.74,null,1540.07,1580.07,1760.09,"ug/l",null,103,2,null,null,null,"Mn_55_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,97971.21,null,92240.71,98271.77,103401.15,"ug/l",null,103,2,null,null,null,"Zn_66_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,466.68,null,300.01,400.01,700.02,"ug/l",null,103,2,null,null,null,"I_127_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,4072.17,null,3970.48,4125.51,4120.51,"ug/l",null,103,2,null,null,null,"Cs_133_112SMPL.d"]
["labman","QC1-4","CS_2019_10_22_JLS.b","2019-10-22T14:11:33","Sample",1,"112SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,46.67,null,60,50,30,"ug/l",null,103,2,null,null,null,"Ir_193_112SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-0.36699999999999999,null,6931.3,null,6921.3,6861.28,7011.32,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.33600000000000002,2.2400000000000002,50121.78,2.2400000000000002,49607.02,49667.22,51091.11,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-2440.4949999999999,null,5377.58,null,5510.97,5320.89,5300.89,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,383.34,null,350,440.01,360,"ug/l",null,103,2,null,null,null,"Mn_55_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,10102.99,null,10102.86,8502.14,11703.96,"ug/l",null,103,2,null,null,null,"Fe_56_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-2376.96,null,-2232.2600000000002,-2352.2800000000002,-2546.34,"ug/l",null,103,2,null,null,null,"In_115_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,4050.49,null,3795.43,3995.47,4360.5600000000004,"ug/l",null,103,2,null,null,null,"Sn_118_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,33.33,null,40,30,30,"ug/l",null,103,2,null,null,null,"Ir_193_113SMPL.d"]
["labman","QC2-4","CS_2019_10_22_JLS.b","2019-10-22T14:13:03","Sample",1,"113SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"U_238_113SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.17100000000000001,9.7200000000000006,39833.300000000003,9.7200000000000006,39071.64,40384.49,40043.769999999997,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-815.86,null,1800.11,null,1970.12,1680.1,1750.1,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,22233.8,null,23495.39,22093.599999999999,21112.400000000001,null,84.735212425607301,null,2,null,null,null,"Sc_45_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,7668.47,null,7601.71,7701.78,7701.93,"ug/l",null,103,2,null,null,null,"Fe_56_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,5,null,10,0,5,"ug/l",null,103,2,null,null,null,"Se_78_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,240627.38,null,245317.73,236594.33,239970.09,null,88.667570683091299,null,2,null,null,null,"Rh_103_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,43.33,null,35,60,35,"ug/l",null,103,2,null,null,null,{"$f":"N7619&\"_\"&P7619&\"_\"&G7619","v":"Sn_118_114SMPL.d"}]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_114SMPL.d"]
["labman","QC3-4","CS_2019_10_22_JLS.b","2019-10-22T14:14:33","Sample",1,"114SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,32.271999999999998,28.57,93.33,28.57,70,90,120,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_114SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Li","lithium",7,9.94,0.35,287896.71999999997,0.35,287743.40999999997,284008.28000000003,291938.46999999997,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Sn","tin",118,null,null,1426.73,null,1400.06,1480.06,1400.06,"ug/l",null,103,1,null,null,null,"Sn_118_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Mg","magnesium",24,null,null,41547.599999999999,null,40444.71,40945.660000000003,43252.42,"ug/l",null,103,2,null,null,null,"Mg_24_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ti","titanium",49,null,null,41.33,null,38,50,36,"ug/l",null,103,2,null,null,null,"Ti_49_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Co","cobalt",59,null,null,4123.8100000000004,null,4155.49,4060.47,4155.4799999999996,"ug/l",null,103,2,null,null,null,"Co_59_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Br","bromine",79,null,null,616.67999999999995,null,550.01,650.01,650.01,"ug/l",null,103,2,null,null,null,"Br_79_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Pd","palladium",105,null,null,28.33,null,25,35,25,"ug/l",null,103,2,null,null,null,"Pd_105_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Sb","antimony",121,null,null,100,null,100,110,90,"ug/l",null,103,2,null,null,null,"Sb_121_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Tb","terbium",159,null,null,13.33,null,0,20,20,"ug/l",null,103,2,null,null,null,"Tb_159_115SMPL.d"]
["labman","QC1-5","CS_2019_10_22_JLS.b","2019-10-22T14:16:02","Sample",1,"115SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Pb","lead",208,null,null,90108.160000000003,null,90034.26,90466.76,89823.47,"ug/l",null,103,2,null,null,null,"Pb_208_115SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Be","beryllium",9,1E-3,86.62,6.67,86.62,10,0,10,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Sb","antimony",121,null,null,33937.29,null,33683.39,35146.480000000003,32981.99,"ug/l",null,103,1,null,null,null,"Sb_121_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Al","aluminium",27,0.504,64.540000000000006,50,64.540000000000006,60,50,40,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","V","vanadium",51,null,null,343.33,null,390,380,260,"ug/l",null,103,2,null,null,null,"V_51_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ga","gallium",71,null,null,11.67,null,10,15,10,"ug/l",null,103,2,null,null,null,"Ga_71_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Rb","rubidium",85,null,null,116.67,null,150,100,100,"ug/l",null,103,2,null,null,null,"Rb_85_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ag","silver",109,null,null,1055.5999999999999,null,933.37,1233.3900000000001,1000.04,"ug/l",null,103,2,null,null,null,"Ag_109_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ba","barium",137,null,null,80,null,70,60,110,"ug/l",null,103,2,null,null,null,"Ba_137_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_116SMPL.d"]
["labman","QC2-5","CS_2019_10_22_JLS.b","2019-10-22T14:17:35","Sample",1,"116SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Bi","bismuth",209,null,null,166.67,null,160,120,220,"ug/l",null,103,2,null,null,null,"Bi_209_116SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","B","boron",10,-1.131,null,626.67999999999995,null,590.01,640.01,650.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Sb","antimony",123,null,null,113.33,null,130,90,120,"ug/l",null,103,1,null,null,null,"Sb_123_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","K","potassium",39,null,null,18192.62,null,18359.48,19110.27,17108.099999999999,"ug/l",null,103,2,null,null,null,"K_39_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Cr","chromium",52,null,null,2910.24,null,2480.17,3110.28,3140.28,"ug/l",null,103,2,null,null,null,"Cr_52_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ge","germanium",74,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"Ge_74_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ag","silver",109,null,null,311.12,null,433.34,133.33000000000001,366.68,"ug/l",null,103,2,null,null,null,"Ag_109_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Cd","cadmium",114,null,null,3.88,null,0.64,4.2300000000000004,6.77,"ug/l",null,103,2,null,null,null,"Cd_114_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ce","cerium",140,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Ce_140_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Bi","bismuth",209,null,null,26.67,null,20,20,40,"ug/l",null,103,2,null,null,null,"Bi_209_117SMPL.d"]
["labman","QC3-5","CS_2019_10_22_JLS.b","2019-10-22T14:19:04","Sample",1,"117SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Th","thorium",232,null,null,110,null,120,120,90,"ug/l",null,103,2,null,null,null,"Th_232_117SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ge","germanium",74,-5.8999999999999997E-2,null,2193.4699999999998,null,2210.15,2140.13,2230.14,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","K","potassium",39,null,null,16090.67,null,15957.19,17558.66,14756.16,"ug/l",null,103,2,null,null,null,"K_39_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ca","calcium",44,null,null,260,null,320,230,230,"ug/l",null,103,2,null,null,null,"Ca_44_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ge","germanium",74,null,null,26.67,null,45,20,15,"ug/l",null,103,2,null,null,null,"Ge_74_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Mo","molybdenum",98,null,null,59.63,null,90,23.9,65,"ug/l",null,103,2,null,null,null,"Mo_98_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ru","ruthenium",101,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Ru_101_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Au","gold",197,null,null,5584.35,null,5400.94,5651.01,5701.09,"ug/l",null,103,2,null,null,null,"Au_197_118SMPL.d"]
["labman","Reagent blank15","CS_2019_10_22_JLS.b","2019-10-22T14:20:33","Sample",1,"118SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,"Annette training","Th","thorium",232,null,null,53.33,null,80,30,50,"ug/l",null,103,2,null,null,null,"Th_232_118SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,349.01900000000001,2.3199999999999998,9422.5400000000009,2.3199999999999998,9082.3799999999992,9692.69,9492.5499999999993,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,1891.5889999999999,37.130000000000003,106.67,37.130000000000003,110,80,130,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,220,null,170,290,200,"ug/l",null,103,2,null,null,null,"Ca_44_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,546.67999999999995,null,550.01,540.01,550.01,"ug/l",null,103,2,null,null,null,"Ni_60_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,133.33000000000001,null,200,100,100,"ug/l",null,103,2,null,null,null,"Rb_85_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0.67,null,0,2,0,"ug/l",null,103,2,null,null,null,"Te_125_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,16.670000000000002,null,50,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_119SMPL.d"]
["labman","Acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:22:04","Sample",1,"119SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,1156.71,null,1350.06,800.02,1320.06,"ug/l",null,103,2,null,null,null,"Hg_202_119SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,3.45,24.27,423.34,24.27,510.01,400.01,360,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,63533.99,null,64932.65,63968.9,61700.43,null,7.6512498287117303,null,1,null,null,null,"Rh_103_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,29483.78,null,29924.44,29553.98,28972.91,"ug/l",null,103,2,null,null,null,"Si_28_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,19443.95,null,18159.169999999998,20161.34,20011.330000000002,"ug/l",null,103,2,null,null,null,"K_39_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,213.33,null,170,240,230,"ug/l",null,103,2,null,null,null,"Ca_44_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1126.72,null,1580.09,1160.04,640.02,null,4.2940293790041997,null,2,null,null,null,"Sc_45_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,10.67,null,10,6,16,"ug/l",null,103,2,null,null,null,"Ti_49_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,116.67,null,150,100,100,"ug/l",null,103,2,null,null,null,"V_51_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2803.56,null,2630.2,2880.24,2900.24,"ug/l",null,103,2,null,null,null,"Cr_52_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,400.01,null,400.01,320,480.01,"ug/l",null,103,2,null,null,null,"Mn_55_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,11403.86,null,12704.84,9902.81,11603.94,"ug/l",null,103,2,null,null,null,"Fe_56_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,118.33,null,130,100,125,"ug/l",null,103,2,null,null,null,"Co_59_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,496.68,null,560.01,510.01,420.01,"ug/l",null,103,2,null,null,null,"Ni_60_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,1250.05,null,1350.06,1025.04,1375.06,"ug/l",null,103,2,null,null,null,{"$f":"N8003&\"_\"&P8003&\"_\"&G8003","v":"Cu_65_120SMPL.d"}]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,5300.85,null,4200.5200000000004,5300.85,6401.17,"ug/l",null,103,2,null,null,null,"Zn_66_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,16.670000000000002,null,5,30,15,"ug/l",null,103,2,null,null,null,"Ga_71_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,8.33,null,10,5,10,"ug/l",null,103,2,null,null,null,"Ge_74_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,3.33,null,0,0,10,"ug/l",null,103,2,null,null,null,"As_75_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Se_78_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,603.34,null,570.01,680.01,560.01,"ug/l",null,103,2,null,null,null,"Br_79_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,166.67,null,450.01,50,0,"ug/l",null,103,2,null,null,null,"Rb_85_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,40,null,40,30,50,"ug/l",null,103,2,null,null,null,"Sr_88_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,450.01,null,450.01,510.01,390,"ug/l",null,103,2,null,null,null,"Zr_90_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,76.67,null,75,85,70,"ug/l",null,103,2,null,null,null,"Mo_98_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,14235.93,null,15276.79,14125.84,13305.15,null,5.2457248068371198,null,2,null,null,null,"Rh_103_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,921.69,null,975.03,995.03,795.02,"ug/l",null,103,2,null,null,null,"Pd_105_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,11.11,null,0,33.33,0,"ug/l",null,103,2,null,null,null,"Ag_109_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,40.71,null,39.69,35.47,46.97,"ug/l",null,103,2,null,null,null,"Cd_114_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-351.01,null,-290,-393.01,-370.01,"ug/l",null,103,2,null,null,null,"In_115_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,610.01,null,525.01,680.01,625.01,"ug/l",null,103,2,null,null,null,"Sn_118_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,25,null,45,15,15,"ug/l",null,103,2,null,null,null,"Sb_121_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,1600.1,null,1600.1,1000.04,2200.17,"ug/l",null,103,2,null,null,null,"I_127_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,15,null,15,20,10,"ug/l",null,103,2,null,null,null,"Cs_133_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,11.67,null,10,5,20,"ug/l",null,103,2,null,null,null,"Ba_137_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,20,null,30,0,30,"ug/l",null,103,2,null,null,null,"Ce_140_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,3.33,null,0,10,0,"ug/l",null,103,2,null,null,null,"Nd_146_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,6.67,null,10,0,10,"ug/l",null,103,2,null,null,null,"Tb_159_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,286.67,null,360,340,160,"ug/l",null,103,2,null,null,null,"W_184_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,53.33,null,70,20,70,"ug/l",null,103,2,null,null,null,"Ir_193_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,17.667000000000002,183.35,3.33,183.35,10,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,49794.33,null,52084.58,49827.839999999997,47470.57,"ug/l",null,103,2,null,null,null,"Au_197_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,126.67,null,100,150,130,"ug/l",null,103,2,null,null,null,"Hg_202_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,36.67,null,50,50,10,"ug/l",null,103,2,null,null,null,"Tl_205_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,5367.59,null,5240.88,5490.97,5370.91,"ug/l",null,103,2,null,null,null,"Pb_208_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,116.67,null,80,170,100,"ug/l",null,103,2,null,null,null,"Bi_209_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,1210.05,null,1070.04,970.03,1590.08,"ug/l",null,103,2,null,null,null,"Th_232_120SMPL.d"]
["labman","Base rinse","CS_2019_10_22_JLS.b","2019-10-22T14:23:40","Sample",1,"120SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,20,null,20,20,20,"ug/l",null,103,2,null,null,null,"U_238_120SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,6.9749999999999996,66.5,166.67,66.5,290,130,80,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,769.77200000000005,10.32,620.01,10.32,660.01,580.01,620.01,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,764.67399999999998,12.89,2820.22,12.89,2770.21,2640.19,3050.25,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,15568.478999999999,5,741534.81,5,741942.1,745136,737526.32,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,26880.09,null,28071.87,26339.26,26229.14,null,1.39054887068419,null,1,null,null,null,"Sc_45_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,177.04400000000001,14.25,1530.07,14.25,1640.08,1360.05,1590.07,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,633.34,null,630.01,670.01,600.01,null,7.6272052431182305E-2,null,1,null,null,null,"Rh_103_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,606.67999999999995,null,790.02,530.01,500.01,"ug/l",null,103,1,null,null,null,"Sn_118_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,656.68,null,730.02,560.01,680.01,"ug/l",null,103,1,null,null,null,"Sb_121_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,476.68,null,460.01,380.01,590.01,"ug/l",null,103,1,null,null,null,"Sb_123_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,2420.547,64.05,16.670000000000002,64.05,10,30,10,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-33591.743000000002,null,60,null,60,100,20,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,33.33,null,0,100,0,"ug/l",null,103,2,null,null,null,"Mg_24_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,21289.263999999999,46.94,446.68,46.94,450.01,420.01,470.01,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,27140.18,null,27050.11,27991.39,26379.05,"ug/l",null,103,2,null,null,null,"Si_28_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"P_31_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,426.67,null,450.01,340,490.01,"ug/l",null,103,2,null,null,null,"S_34_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,18859.98,null,20862.04,16507.54,19210.349999999999,"ug/l",null,103,2,null,null,null,"K_39_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,200,null,170,240,190,"ug/l",null,103,2,null,null,null,"Ca_44_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,553.35,null,340,720.02,600.02,null,2.10885923119451,null,2,null,null,null,"Sc_45_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,7.33,null,8,4,10,"ug/l",null,103,2,null,null,null,"Ti_49_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,56.67,null,60,60,50,"ug/l",null,103,2,null,null,null,"V_51_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2713.55,null,2520.19,2920.25,2700.21,"ug/l",null,103,2,null,null,null,"Cr_52_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,463.34,null,550.01,590.01,250,"ug/l",null,103,2,null,null,null,"Mn_55_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,10002.969999999999,null,9502.7099999999991,10203,10303.19,"ug/l",null,103,2,null,null,null,{"$f":"N8067&\"_\"&P8067&\"_\"&G8067","v":"Fe_56_121SMPL.d"}]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,110,null,115,120,95,"ug/l",null,103,2,null,null,null,"Co_59_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,416.67,null,360,330,560.01,"ug/l",null,103,2,null,null,null,"Ni_60_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,466.68,null,500.01,475.01,425.01,"ug/l",null,103,2,null,null,null,"Cu_65_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,666.71,null,1000.08,400.01,600.04,"ug/l",null,103,2,null,null,null,"Zn_66_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,11.67,null,10,15,10,"ug/l",null,103,2,null,null,null,"Ga_71_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Ge_74_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"As_75_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,0,0,5,"ug/l",null,103,2,null,null,null,"Se_78_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,263.33,null,220,200,370,"ug/l",null,103,2,null,null,null,"Br_79_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,100,null,100,0,200,"ug/l",null,103,2,null,null,null,"Rb_85_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Sr_88_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,6.67,null,10,10,0,"ug/l",null,103,2,null,null,null,"Y_89_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,36.67,null,70,40,0,"ug/l",null,103,2,null,null,null,"Zr_90_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,21.67,null,10,30,25,"ug/l",null,103,2,null,null,null,"Mo_98_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,180,null,210,220,110,null,6.6327292022484996E-2,null,2,null,null,null,"Rh_103_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,76.67,null,75,85,70,"ug/l",null,103,2,null,null,null,"Pd_105_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ag_109_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,2.38,null,-1.36,9.18,-0.68,"ug/l",null,103,2,null,null,null,"Cd_114_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-21,null,-30,-18,-15,"ug/l",null,103,2,null,null,null,"In_115_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,35,null,50,30,25,"ug/l",null,103,2,null,null,null,"Sn_118_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,56.67,null,55,70,45,"ug/l",null,103,2,null,null,null,"Sb_121_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,2266.86,null,2200.17,2100.16,2500.2600000000002,"ug/l",null,103,2,null,null,null,"I_127_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,15,null,30,10,5,"ug/l",null,103,2,null,null,null,"Cs_133_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ba_137_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,40,null,20,20,80,"ug/l",null,103,2,null,null,null,"W_184_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ir_193_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,-1.034,null,0,null,0,0,0,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,866.7,null,1000.05,950.04,650.02,"ug/l",null,103,2,null,null,null,"Au_197_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,63.33,null,70,60,60,"ug/l",null,103,2,null,null,null,"Hg_202_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,13.33,null,10,10,20,"ug/l",null,103,2,null,null,null,"Tl_205_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,76.67,null,80,40,110,"ug/l",null,103,2,null,null,null,"Pb_208_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,13.33,null,10,20,10,"ug/l",null,103,2,null,null,null,"Bi_209_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"Th_232_121SMPL.d"]
["labman","2% Nitric acid rinses ","CS_2019_10_22_JLS.b","2019-10-22T14:25:08","Sample",1,"121SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,6.67,null,20,0,0,"ug/l",null,103,2,null,null,null,"U_238_121SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,-2.1000000000000001E-2,null,176.67,null,110,210,210,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0,"N/A",0,"N/A",0,0,0,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,-1.3220000000000001,null,433.34,null,460.01,510.01,330,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,-1.4179999999999999,null,1960.11,null,1890.1,2080.12,1910.1,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,0.26300000000000001,3.53,43575.17,3.53,44043.01,42829.97,43852.53,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,1600313.41,null,1583649.2,1672701.69,1544589.35,null,82.786702338103296,null,1,null,null,null,"Sc_45_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,-5.8000000000000003E-2,null,2250.15,null,2470.1799999999998,2230.14,2050.12,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,757744.49,null,760006.94,755363.58,757862.96,null,91.2533924068983,null,1,null,null,null,"Rh_103_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1276.72,null,1490.07,1160.04,1180.04,"ug/l",null,103,1,null,null,null,"Sn_118_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,230,null,100,230,360,"ug/l",null,103,1,null,null,null,"Sb_121_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,156.66999999999999,null,190,100,180,"ug/l",null,103,1,null,null,null,"Sb_123_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,-1.764,null,36.67,null,60,40,10,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-24.093,null,80,null,80,70,90,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,66.67,null,200.01,0,0,"ug/l",null,103,2,null,null,null,"Mg_24_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,-0.80700000000000005,null,10,null,30,0,0,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,30782.560000000001,null,30956.21,30154.799999999999,31236.67,"ug/l",null,103,2,null,null,null,"Si_28_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,20,null,40,0,20,"ug/l",null,103,2,null,null,null,"P_31_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,416.67,null,460.01,320,470.01,"ug/l",null,103,2,null,null,null,"S_34_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,16390.88,null,15957.19,17458.53,15756.93,"ug/l",null,103,2,null,null,null,"K_39_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,260,null,410.01,200,170,"ug/l",null,103,2,null,null,null,"Ca_44_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,23221.64,null,23535.439999999999,23134.83,22994.639999999999,null,88.499969003057004,null,2,null,null,null,"Sc_45_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,14.67,null,8,16,20,"ug/l",null,103,2,null,null,null,"Ti_49_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,316.67,null,260,390.01,300,"ug/l",null,103,2,null,null,null,{"$f":"N8131&\"_\"&P8131&\"_\"&G8131","v":"V_51_122SMPL.d"}]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2673.54,null,2710.21,2780.22,2530.1799999999998,"ug/l",null,103,2,null,null,null,"Cr_52_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,440.01,null,460.01,450.01,410.01,"ug/l",null,103,2,null,null,null,"Mn_55_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,13338.67,null,14305.95,12104.87,13605.18,"ug/l",null,103,2,null,null,null,"Fe_56_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,86.67,null,80,80,100,"ug/l",null,103,2,null,null,null,"Co_59_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,480.01,null,570.01,360,510.01,"ug/l",null,103,2,null,null,null,"Ni_60_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,783.35,null,925.03,750.02,675.01,"ug/l",null,103,2,null,null,null,"Cu_65_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,1066.71,null,1300.06,900.03,1000.03,"ug/l",null,103,2,null,null,null,"Zn_66_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,16.670000000000002,null,10,15,25,"ug/l",null,103,2,null,null,null,"Ga_71_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,40,null,45,35,40,"ug/l",null,103,2,null,null,null,"Ge_74_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,6.67,null,0,20,0,"ug/l",null,103,2,null,null,null,"As_75_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,3.33,null,10,0,0,"ug/l",null,103,2,null,null,null,"Se_78_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,496.68,null,560.01,490.01,440.01,"ug/l",null,103,2,null,null,null,"Br_79_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,116.67,null,0,100,250,"ug/l",null,103,2,null,null,null,"Rb_85_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,16.670000000000002,null,20,20,10,"ug/l",null,103,2,null,null,null,"Sr_88_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"Y_89_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,40,null,50,30,40,"ug/l",null,103,2,null,null,null,"Zr_90_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,58.33,null,80,35,60,"ug/l",null,103,2,null,null,null,"Mo_98_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ru_101_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,239924,null,236280.19,241987.46,241504.35,null,88.4083845066816,null,2,null,null,null,"Rh_103_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,196.67,null,165,195,230,"ug/l",null,103,2,null,null,null,"Pd_105_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,44.44,null,0,133.33000000000001,0,"ug/l",null,103,2,null,null,null,"Ag_109_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,1.02,null,0.64,1.55,0.87,"ug/l",null,103,2,null,null,null,"Cd_114_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-46,null,-25,-44,-69,"ug/l",null,103,2,null,null,null,"In_115_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,85,null,50,90,115,"ug/l",null,103,2,null,null,null,"Sn_118_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,11.67,null,5,25,5,"ug/l",null,103,2,null,null,null,"Sb_121_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Te_125_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,666.69,null,1000.04,1000.04,0,"ug/l",null,103,2,null,null,null,"I_127_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Cs_133_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,5,null,10,0,5,"ug/l",null,103,2,null,null,null,"Ba_137_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Ce_140_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,13.33,null,0,40,0,"ug/l",null,103,2,null,null,null,"W_184_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,90,null,90,40,140,"ug/l",null,103,2,null,null,null,"Ir_193_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,4.9589999999999996,111.4,16.670000000000002,111.4,30,0,20,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,8185.46,null,8252.15,8102.11,8202.1299999999992,"ug/l",null,103,2,null,null,null,"Au_197_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,30,null,30,20,40,"ug/l",null,103,2,null,null,null,"Hg_202_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,6.67,null,0,10,10,"ug/l",null,103,2,null,null,null,"Tl_205_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,93.33,null,70,60,150,"ug/l",null,103,2,null,null,null,"Pb_208_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,26.67,null,30,50,0,"ug/l",null,103,2,null,null,null,"Bi_209_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,33.33,null,30,40,30,"ug/l",null,103,2,null,null,null,"Th_232_122SMPL.d"]
["labman","Reagent Blank","CS_2019_10_22_JLS.b","2019-10-22T14:45:38","Sample",1,"122SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"U_238_122SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Li","lithium",7,19.997,3.51,688697.31,3.51,688999.44,675540.93,701551.55,"ug/l",null,103,1,1,3.7499500373748801E-2,1.0090470682887799E-3,"Li_7_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Be","beryllium",9,0.48899999999999999,4.5199999999999996,5467.48,4.5199999999999996,5560.84,5320.77,5520.82,"ug/l",null,103,1,1,1.2204884423180001E-2,0,"Be_9_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",10,48.02,2.39,58161.02,2.39,56505.86,58170.96,59806.239999999998,"ug/l",null,103,1,1,1.27304377014654E-3,2.2549556097869899E-3,"B_10_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"B","boron",11,48.77,1.64,271105.5,1.64,267243.02,271904.74,274168.75,"ug/l",null,103,1,1,5.8363444394679701E-3,1.08603656272137E-2,"B_11_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,21.725000000000001,2.36,1536143.1,2.36,1518959.04,1527516.23,1561954.04,"ug/l",null,103,1,1,7.5341777328951101E-2,3.77238202851055E-2,"Al_27_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,2055088.88,null,2058595.13,2001900.44,2104771.06,null,106.312944510397,null,1,null,null,null,"Sc_45_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,4.399,1.39,58701.120000000003,1.39,57864.9,59088.92,59149.54,"ug/l",null,103,1,1,1.36914385749496E-2,3.75824085177478E-3,"Ge_74_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,917638.97,null,895001.78,937234.67,920680.46,null,110.50910927627901,null,1,null,null,null,"Rh_103_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,1700.09,null,1760.1,1770.1,1570.08,"ug/l",null,103,1,null,null,null,"Sn_118_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,913.36,null,870.02,920.03,950.03,"ug/l",null,103,1,null,null,null,"Sb_121_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",123,null,null,1530.07,null,1460.07,1500.07,1630.08,"ug/l",null,103,1,null,null,null,"Sb_123_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,582.26300000000003,6.53,5730.98,6.53,5801,5420.87,5971.06,"ug/l",null,103,1,1,1.0619264545473301E-5,6.7102627168610498E-5,"Te_125_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,-16418.009999999998,null,41544.089999999997,null,41283.449999999997,42265.99,41082.83,"ug/l",null,103,1,-1,-2.75567080550751E-6,3.9170091822160301E-5,"Hg_202_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mg","magnesium",24,null,null,266.67,null,0,500.01,300.01,"ug/l",null,103,2,null,null,null,"Mg_24_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Al","aluminium",27,23.878,13.26,953.36,13.26,840.02,1080.03,940.03,"ug/l",null,103,2,1,1.30338258777746E-4,1.4750865401231601E-4,"Al_27_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Si","silicon",28,null,null,57189.34,null,56246.33,56727.77,58593.93,"ug/l",null,103,2,null,null,null,"Si_28_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"P","phosphorus",31,null,null,43.33,null,60,10,60,"ug/l",null,103,2,null,null,null,"P_31_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"S","sulfur",34,null,null,366.67,null,310,390,400.01,"ug/l",null,103,2,null,null,null,"S_34_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"K","potassium",39,null,null,17491.97,null,17608.59,17758.93,17108.400000000001,"ug/l",null,103,2,null,null,null,"K_39_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ca","calcium",44,null,null,233.33,null,260,170,270,"ug/l",null,103,2,null,null,null,{"$f":"N8195&\"_\"&P8195&\"_\"&G8195","v":"Ca_44_123SMPL.d"}]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sc","scandium",45,null,null,28429.11,null,28923.24,27881.56,28482.54,null,108.34618097336499,null,2,null,null,null,"Sc_45_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ti","titanium",49,null,null,17.329999999999998,null,14,8,30,"ug/l",null,103,2,null,null,null,"Ti_49_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"V","vanadium",51,null,null,500.01,null,490.01,450.01,560.01,"ug/l",null,103,2,null,null,null,"V_51_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cr","chromium",52,null,null,2803.56,null,2970.25,2780.22,2660.2,"ug/l",null,103,2,null,null,null,"Cr_52_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mn","manganese",55,null,null,483.34,null,470.01,360,620.01,"ug/l",null,103,2,null,null,null,"Mn_55_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Fe","iron",56,null,null,288548.08,null,285969.39,295836.73,283838.12,"ug/l",null,103,2,null,null,null,"Fe_56_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Co","cobalt",59,null,null,145,null,110,130,195,"ug/l",null,103,2,null,null,null,"Co_59_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ni","nickel",60,null,null,1043.3699999999999,null,930.03,1120.04,1080.04,"ug/l",null,103,2,null,null,null,"Ni_60_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cu","copper",65,null,null,1033.3699999999999,null,875.03,1075.04,1150.05,"ug/l",null,103,2,null,null,null,"Cu_65_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zn","zinc",66,null,null,209433.4,null,214795.46,195876.52,217628.23,"ug/l",null,103,2,null,null,null,"Zn_66_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ga","gallium",71,null,null,15,null,5,20,20,"ug/l",null,103,2,null,null,null,"Ga_71_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ge","germanium",74,null,null,2243.48,null,2240.14,2145.14,2345.16,"ug/l",null,103,2,null,null,null,"Ge_74_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"As","arsenic",75,null,null,23.33,null,20,10,40,"ug/l",null,103,2,null,null,null,"As_75_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Se","selenium",78,null,null,1.67,null,5,0,0,"ug/l",null,103,2,null,null,null,"Se_78_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Br","bromine",79,null,null,626.67999999999995,null,510.01,680.01,690.01,"ug/l",null,103,2,null,null,null,"Br_79_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rb","rubidium",85,null,null,100,null,50,50,200.01,"ug/l",null,103,2,null,null,null,"Rb_85_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sr","strontium",88,null,null,20,null,10,30,20,"ug/l",null,103,2,null,null,null,"Sr_88_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Y","yttrium",89,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Y_89_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Zr","zirconium",90,null,null,523.34,null,560.01,560.01,450.01,"ug/l",null,103,2,null,null,null,"Zr_90_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Mo","molybdenum",98,null,null,58.74,null,232.38,-96.38,40.21,"ug/l",null,103,2,null,null,null,"Mo_98_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ru","ruthenium",101,null,null,12055.89,null,11749.01,12239.35,12179.32,"ug/l",null,103,2,null,null,null,"Ru_101_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Rh","rhodium",103,null,null,292507.69,null,286096.05,290001.87,301425.15000000002,null,107.784683185847,null,2,null,null,null,"Rh_103_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pd","palladium",105,null,null,6351.18,null,6651.3,6571.26,5830.99,"ug/l",null,103,2,null,null,null,"Pd_105_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ag","silver",109,null,null,477.79,null,633.35,433.34,366.67,"ug/l",null,103,2,null,null,null,"Ag_109_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cd","cadmium",114,null,null,4.5,null,9.27,4.18,0.05,"ug/l",null,103,2,null,null,null,"Cd_114_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"In","indium",115,null,null,-37,null,0,-54,-57,"ug/l",null,103,2,null,null,null,"In_115_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sn","tin",118,null,null,128.33000000000001,null,100,140,145,"ug/l",null,103,2,null,null,null,"Sn_118_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Sb","antimony",121,null,null,71.67,null,85,55,75,"ug/l",null,103,2,null,null,null,"Sb_121_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Te","tellurium",125,null,null,189.33,null,176,188,204,"ug/l",null,103,2,null,null,null,"Te_125_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"I","iodine",127,null,null,833.38,null,500.02,700.04,1300.08,"ug/l",null,103,2,null,null,null,"I_127_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Cs","cesium",133,null,null,5,null,5,10,0,"ug/l",null,103,2,null,null,null,"Cs_133_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ba","barium",137,null,null,21.67,null,35,15,15,"ug/l",null,103,2,null,null,null,"Ba_137_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ce","cerium",140,null,null,26.67,null,0,0,80,"ug/l",null,103,2,null,null,null,"Ce_140_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Nd","neodymium",146,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Nd_146_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Gd","gadolinium",156,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Gd_156_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tb","terbium",159,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Tb_159_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Er","erbium",166,null,null,0,null,0,0,0,"ug/l",null,103,2,null,null,null,"Er_166_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"W","tungsten",184,null,null,36.67,null,60,30,20,"ug/l",null,103,2,null,null,null,"W_184_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Ir","iridium",193,null,null,110,null,70,180,80,"ug/l",null,103,2,null,null,null,"Ir_193_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pt","platinum",195,8949.9290000000001,2.91,30562.51,2.91,29577.25,29597.29,32513,"ug/l",null,103,2,1,1.1667476055621E-5,1.2067361848075E-5,"Pt_195_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Au","gold",197,null,null,171733.25,null,175456.13,172272.29,167471.32999999999,"ug/l",null,103,2,null,null,null,"Au_197_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Hg","mercury",202,null,null,25944.44,null,26171.45,25861.05,25800.82,"ug/l",null,103,2,null,null,null,"Hg_202_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Tl","thallium",205,null,null,70,null,80,80,50,"ug/l",null,103,2,null,null,null,"Tl_205_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Pb","lead",208,null,null,216.67,null,200,220,230,"ug/l",null,103,2,null,null,null,"Pb_208_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Bi","bismuth",209,null,null,523.34,null,490.01,520.01,560.01,"ug/l",null,103,2,null,null,null,"Bi_209_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"Th","thorium",232,null,null,113.33,null,150,30,160,"ug/l",null,103,2,null,null,null,"Th_232_123SMPL.d"]
["labman","NMS 55","CS_2019_10_22_JLS.b","2019-10-22T14:47:07","Sample",1,"123SMPL.d","N:\\Lab\\MassICP_Analysis\\WonderWoman\\D-DRIVE\\Validation\\Container Screening","G3663A",1,1,1,null,"U","uranium",238,null,null,13.33,null,10,20,10,"ug/l",null,103,2,null,null,null,"U_238_123SMPL.d"]

{"sheet":"SQ","index":1,"n_rows":125,"n_cols":117,"rows_kept":125,"cut":false,"hidden":false,"merged":[]}
["Sample",null,null,null,null,null,null,null,null,"24  Mg  [ He ] ",null,"27  Al  [ He ] ",null,"28  Si  [ He ] ",null,"31  P  [ He ] ",null,"34  S  [ He ] ",null,"39  K  [ He ] ",null,"44  Ca  [ He ] ",null,"45  Sc  [ He ] ",null,"49  Ti  [ He ] ",null,"51  V  [ He ] ",null,"52  Cr  [ He ] ",null,"55  Mn  [ He ] ",null,"56  Fe  [ He ] ",null,"59  Co  [ He ] ",null,"60  Ni  [ He ] ",null,"65  Cu  [ He ] ",null,"66  Zn  [ He ] ",null,"71  Ga  [ He ] ",null,"74  Ge  [ He ] ",null,"75  As  [ He ] ",null,"78  Se  [ He ] ",null,"79  Br  [ He ] ",null,"85  Rb  [ He ] ",null,"88  Sr  [ He ] ",null,"89  Y  [ He ] ",null,"90  Zr  [ He ] ",null,"98  Mo  [ He ] ",null,"101  Ru  [ He ] ",null,"103  Rh  [ He ] ",null,"105  Pd  [ He ] ",null,"109  Ag  [ He ] ",null,"114  Cd  [ He ] ",null,"115  In  [ He ] ",null,"118  Sn  [ He ] ",null,"121  Sb  [ He ] ",null,"125  Te  [ He ] ",null,"127  I  [ He ] ",null,"133  Cs  [ He ] ",null,"137  Ba  [ He ] ",null,"140  Ce  [ He ] ",null,"146  Nd  [ He ] ",null,"156  Gd  [ He ] ",null,"159  Tb  [ He ] ",null,"166  Er  [ He ] ",null,"184  W  [ He ] ",null,"193  Ir  [ He ] ",null,"195  Pt  [ He ] ",null,"197  Au  [ He ] ",null,"202  Hg  [ He ] ",null,"205  Tl  [ He ] ",null,"208  Pb  [ He ] ",null,"209  Bi  [ He ] ",null,"232  Th  [ He ] ",null,"238  U  [ He ] "]
[null,"Rjct","Data File","Acq. Date-Time","Type","Sample Name","SQBlk","SQISTD","SQStd","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS","Conc.","CPS"]
[null,false,"001SMPL.d","2019-10-22T11:22:58","Sample","2% Nitric acid rinses ",false,false,false,"<0.002",0,3.6567695034817498,490.00711346655424,891.36190962184696,58362.990666781996,0.55555634200571802,10.000014156102923,849.68450938158105,433.33909978460616,77.176073623918896,20728.629003360023,19.9746438949173,296.66980301335752,0.17399542529463899,380.00600884349132,4.8287087693035198E-2,6.6666698455810547,7.6207544982561303E-3,53.333545525865567,0.117981180044366,1226.7094492905917,9.8959606735844696E-2,380.00488986564346,0.90403326478725599,7134.933376844725,7.9969448907705101E-3,165.00096393126799,5.38841190358854E-2,353.33755602432956,5.4308793918635002E-2,525.01158298329574,"<0.000",0,7.1126285550350103E-3,11.666679605841443,1.7937208632075999E-2,40.000075697897735,"<0.000",0,2.53585867885067E-2,3.3333380520343074,1.7154649740647201,80.000365972513464,7.8550858553830899E-2,166.66905457780135,"<0.000",0,2.6062101801452802E-4,6.6666856408116368,3.6328311990672501E-4,3.3333380520343074,"<0.000",0,"<0.000",0,8.6908666503242098E-5,6.6666768987972,5.9943075314910599E-4,3.3333384493986,"<0.000",0,"<0.000",-0.2562541339359492,"<0.000",-35.0001015216097,1.4702019460851199E-2,58.333502536016169,7.2483974038567199E-3,40.000111460684067,"<0.005",0,7.8617516061633405E-2,100.00148043039771,7.2982138130533905E-4,5.0000062833228762,"<0.001",0,8.7634017385859498E-5,3.3333384493986,"<0.000",0,"<0.000",0,1.0268299467311201E-3,66.666934291517734,"<0.000",0,8.4555863367139302E-5,3.3333384493986,"<0.000",0,2.2438809648298001E-4,6.6666872302688072,4.04092376138687E-2,366.67342210824495,5.6680493558219696E-4,3.3333384493986,6.1091781303237698E-5,3.3333384493986,2.0709293902157798E-3,76.667048583422385,2.20449426210491E-4,10.000026871760285,"<0.000",0,"<0.000",0]
[null,false,"002SMPL.d","2019-10-22T11:24:26","Sample","2% Nitric acid rinses ",false,false,false,0.12126604149300201,66.668480138024492,3.2338762513137298,433.33941767604028,910.25832704624395,59600.256273345301,1.8518559303547999,33.333406746386387,666.675071505814,340.00428646796519,79.848084656795606,21446.301240267134,17.5058666370604,260.00273317094542,0.17094287279902801,373.3392341930782,7.2430645354592801E-2,10.000006675720215,9.5259375584959092E-3,66.666892965631305,0.124072864083328,1290.0477407865562,9.5487265356074796E-2,366.67109896732723,0.89134909668528695,7034.8256729876775,5.4928421296273197E-3,113.33381166060053,4.5750591845245298E-2,300.00309179221949,5.9480903260690401E-2,575.01117084200314,"<0.000",0,5.0804483449422297E-3,8.333341553807136,1.4200294265425999E-2,31.666735733547778,1.3544648728298699E-3,3.3333380520343074,2.5358568650658001E-2,3.333335667848552,1.28659790403408,60.000235935048721,3.92755978263611E-2,83.334884916777327,"<0.000",0,1.3031032259711901E-4,3.3333380520343074,1.08984935972018E-3,10.000014156102923,1.6160965057211899E-4,1.666667833924276,"<0.000",0,3.0418116676584603E-4,23.333433121441306,"<0.000",0,"<0.000",0,"<0.000",0.78833285458387359,"<0.000",-12.000021040439293,5.0406866036726804E-3,20.000035067398823,6.6443623787055001E-3,36.666758308806642,"<0.005",0,0.13103074290694,166.67110497762746,1.9461914134977401E-3,13.333357373873035,"<0.001",0,1.75268034771719E-4,6.6666768987972,"<0.000",0,"<0.000",0,6.6743923647380801E-4,43.33349242806198,"<0.000",0,1.6911198880982699E-4,6.6666872302688072,"<0.000",0,"<0.000",0,4.59199608435596E-2,416.67772469445941,1.7004167662275499E-3,10.000026871760285,6.7201137131357605E-4,36.666819900271996,2.1609697350177901E-3,80.0003961721997,2.20449172174856E-4,10.000015348195801,"<0.000",0,"<0.000",0]
[null,false,"003SMPL.d","2019-10-22T11:25:54","Sample","2% Nitric acid rinses ",false,false,false,"<0.002",0,2.9851132987144702,400.00518202773952,892.48451952705796,58436.494897457538,1.4814849915327299,26.666729847589185,712.42717853155796,363.33786105109465,79.661195512964596,21396.104909392558,16.608120287801398,246.66911711291792,0.18620561198833099,406.67305658251462,4.3458387558131702E-2,6.0000040531158447,8.0970454371164704E-3,56.666848609842134,0.141065652203785,1466.7302739610732,7.3785411543202603E-2,283.33598032589782,0.93360732125216395,7368.3417377746982,5.9775101671299203E-3,123.33396727839163,3.8125500152966299E-2,250.00262205795514,3.7067614478474301E-2,358.33841171126988,1.4535675001192299E-2,33.333802968274121,4.0643575131326E-3,6.6666713356971039,7.4738388901196499E-3,16.666702578464609,"<0.000",0,2.53585867885067E-2,3.3333380520343074,1.2865971201234701,60.000199377533797,4.7130627498595597E-2,100.00167116526524,"<0.000",0,5.2124166320876696E-4,13.333361744880252,1.4531387157860399E-3,13.333409428595361,7.7262392732444201E-4,7.9680108386673902,3.3446732795692E-4,3.3333384493986,1.3036310853805901E-4,10.000023692845945,"<0.000",0,"<0.000",0,"<0.000",-0.19625189270200138,"<0.000",-19.0000469237559,7.9810928393357801E-3,31.6667448729265,6.3423429219636503E-3,35.00007100403203,"<0.005",0,0.13103074290694,166.67110497762746,1.94619129749591E-3,13.33335657914445,1.1611022972144E-3,1.666667833924276,"<0.000",0,"<0.000",0,"<0.000",0,5.1341370033093598E-5,3.3333384493986,"<0.000",0,1.6911198880982699E-4,6.6666872302688072,"<0.000",0,"<0.000",0,3.1225418027561E-2,283.33944318208842,1.70041669865924E-3,10.000026474395993,3.0545911043169201E-4,16.666703373193194,1.4406436062514699E-3,53.333490689593191,5.1438165577696502E-4,23.333380669354685,"<0.000",0,"<0.000",0]
[null,false,"004SMPL.d","2019-10-22T11:27:22","Sample","2% Nitric acid rinses ",false,false,false,"<0.002",0,3.3582566363442599,450.00638927013097,926.29565757115802,60650.319734260644,1.29629958834908,23.333392590283466,725.49881196866204,370.00439410401788,73.2006220676314,19660.867240398638,14.588195978279099,216.66855489211449,0.131259087759693,286.66984766716979,8.2088069800719796E-2,11.333341598510742,1.19074376758527E-2,83.333726276948155,0.132088331793105,1373.3885751159703,0.120661611216011,463.34058706948281,0.94205654301074604,7435.0258263827172,4.76584276958827E-3,98.333633864914873,5.2359059781515498E-2,343.33719377710628,5.5170671290644999E-2,533.34348599138696,"<0.000",0,1.1176987763948499E-2,18.333353723088596,7.4738310497676503E-3,16.666685094435735,2.70893362080928E-3,6.6666856408116368,3.8037871113835699E-2,5.0000058859585836,1.6439874926123701,76.667027920479157,5.4985994269721498E-2,116.66917266950645,"<0.000",0,2.6062101801452802E-4,6.6666856408116368,2.5429839180667901E-3,23.333385437726196,1.6160965057211899E-4,1.666667833924276,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0.47041507039843589,"<0.000",-19.000052645801713,7.9810952429182104E-3,31.666754409669522,1.1174609586015899E-2,61.666821514564425,"<0.005",0,0.18344242028136201,233.33875859789259,1.7029164790445699E-3,11.666680797934321,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,1.0268274006618699E-4,6.6666768987972,"<0.000",0,8.4555863367139302E-5,3.3333384493986,"<0.000",0,"<0.000",0,4.2246118279338697E-2,383.3412772667196,3.9676471167808397E-3,23.333443055548624,1.2218356260647499E-4,6.6666768987972,1.8008094291697E-3,66.667045553519642,2.2044941745064101E-4,10.000026474395993,"<0.000",0,"<0.000",0]
[null,false,"005SMPL.d","2019-10-22T11:28:55","Sample","2% Nitric acid rinses ",false,false,false,6.0632196774693302E-2,33.333787073701828,3.1094923552434901,416.67197560262713,930.02417614940498,60894.448962393661,0.37037089467047901,6.6666761040686149,921.58192775977602,470.00678315748593,76.866184949536702,20645.396376174111,10.3239332990896,153.33436101673502,0.15262725798271301,333.33793143424492,3.3800960809498001E-2,4.6666688124338789,3.81037406951421E-3,26.666750510532399,0.136576983437772,1420.05933552892,9.8091431983936397E-2,376.67109881831561,0.86177207572249304,6801.3939152472849,5.0889553081547396E-3,105.00041487315673,4.8292311699338203E-2,316.67006338631023,3.8791628181296403E-2,375.00472112258893,1.4535675001192299E-2,33.333802968274121,5.0804512519951796E-3,8.333346322178647,2.9895312082162998E-3,6.6666713356971039,"<0.000",0,"<0.008",0,1.1436421572187501,53.333523273465197,7.0695827945210593E-2,150.00226634249191,"<0.000",0,"<0.000",0,7.2656623981345002E-4,6.6666761040686149,9.6965817314772498E-4,10.000009785095704,"<0.000",0,4.3454428048405702E-4,33.333457211651542,"<0.000",0,"<0.000",0,1.3894414684065799E-4,1.1370828417378396,"<0.000",-15.666695122917069,7.9810872309767893E-3,31.666722620526116,4.8322611195781599E-3,26.66671991348187,"<0.005",0,5.2412452109864303E-2,66.668639083747451,7.2982138130533905E-4,5.0000062833228762,"<0.001",0,"<0.000",0,9.6520668520536904E-4,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.46941697121966E-2,133.33489596847213,1.70041480674659E-3,10.000015348195801,2.4436732912845501E-4,13.333364923794592,1.8008056509426799E-3,66.666905681288654,3.6741554099372801E-4,16.666703770557486,"<0.000",0,"<0.000",0]
[null,false,"006SMPL.d","2019-10-22T11:30:24","Sample","deionized (DI) Water rinse bottle",false,false,false,0.12126439354938701,66.667574147403656,0.248757448807266,33.333498140173681,916.64229619299294,60018.25431399166,0.37037089467047901,6.6666761040686149,764.71631774310197,390.00532204898218,80.283283174379505,21563.190687866794,13.914889587384,206.66839291649475,0.16178518911235401,353.33885302138054,4.3458383528745002E-2,6.0000034968058271,6.6681535428523E-3,46.66680584351014,0.11798143931044699,1226.7121450099523,9.8959547131200795E-2,380.00466098381094,0.77726601844274001,6134.4438016662853,4.6850677163350601E-3,96.667002191141208,4.9308985873061199E-2,323.33676174277917,2.4137042293099101E-2,233.33655322598659,2.9071758934479799E-2,66.668543716313678,6.0965391767518603E-3,10.000011771917167,1.8684600700000598E-2,41.666764194765229,1.3544648728298699E-3,3.3333380520343074,"<0.008",0,0.21443264885548899,10.000023692845945,8.64060005923626E-2,183.33607924487373,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,8.6908666503242098E-5,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",-0.36333373938996977,"<0.000",-6.3333411067723056,3.3604555992640798E-3,13.333348234494304,3.0201592844015602E-4,1.666667833924276,"<0.005",0,"<0.001",0,1.4596431686171099E-3,10.000015348195801,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,1.0268274006618699E-4,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,9.1839056366564403E-3,83.334759747020499,1.13361176307704E-3,6.6666880249973923,1.2218356260647499E-4,6.6666768987972,9.9044282679816901E-4,36.666787713764293,"<0.000",0,"<0.000",0,"<0.000",0]
[null,false,"007SMPL.d","2019-10-22T11:31:53","Sample","deionized (DI) Water rinse bottle",false,false,false,6.0632196774693302E-2,33.333787073701828,0.24875692096514601,33.333427409329602,941.92623745106903,61673.750708593667,0.18518544733523901,3.3333380520343074,790.86095591369303,403.33908751598364,72.579325972354098,19493.994067723423,12.568285581023201,186.66819930074317,0.15873229892987201,346.67134086284028,8.2088076708239796E-2,11.333342552185059,8.5733547723141207E-3,60.00028043984949,0.123431598783213,1283.3801841232153,0.105904117304191,406.67181044809308,0.85755620527854404,6768.1208533870367,4.4427350071738298E-3,91.666951403017563,4.5750576332102499E-2,300.00299006696059,1.6378616443893899E-2,158.33464022728742,"<0.000",0,5.0804483449422297E-3,8.333341553807136,6.7264537715979403E-3,15.000029578804529,"<0.000",0,2.5358568650658001E-2,3.333335667848552,0.35738802075704401,16.666718870400604,6.2840685906678906E-2,133.33524167541955,"<0.000",0,"<0.000",0,"<0.000",0,1.6160965057211899E-4,1.666667833924276,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",-0.45416743952779809,"<0.000",-10.000017017125829,4.2005719527371602E-3,16.66669502854305,"<0.000",0,"<0.005",0,0.13102996817149801,166.67011951414517,1.2163686981712799E-3,8.3333419511714286,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,1.0268289919551499E-4,6.6666872302688072,"<0.000",0,"<0.000",0,4.2078875317697298E-5,3.3333384493986,2.2438809648298001E-4,6.6666872302688072,2.0204603917815499E-2,183.33657595025795,2.83402656982363E-3,16.666703373193194,1.8327554782521699E-4,10.000026474395993,1.6207255624202299E-3,60.00023275613438,3.6741624178168698E-4,16.666735559700893,"<0.000",0,"<0.000",0]
[null,false,"008SMPL.d","2019-10-22T11:33:21","Sample","Reagent blank1",false,false,false,0.18189823826769599,100.00226721172633,0.49751426894863698,66.666912039117335,853.05908905937804,55855.06752706963,2.5926047397982601,46.666885316368656,810.46765185634001,413.3385024467334,76.679407533940406,20595.230054252024,158.684039716055,2356.8261367555538,10.3300517386221,22560.832997150723,6.7601934282782605E-2,9.33333937327067,7.0016461611079203E-2,490.00740592667353,0.14202754403919199,1476.7315453777976,0.148440276026788,570.01065994286546,1.0645877313251699,8402.0830125082211,5.3312894809754603E-3,110.00049586096661,7.7776489080443503E-2,510.00842288123931,0.21034736479616101,2033.460789673396,1.8026566805453299,4133.9258482191281,6.0965377232253797E-3,10.000009387731412,6.7264528806488498E-3,15.000027591983065,2.70893362080928E-3,6.6666856408116368,1.2679284325329001E-2,1.666667833924276,11.0791944576057,516.67601769155237,2.3565086207546201E-2,50.000352784999144,0.24196540553264401,703.34891224877663,7.8186268122329503E-4,20.000047385691889,6.5391075912794704E-3,60.000189840790775,4.0726140395498399E-2,420.00553776812257,"<0.000",0,3.2222914237855602,247178.75282716623,"<0.000",0,2.4036156892928702E-3,66.667423148966847,1.1482109047059299E-3,9.3966600833835869,"<0.000",-65.333713034783088,3.02442303871262E-2,120.0006498893005,1.14766362433735E-2,63.33354855576831,"<0.005",0,0.31447858633639603,400.0167618198962,1.2163686981712799E-3,8.3333419511714286,0.10682198160973801,153.33425929147612,1.66504985089375E-3,63.333564450339999,9.6520530447024504E-4,6.6666768987972,"<0.000",0,1.0268289919551499E-4,6.6666872302688072,"<0.000",0,"<0.000",0,7.5742211332552897E-4,60.000278850392327,2.2438774874516601E-4,6.6666768987972,0.27737266729454801,2516.8795830307313,"<0.000",0,1.2218356260647499E-4,6.6666768987972,1.37763733012802E-2,510.00960543737398,1.54314727993116E-3,70.000246912237301,5.7182482448983598E-5,3.3333384493986,3.2257864996790601E-4,26.666752894718154]
[null,false,"009SMPL.d","2019-10-22T11:34:53","Sample","Reagent blank2",false,false,false,0.24252878709877301,133.33514829480731,0.323384099857709,43.333469380933003,894.06752353457796,58540.143984454713,0.55555682767318704,10.00002289811736,830.07604321996803,423.33878204218377,71.461500133418696,19193.758567034769,19.525746427330301,290.00263417756605,11.176736033993601,24409.991498241947,8.2088065195706494E-2,11.333340962727865,5.3345735264113098E-2,373.33799441648529,0.14106560328565701,1466.7297653347784,0.12066154498862899,463.34033275633556,0.98431584704152697,7768.5504106604894,4.2811815881748597E-3,88.333619708811952,6.2526010946172494E-2,410.00555063279148,6.2067102619810702E-2,600.01236349371891,0.130823123136733,300.0089235605804,3.04826813484945E-3,5.0000035017728282,2.0926741978570099E-2,46.666751801966349,2.70893362080928E-3,6.6666856408116368,"<0.008",0,10.721791474380799,500.008601049454,2.3565086207546201E-2,50.000352784999144,4.5869300791244599E-3,13.333361744880252,"<0.000",0,7.2656623981345002E-4,6.6666761040686149,1.42217170644061E-2,146.66746874649121,"<0.000",0,3.3530280809702102,257207.43106314391,"<0.000",0,4.0060008386649402E-4,11.111166990466836,"<0.000",0.42665531992931394,"<0.000",-60.66694288452053,2.6883732925468199E-2,106.6671355565319,"<0.000",0,"<0.005",0,0.28827042344285903,366.67997861931673,9.7309544574473799E-4,6.6666768987972,2.6705405710110101E-2,38.333436076838233,1.75268034771719E-4,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,8.4555863367139302E-5,3.3333384493986,5.4702602120175296E-4,43.333450704811248,4.4877583185361502E-4,13.333363731701716,0.16347979946803001,1483.4157003729076,4.5344657687297601E-3,26.666862169898618,1.8327534390971299E-4,10.000015348195801,4.6821161338892799E-3,173.33474854626135,2.20449426210491E-4,10.000026871760285,5.7182482448983598E-5,3.3333384493986,"<0.000",0]
[null,false,"010SMPL.d","2019-10-22T11:36:21","Sample","Reagent blank3",false,false,false,6.0632196774693302E-2,33.333787073701828,0.22388117489825801,30.000077436366514,879.10539880315798,57560.480913115338,1.1111131696789001,20.000037054220282,692.81944788345902,353.33791842056434,73.697969889211606,19794.449294979484,22.2189933943976,330.00359997147916,11.045254460074201,24122.835740802151,0.101402923297987,14.000012079874674,4.9535300554820402E-2,346.67081952088841,0.14042435999869901,1460.06243755327,0.104167910189246,400.00477512670403,0.84487574279893196,6668.0423955430124,5.0081782519990599E-3,103.33374187349665,5.1342411786221002E-2,336.67066708201168,5.2584866535413499E-2,508.34610803702225,4.3607025003576799E-2,100.00140890482237,6.0965391767518603E-3,10.000011771917167,1.2705516366220301E-2,28.333372647563028,2.7089297456597398E-3,6.6666761040686149,"<0.008",0,10.5073481575243,490.0080798565138,2.3565086207546201E-2,50.000352784999144,6.8804033207392297E-3,20.000066459177933,"<0.000",0,1.4531335189867599E-3,13.333361744880252,4.6866891524939402E-3,48.333462948598516,"<0.000",0,3.38288022806239,259497.35941443802,"<0.000",0,"<0.000",0,"<0.000",-3.1791863698881913,"<0.000",-65.000427549077671,2.94041358661935E-2,116.66738972066757,1.81209657872714E-3,10.000012566645752,"<0.005",0,0.104824129484287,133.33629270401258,4.8654772287236899E-4,3.3333384493986,1.1611022972144E-3,1.666667833924276,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,4.2277984098679298E-4,16.666712909936216,7.1534191875116204E-4,56.666835894184771,5.6096969285166702E-4,16.666701783736023,0.12674243314283701,1150.0608383381057,3.40083339731848E-3,20.000052948791986,6.1091781303237698E-5,3.3333384493986,1.2605661796616899E-3,46.666916310783478,2.9393272011804598E-4,13.333376049994785,5.7182482448983598E-5,3.3333384493986,"<0.000",0]
[null,false,"011SMPL.d","2019-10-22T11:37:49","Sample","Reagent blank4",false,false,false,0.12126604149300201,66.668480138024492,0.14925400786731599,20.000037054220282,874.45601744644,57256.057089526592,3.7037159226557002,66.66688660780261,934.656091862023,476.67460684963157,69.597199021227894,18693.028168472778,17.9547326415219,266.66943470632879,11.2836984669378,24643.5974517922,8.6916773965010005E-2,12.000007947285971,7.0969137514523897E-2,496.67467012989511,0.15805845499967999,1643.412959723478,8.7674589969927594E-2,336.67042548452179,0.79839900551825205,6301.2324152168922,4.5235156839608702E-3,93.333699107164648,0.12607027743706101,826.68816924063105,6.1205026301673397E-2,591.67853724238,4.3607433935672103E-2,100.00234668458779,4.0643618737120297E-3,6.6666784882543704,1.3452904692158801E-2,30.000052799780377,1.3544648728298699E-3,3.3333380520343074,"<0.008",0,11.508056006179,536.67589023695552,1.57100574716975E-2,33.333568523332765,5.73367162116337E-3,16.666728407143626,"<0.000",0,1.4531335189867599E-3,13.333361744880252,2.4241468777711499E-3,25.000039363900232,"<0.000",0,3.3855352667351499,259701.02477598647,8.9914587962204703E-4,5.0000062833228762,"<0.000",0,"<0.000",-0.39250573447585779,"<0.000",-38.000136762855391,1.5962203705539701E-2,63.333561271425658,1.20806472184682E-3,6.6666768987972,"<0.005",0,0.55033733240483296,700.0290868189478,1.21636904617679E-3,8.3333443353571841,1.1611030446504401E-2,16.666689068078664,"<0.000",0,"<0.000",0,"<0.000",0,5.1341370033093598E-5,3.3333384493986,"<0.000",0,8.4555863367139302E-5,3.3333384493986,5.8910543826743605E-4,46.666832069553458,4.4877619296596099E-4,13.333374460537614,0.11388350694990899,1033.3789420634769,5.6680493558219696E-4,3.3333384493986,6.1091781303237698E-5,3.3333384493986,2.3410489648522902E-3,86.667037308210567,7.3483130098369296E-4,33.333417475222291,"<0.000",0,"<0.000",0]
[null,false,"012CALS.d","2019-10-22T11:39:17","CalStd","Reagent blank5",false,true,false,0.12126604149300201,66.668480138024492,0.27363290426444797,36.666809171436093,894.88440550585904,58593.630311782705,2.9629733385861599,53.333520094550842,882.36519489784303,450.00624939790004,70.653812937119497,18976.822832205409,10.3239341552291,153.33437373239238,12.014261329864,26239.146744422917,6.7601934282782605E-2,9.33333937327067,6.0966693711886601E-2,426.67296727488991,0.13240900235471201,1376.7227476330188,8.07300529106285E-2,310.00340317681326,0.76460427741173198,6034.5130997663191,3.5541880057390601E-3,73.333561122414054,0.13420393455027199,880.02348542178436,8.0170527728944299E-2,775.02099815803012,4.3607433935672103E-2,100.00234668458779,1.1176989217475E-2,18.333356107274351,1.2705523850192701E-2,28.33338933686332,4.0633946184896099E-3,10.000014156102923,2.53585867885067E-2,3.3333380520343074,10.5073413409102,490.00776196507968,5.4985656234444402E-2,116.66845542693164,2.2934633991517198E-3,6.6666761040686149,"<0.000",0,2.1796987194403501E-3,20.000028312205846,2.5857569521811402E-3,26.666711568831726,"<0.000",0,3.5378052101102901,271381.49986235052,"<0.000",0,1.2018039764904399E-3,33.333604286120433,"<0.000",0.84999732181437038,"<0.000",-36.666767766076802,1.68023038348315E-2,66.666843692459011,6.0403185688031203E-4,3.333335667848552,"<0.005",0,0.157235806858709,200.00394632427773,2.4327396583783802E-3,16.666699399550268,2.6705397958921601E-2,38.33342495063804,1.7526828549426499E-4,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,4.2278011314216999E-4,16.666723638772115,6.3118482021982502E-4,50.000210652745608,1.1219387437258301E-4,3.3333384493986,0.106535821383552,966.70604323435418,2.26722352615409E-3,13.333376049994785,"<0.000",0,2.1609698745545902E-3,80.000401337935514,1.4696636005902299E-4,6.6666880249973923,5.7182482448983598E-5,3.3333384493986,"<0.000",0]
[null,false,"013CALS.d","2019-10-22T11:41:11","CalStd","Standard NMS-59",false,false,true,1069.6701833038901,219068.38868562257,39.6507501332621,1820.0936109807053,2782.3138794364199,56948.436651402175,2674.17545825974,13605.112915757602,2730.9694589585702,483.34079022197733,134.70555697659199,19544.153268877115,504.55870153476701,4530.5800716492522,16.789927684734,24530.12821520177,0.19038959718726101,19.333348592122395,2.65844881713203,14929.573345240357,2.6505854965681799,24116.280565118581,2.6427221760043298,6377.819322280794,1.1164559361340001,6467.8993399220053,1.3134977674383099,22771.21141198459,7.78181228874162E-2,433.33927780380935,26.1126889374892,216656.23042858162,0.287673900601852,533.34647417097858,2.5955422526212102,3228.6357144221306,0.14549895243338501,238.33541514969741,12.8990780574675,22557.812163862403,51.439045818593101,3847.0917192038141,17.4262574176018,570.00997329736776,127.418116461905,261902.68326283534,25.404990086742401,58411.161238980865,4.79841554613394E-4,10.000014156102923,3.4656337694234699E-3,26.666723489760503,0.249496126979006,2281.8199326323479,"<0.000",0,3.7056140145440501,272969.55322410335,6.1042465573356004E-4,3.3333384493986,4.9709115294545496,138381.80105314372,4.4590846567736202E-3,37.683802017315116,1.2427278823636401,15434.371478244093,1.97008070741496E-2,85.000330929949925,1.62784136938172E-3,10.00001216928146,"<0.005",0,0.69753253640384905,1033.427484087267,9.4408506817000498E-2,786.68560127389037,12.427278823636399,22156.32291203256,3.5611721714303303E-4,16.666701783736023,"<0.000",0,3.21431645611022E-3,63.333511998253378,"<0.000",0,"<0.000",0,8.9225292599640603E-4,40.000134507813037,6.7643428787390295E-4,60.000193019705115,2.01443580462859E-4,6.6666872302688072,0.155824268951961,1566.7619419615903,1.0284529986409299E-3,6.6666880249973923,1.1781652031980401,70482.50546014549,11.879440704499499,479627.77931132243,1.33221988234556,65555.171549193488,1.23420020498147,75110.42531305911,0.24854557647272699,21211.177255543411]
[null,false,"014CALS.d","2019-10-22T11:42:59","CalStd","Reagent blank6",false,false,false,0.307694064192999,66.668480138024492,0.68638239283458902,33.333434561886868,2680.6799142601199,58048.667754799106,6.1929366120365898,33.333467940487445,2029.4634630676201,380.00512510530461,121.672341527315,18676.47136471729,30.176408589916299,286.66935811436139,16.789927684734,25952.03574609703,9.9500223474487406E-2,10.666673024495443,6.9868301631451593E-2,413.338909347769,0.17322802382862401,1656.7466488965356,0.143536337393778,363.33805814378371,0.98425162482287298,5967.7614910151897,4.3266910605888704E-3,78.333597605423165,0.114891176098735,666.68010989804895,0.101431742746857,875.02707211982124,0.10398349272063601,200.00563114894103,1.42378699552532E-2,18.333356504638644,2.4627566682771499E-2,41.666756644843666,9.2478833695864796E-3,16.666709333657582,2.1672759971715298E-2,1.666667833924276,12.813815158795199,430.00567749134188,0.198492509759465,416.68061154615253,2.5610308545776801E-2,60.000208914276833,"<0.000",0,1.2812771062921599E-3,10.000023692845945,6.2975706328940104E-3,58.150802362817487,1.7785453828287999E-4,1.666667833924276,3.7056140145440501,273678.45946353982,"<0.000",0,4.3394333365820902E-2,1211.1623822826348,"<0.000",-1.72584015102426,"<0.000",-19.666795993842747,1.46410106240838E-2,63.333583523826057,5.4120843543441297E-4,3.3333384493986,"<0.004",0,0.42639063093948998,633.35849270280812,1.1969728230044E-3,10.000017732381556,3.7295907161409101E-3,6.666674117247152,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,2.2248524766043599E-4,10.000025282303115,5.9971772489967104E-4,53.333529631293878,"<0.000",0,7.6055195774686807E-2,766.69602518345971,5.1289364821306401E-4,3.3333384493986,1.55609977346781E-3,93.3338087797094,9.7993489437569106E-3,396.67238280168596,3.8717150232133402E-2,1910.120821146357,6.4636588465412995E-2,3943.8413922412983,3.5062098544097502E-4,30.000123927988749]
[null,false,"015CALS.d","2019-10-22T11:44:47","CalStd","Reagent blank7",true,false,false,1.3828302817893601,300.00589564455817,0.82259918414986699,40.000155170756251,2685.2467275221502,58222.515408458574,8.0404548726074605,43.333518654105283,2115.7466584167901,396.671823312762,123.252319353678,18943.382686601191,26.282744765781899,250.00224694606291,16.789927684734,25985.489521640196,9.94148716650603E-2,10.666673024495443,6.2516614457800407E-2,370.00471835328062,0.148132154682169,1416.7257479321099,0.16599502736348601,420.0056903560108,1.2052774769226,7301.5619002631729,6.0836877696260998E-3,110.00049586096661,0.15479759838575699,896.69032985930153,5.7117778899080998E-2,491.67490309743613,"<0.001",0,6.4914283858002004E-3,8.3333391696213805,2.6690250831454099E-2,45.000121717649876,1.8568131827256199E-3,3.3333380520343074,2.17670724238745E-2,1.666667833924276,17.2668370920799,576.6775808233383,8.7830269983792605E-2,183.33607924487373,2.8629725578135702E-3,6.6666761040686149,"<0.000",0,1.2902390816559499E-3,10.000023692845945,1.2734268885151901E-3,11.66668040057003,"<0.000",0,3.7056140145440501,271179.05845911714,3.0722727531929002E-4,1.666667833924276,9.2407561654441004E-3,255.55971504922732,"<0.000",0.98624575973186057,"<0.000",-33.666781162220509,1.43871025617795E-2,61.666866019365195,5.4619618528674904E-4,3.333335667848552,"<0.005",0,0.11324010669194,166.67011951414514,2.0133396277001401E-4,1.666667833924276,2.8229745764115602E-3,5.0000062833228762,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,2.24536329734475E-4,10.000046739974914,4.9176136752992203E-4,43.333461433647159,1.01386661742125E-4,3.3333384493986,0.115134670457986,1150.0480273128358,5.1762088233943697E-4,3.3333384493986,5.0478442258686501E-4,30.000122338531582,4.8201519415316998E-3,193.33523144321794,8.0462608015956904E-3,393.33907494933788,2.7844426614840299E-2,1683.4290825809676,"<0.000",0]
[null,false,"016SMPL.d","2019-10-22T11:46:41","Sample","Reagent blank8",false,false,false,0.24348146372642901,333.34058076159465,-0.31469628138390998,23.333375900983174,132.34087066977301,57717.52192640023,-1.4938918886458901,33.33344250917272,104.877528418785,393.33829929423882,3.8733226377968299,18459.463404951417,-1.05986090498706,226.66862785813271,0,24550.163208392809,0.104164980413329,20.666693369547527,2.1293714766725001E-2,470.00711376457747,3.07465054443169E-2,1623.4090435990338,-3.1525885268885799E-2,323.33672359580709,-0.10079983526627501,6367.9362130945565,-1.1515230564916E-3,85.000310267006697,-3.4227136728298302E-2,666.68088555314807,2.2884953345754199E-2,658.34781210901917,5.4446686635820002E-2,100.00234668458779,7.0489959532899996E-3,16.666688273350076,7.5192985825211402E-5,43.333434412875256,5.8662692204684503E-3,13.333380818366296,2.3433044853746698E-2,3.3333380520343074,-1.4356955996298999,510.00806684283316,2.9492124366048301E-3,183.33607924487373,8.9555424329226998E-3,26.666733026503532,"<0.000",0,4.7976054489501098E-4,13.333361744880252,1.33898624267938E-3,23.333384245633315,"<0.000",0,0,265595.39317773684,"<0.000",0,1.42493953597266E-3,288.89396041643039,"<0.000",-1.4679285673112354,"<0.000",-40.333593164877193,4.27266157878688E-3,78.333778803540582,2.9032228075965899E-4,5.0000035017728282,"<0.005",0,0.118002129320479,333.34122449177261,4.1536632159549797E-4,5.0000062833228762,-9.01425683407502E-4,3.3333384493986,7.3200894693594294E-5,3.3333384493986,"<0.000",0,"<0.000",0,4.4355605204687802E-5,3.3333384493986,"<0.000",0,"<0.000",0,-2.8283772652294501E-5,40.000194907185509,2.1314721478836401E-6,3.3333384493986,-1.12092088454458E-2,1016.7079199114723,1.08820477458432E-5,3.3333384493986,-2.7571984652205298E-4,13.3333537975944,-2.0200059836029799E-3,110.00067070125533,-4.4955499546754599E-3,170.00147248305669,-1.24195326400373E-2,913.36238771633725,4.0143418021298198E-5,3.3333384493986]
[null,false,"017SMPL.d","2019-10-22T11:48:10","Sample","QC1-1",false,false,false,193.47529205639501,40645.509738250541,8.7313481476630095,446.6726872324046,71.8594432687687,57476.909950818961,534.93331460431898,2813.5545438495096,214.243832094027,420.00551233680773,581.65424413090602,104166.27633545916,435.93143672398901,4227.1690247381375,0,24984.11990765552,1.04324168613828E-2,11.333342552185059,0.50031823964323496,3203.6258474445472,0.45311000716476502,5530.8625477920586,0.45988192477422901,1523.4005001674022,20.046022447567498,123861.79476071436,0.23400542363395199,4177.1635323192149,0.95817402841591404,6204.4294322863798,5.07564343002461,42526.097548839643,49.404099160137299,91236.252853091268,0.50055946428345899,626.67841931166618,1.23219088706818E-2,63.333550145225473,2.5185330980854501,4357.2265992975244,10.3067051313795,761.6841011988505,-0.68862310781465697,533.34269682552372,24.6585001079258,49771.432590342789,1.4962672104679799E-2,40.000161528584933,"<0.000",0,9.40645522963228E-4,16.666709333657582,1.3680360237351499E-3,23.333372722068834,"<0.000",0,0,261573.02344600824,1.12826465237853E-5,1.666667833924276,0.92861811224980495,25018.379601088043,0.224391282569079,1818.1136181471311,0.246646541256343,2902.9216778836394,-2.6966187559756601E-3,48.333453411855508,"<0.000",0,"<0.005",0,-1.9319581035743199E-2,133.3382636309772,0.479718774309917,3832.1081989192735,2.3797567378469999,4070.5006310765843,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,2.6073484686183499E-4,6.6666872302688072,3.9621506252572603E-4,26.666739384332207,2.14145038739902E-4,60.000227987762877,3.1905363586779898E-4,13.333363731701716,-1.30712173756754E-2,983.36991282882582,2.7021647651161301E-3,20.000064074992178,0.22368362264776701,12851.891084533569,2.3438021852834701,90865.760364744652,0.25984102657326602,12631.702692632338,0.218682335792486,14376.6408176724,4.9000755301144901E-2,4007.1929334522342]
[null,false,"018SMPL.d","2019-10-22T11:49:40","Sample","QC2-1",false,false,false,-0.75666021823349106,133.33514829480731,1.55014358437079E-2,40.000093181926609,582.96398039961798,69551.781190160196,-2.9992407880433301,26.666739384332207,311.59439273812501,446.67323400567125,-4.3124248217012697,17942.436619306431,87.615925594226994,1063.3672682935903,0,25504.824652500578,4.9748845496121596,534.00813802083337,1.8504090504882501E-2,470.00739350903945,2.6588800937474202E-2,1636.7448690025278,7.6299917482125197E-3,430.00609710803491,-1.24538375525829E-2,7068.2162275263327,-1.74585720719811E-3,76.666881690419487,9.7938546619919797E-2,1430.0648285438892,3.7100398333140203E-2,791.68811024525019,0.10687509312948799,200.00469336917558,-1.16406176650329E-3,6.6666737198828594,0.56798150017591098,976.6944231088795,1.9580234219967402E-3,6.6666856408116368,6.08524703366937E-4,1.666667833924276,-2.68791319816859,473.34044704824026,-2.2033149240918599E-2,133.33476285142913,4.8783012043885696,11026.832924880093,"<0.000",0,2.4624686572192398,18499.885792038553,0.65732553977758001,5837.663487230785,"<0.000",0,0,261823.01042144978,0.56169252250319501,2943.5883146983142,2.6130745817911E-2,944.47409303718177,"<0.000",-36.157323766188327,"<0.000",-2612.0166783529025,1.0570269650828901,4433.9168650818347,0.48006782623745398,2831.9037460500435,0.37310194280966003,79.333534240722656,0.168252446082372,400.01479089293161,8.4130913455945098E-4,8.3333447327214767,1.0821678828490699E-2,23.333385437726196,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,0.40751748870843202,17532.851301240156,1.7573553367172501E-5,43.333532164491238,0.39720220550103602,12611.674209857869,9.1721517388611797,89567.480436233993,0.54101904549429503,3367.0313015525844,-9.8145149100753895E-5,23.333402127026488,6.6280077568774097E-3,443.34071452407966,3.39509303925444E-3,540.01032978280887,8.2790426594031893E-2,6458.0317234292861,1.62887511982266E-4,13.333376049994785]
[null,false,"019SMPL.d","2019-10-22T11:51:09","Sample","QC3-1",false,false,false,-0.270415038546232,233.34103946901701,-0.32630139929915902,23.33345378438452,55.834901716241298,57463.493609873054,-2.2831939565882702,30.000088562566713,90.909365010202706,400.00518202773964,-5.9919634294538602,17425.166201657295,11.4119976603869,346.67080044740237,0,25124.31277666023,4.1891614090905999E-2,14.666678428649902,3.8205117380328401E-3,380.00522047273489,3.1111120442590099E-2,1660.0828304874474,2.6083795229040201E-3,413.33912551394423,-2.4412958316584001E-2,6934.8776777930543,-2.3773351029348E-3,65.000199303027998,-3.6202088510536898E-2,666.67972842832808,3.7603204277101399E-2,791.68874602814196,0.35772335251415699,666.69676999318335,2.8584174553579599E-3,11.666680003205736,1.7708023099304801E-3,46.666770080723815,5.1459677990209801E-5,3.3333380520343074,"<0.013",0,323.12010784299298,11073.522377309366,-5.5077159457943198E-2,66.667621831120542,2.6481479303063701E-2,66.666913628574505,3.2445135100290899E-4,6.6666856408116368,3.1131680597503001E-3,33.333409130572143,1.61752725038453E-2,156.30234133285771,3.6659658821285197E-4,3.3333384493986,0,265577.44395925436,"<0.000",0,-1.8563436428617899E-3,200.00270654763867,4.0788660016112901E-4,4.3195826323889861,"<0.000",-37.000097677110169,3.0344704065749298E-4,61.666829461850284,8.48095951229209E-4,8.3333419511714286,"<0.005",0,38.153483598572002,55158.671276837696,2.09827118125033E-4,3.333335667848552,5.8245804257145897E-3,15.000024015704433,5.1244176683480495E-4,23.333408882219459,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,7.5707878215819403E-3,340.00447084499689,-6.6879015712861603E-5,36.666774203378345,2.0726452566949501E-3,70.000242143865776,0.62105767522797095,7201.7193178720127,0.13902041081187999,880.02549449564776,1.2518687576999001E-4,36.666789303221471,-8.3179478410282896E-4,156.66765933234009,-4.0775566738809601E-3,190.00137681760327,-2.3790987713739099E-2,240.00219623244507,4.0146131139437903E-5,3.3333384493986]
[null,false,"020SMPL.d","2019-10-22T11:52:43","Sample","acid rinse",false,false,false,11.0936330682948,100.00136122110548,71.540760007783803,130.00068152943231,59296.648170891,49650.671460668178,611.38088345815402,123.33405390380739,78733.251613037806,560.01007517163828,3636.0121946528202,21346.114063548612,628.210705130919,230.0022090971141,7.1054273576010003E-15,960.02912998156046,5.4199032356024901,13.33334763844808,2.40770869063308,236.66843697426066,15.7970473489094,1926.7723305010943,16.112348281662001,426.67243321728057,178.76573792929301,9569.4411827250588,0.59925845211397699,83.33360229928887,11.5735567530807,456.67294139654024,10.7234782460794,558.3455369500067,2839.2860979117499,29658.618826984293,2.0824037067515899,13.333349823951474,13.5869659239501,105.00044984121446,2.18999644449169,16.666728407143626,5.4981460473739796,1.666667833924276,18735.284703505899,2310.1562664386188,29.248496904117498,200.00284562514628,36.6075698226336,263.33596790826368,0.32429119710526,20.000066459177933,1.3878455623542301,30.000082204738025,3.5116135077276498,81.666940425828983,"<0.046",0,0,606.67852133494841,0.96098617035791101,11.666691526770222,13.639614478065401,844.46875022571646,0.212405365421205,3.9917102755357803,"<0.036",-593.68499745401243,104.620730721415,1003.3639122541163,5.1264898274214401,70.00019128123634,"<2.030",0,1305.9919950597,4300.6897933935325,0.35977567906087699,6.6666745146114446,7.9888555680972004,31.666791364548732,5.1275011435927196,533.34333260839185,1.0690037049501999,20.000060896077837,1.0656612935109799,46.666855911411005,"<0.006",0,"<0.017",0,1.37143848726031,136.66751975813227,3.3325515712291899E-2,6.6666872302688072,0.13585499800329001,10.000015348195801,836.45013263067699,18694.426890834777,96.023031749698802,1383.4002479404173,0.124847049463829,16.666724036136408,2.4097680688166299,216.66852310297108,3.4971077075436598,383.33860461902714,1.15509907939005,160.00111977257652,8.7871261362302E-2,16.666714896757679]
[null,false,"021SMPL.d","2019-10-22T11:54:13","Sample","base rinse",false,false,false,"<0.067",0,29.197767078346899,100.00054846206483,34955.5373048213,55908.57501325148,208.64778498068901,80.000310341512503,28250.4165367852,390.00538562726894,1771.0389914147399,19944.45591047432,265.30601443026097,190.00147218503346,0,1780.0979349002548,2.2530523988704898,17.333348910013836,0.395631007041648,186.66814525919935,2.54884848487491,1780.0917805220915,1.8715215005998,356.67114997896829,23.395808556816998,10336.461107502755,7.8761723310717402E-2,106.66712284087454,0.77807042734186904,376.67095894608457,1.6884696014107601,1050.0372485576452,56.0645027588345,7535.1007238425482,0.30775249888131401,28.333371058105858,0.35235965006294401,45.000101849435246,2.4446168795665399E-2,3.3333380520343074,1.8235792339638801,10.000021706024482,278.12525733474001,700.01551538682691,1.0316642867376999,166.66881417239537,0.117248572441203,20.000056922434911,"<0.001",0,0.79448229786027402,443.33984985937894,9.1876739431023299E-2,61.666862840450868,"<0.001",0,0,19751.36201036927,0.87285255049099597,345.00382227196559,1.79148851608848E-3,22.222333980933673,9.6493933101501297E-2,59.077244984154291,"<0.001",-427.6767357834666,2.36669296464992,743.35015778715172,0.100691523684122,45.000116949278357,"<0.062",0,14.8131320132146,1600.1203258859193,5.3271510532154299E-3,3.333335667848552,0.12637185377709001,16.66669979691456,6.8903083515204798E-3,23.333408882219459,5.4725391207253501E-3,3.3333384493986,"<0.001",0,5.3680320490913404E-3,30.000086973109536,"<0.001",0,9.6370517809485606E-2,313.3366538087032,5.7405850055667297E-3,40.000134905177333,5.4666264729550697E-3,13.333364129066007,83.523980578749303,60849.944393956415,7.7656780888471097E-2,36.666806787250344,4.1155474570471404E-3,20.000053346156278,2.0528060420324499,6011.1512369242628,5.3743020254950501E-2,220.00220288829701,1.30922996893994,5887.7995948821363,4.3184581648414997E-3,26.66676243146118]
[null,false,"022SMPL.d","2019-10-22T11:55:41","Sample","nitric",false,false,false,"<0.359",0,784.73199675303294,490.00763480850605,203326.372798581,57299.595500336167,666.97231164975699,46.666840016839302,161432.284773269,393.33815942200789,10373.5765550786,20695.390273775938,1612.8242851837299,200.00165005522493,-1.4210854715202001E-14,333.33734651400624,15.6829822765829,10.666673819224039,2.5015415230230298,63.333586702740398,61.633061939734802,1843.4300465867411,64.366271462881301,406.67256702970616,827.45012976282203,10403.162107935694,2.7475282239642,88.333643550669493,45.855457625475402,413.33879490685268,54.766540934598297,650.01702172098851,140.27421209421601,333.34178080180294,14.991798123456499,21.666695748765832,9.5624676440297698,16.666683107614272,5.8555259600073999,10.000042766331989,98.629465187157905,6.6666737198828594,7498.8441899683003,206.66839927432341,43.922275059922498,66.667858262883442,6.2659257629349101,10.000023692845945,0.244122196508332,3.3333380520343074,8.3773158238667396,40.000161528584933,2.9252080784561398,15.000020042061507,"<0.210",0,0,133.33404978115286,8.7475824515193796,23.333371927340252,"<0.074",0,"<0.242",-2.9168580335835026E-2,"<0.165",-30.00005963444621,23.7106777292727,50.000099390743685,14.4408112360901,43.333455870547056,"<9.235",0,1335.73507663436,966.71742374811777,1.6377171821038301,6.666674117247152,7.6524507030648303,6.666674117247152,"<0.044",0,"<0.243",0,"<0.104",0,"<0.027",0,"<0.077",0,0.76088981745585305,16.666724036136408,"<0.023",0,1.23712115887103,20.000061690806422,315.502055945341,1550.0846891372457,10.527048305285399,33.333469132580326,0.341710298293947,10.000026474395993,4.5587977819589796,90.000407149388295,1.7948292336820899,43.333487659690462,0.757095084822345,23.333380271990393,"<0.024",0]
[null,false,"023SMPL.d","2019-10-22T11:57:10","Sample","Reagent blank9",false,false,false,-1.2307462606517601,33.333787073701828,-7.6220326556934698E-2,36.666790097950049,-68.992866237710103,57309.525513209024,-0.69399984729588604,40.00012258688426,506.440951394781,496.67421236623005,-0.28660989168790302,19093.546976253197,7.9560095942767403E-2,253.33589812115983,0,26252.524022869406,1.75166388799218E-2,12.666673342386881,7.2611136763737101E-3,416.67286569864251,2.71314684300808E-2,1690.0828300404128,1.3128716909602701E-2,456.67386963952777,0.72434564216954,11770.834613684594,-1.5952071243863499E-3,81.66692771017162,1.4586506130675801E-2,986.69546564377151,2.5728921855506699E-2,716.68318125145788,0.50175602909621098,966.71688333265968,-2.3440139736101599E-5,8.3333439379928915,-8.9493081450239403E-3,30.000051607687499,1.84830101125032E-3,6.6666856408116368,-3.55644860698286E-5,1.666667833924276,-1.1136270764227301,540.01046965503986,-5.58807760665333E-2,66.667858262883442,2.8954106089723301E-6,6.6666856408116368,"<0.000",0,3.0206012065947399E-3,33.333428204058187,1.8724770371517801E-4,13.333355387051574,"<0.000",0,0,269654.00639850885,2.9044572289762999E-2,158.33420517542652,-6.00841933197197E-3,88.889860444620453,6.9639852830275599E-5,1.562078553208553,"<0.000",-37.333468914029957,8.6345116694569598E-4,65.000227913257064,8.2701738826260196E-4,8.3333419511714286,"<0.005",0,0.68394796714292305,1166.7263291789516,1.2159751348364499E-3,11.666680797934321,2.8549055151196301E-3,10.00001216928146,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,3.0234454798714201E-4,23.333429545162673,-1.1134348752471899E-4,33.333477477230467,1.02533780209501E-4,6.6666872302688072,1.0229588955966,11304.16666502133,2.6056740178026002E-3,20.000062088170719,-4.4838020208632197E-4,3.3333384493986,-1.7278489935235899E-3,123.33406264582182,-7.0862313579729804E-3,46.666826506453354,-2.6181032653687101E-2,100.00055164097917,"<0.000",0]
[null,false,"024SMPL.d","2019-10-22T11:58:42","CalStd","Blank control1",true,false,false,0.15036268865307301,33.333787073701828,0.60376099811740302,30.000059952337647,2518.7331151937201,55804.932436670875,4.8422006866449996,26.666731437046355,1966.12378089412,376.67102252437144,120.936205129892,18993.387773981915,31.208131511031802,303.33646958068306,16.789927684734,26553.059500161919,0.140345387072493,15.333346048990885,5.8279876019621202E-2,350.00422909849539,0.18544622907031399,1793.424808480966,0.18172762913964,463.3399767179294,1.5511288539275701,9436.0325697171029,4.6055492089780601E-3,83.333605478203211,9.9825361812703906E-2,576.67674158995226,0.107151520405879,916.69451855185298,0.40350090951414402,766.69529403313425,3.9459117120023597E-3,5.0000035017728282,2.2107971129237201E-2,36.666742414234939,5.6820531075426799E-3,10.000023692845945,4.4557379315151302E-2,3.3333380520343074,18.552606316371101,603.34522619825759,4.9532263687237503E-2,100.00094796222579,7.4251240425254397E-3,16.66669979691456,"<0.000",0,8.9832144026047601E-3,66.666969259575467,1.5300607729219701E-3,13.33335538705157,"<0.000",0,3.7056140145440501,255390.90765271228,8.4817288355203101E-3,43.333418518303553,4.26604573543061E-4,11.111166990466836,"<0.000",-0.39250439178191332,"<0.000",-38.00010719895203,1.5689378859270301E-2,63.333511998253385,1.73988722856479E-3,10.000017732381556,"<0.005",0,0.76956566717627695,1066.7249282214154,6.4134136960588401E-4,5.0000062833228762,9.9916376125462805E-3,16.666697015364512,7.6125734411597895E-5,3.3333384493986,"<0.000",0,"<0.000",0,4.6127892775240001E-5,3.3333384493986,"<0.000",0,3.9736076352837798E-4,16.666702181100316,5.6232806665063701E-4,46.666799485681473,"<0.000",0,0.64679774706176996,6084.5384050430002,2.19848512955208E-3,13.333385586737807,1.19108777671899E-4,6.6666880249973923,5.5593519418243899E-3,210.00214263793615,2.17209290279733E-4,10.000015348195801,1.0537742244832299E-3,60.000297129149772,8.3494943377644405E-5,6.6666880249973923]
[null,false,"025SMPL.d","2019-10-22T12:00:10","Sample","19286427-001",false,false,false,63.4912044107095,14339.180313181438,13.8518457450031,730.01583283122636,45.0398522725541,57730.974074637285,4.6865062831858504,53.333472410835732,10018.263011302601,2333.4906166768478,319.40276756943803,70286.531112042023,195.89632945595,2243.4759187180412,0,26986.928308800929,0.154232974978737,32.666704813639321,6.65653229020398E-2,760.01705308729822,4.8473380143762998E-2,2290.1504683968851,7.8188186051413894E-2,670.01321429987922,3.0380030934308202,28189.524962767278,1.87939985738687E-3,118.33382748066644,3.3661305492135599E-2,776.68524538449583,0.36544691964916598,4067.1507155010017,361.19103426047297,690290.47766877792,0.62262261292667698,796.68616061380271,1.4991230801282399E-2,61.666837409136143,1.8864234753814099E-3,13.333361744880252,"<0.013",0,832.71821376140394,27645.161534669576,4.9907812796367301E-2,200.00357677547186,1.63736512611301,3673.7303122818921,1.1111654022263299E-2,223.33516741789629,0.11544028701620999,916.69315243721269,1.7335742135113501,14976.31328516843,1.92135421148129E-4,1.666667833924276,0,252109.27358702908,-1.21146292861823E-3,36.666736056406258,"<0.000",0,7.4686955960942905E-4,5.4419996185800201,"<0.000",-69.335280418366963,1.40072946775499E-2,118.33658042048542,9.5788569559229306E-2,553.34298898761983,"<0.005",0,-0.184888111688673,800.03055135477314,1.0911638958172E-3,13.333349029222893,98.510793639238301,162227.36737429051,9.4273514814212191,407497.09028198686,4.7161811678264696E-3,36.666781355935619,2.43622575148809E-2,443.34016775081301,7.0152662065572999E-4,53.333507378893479,2.5699656656711402E-3,63.333602597312087,9.7466056210821408E-3,420.00637700150838,1.70076972161847E-4,60.000237524505884,8.7244616198562796E-4,26.666738192239333,2.4412296507629798,28676.336802261649,8.9114090888553696E-2,546.67711079105072,2.2337520332230401E-3,130.00078961251987,4.6201498112512902E-2,1930.1201596341507,7.6172898091019401E-5,13.3333537975944,-1.64200149031535E-4,50.000153432287476,5.07605638477649E-3,406.67294849942704]
[null,false,"026SMPL.d","2019-10-22T12:01:38","Sample","19286427-002",false,false,false,61.352194781947802,14038.951170321705,17.574010773278399,930.02501055561004,51.872284134678502,58643.889100096283,7.5023627215957296,70.000259627894664,10961.6361811192,2550.1841864459088,282.77490472845398,65285.045811448021,212.27721886086201,2436.8377322429756,0,27340.816649881755,6.1196400137132997E-2,22.666688283284504,4.95877514295553E-2,666.68082197486126,0.185779596184666,3693.7221739631568,0.14874406432760201,866.68921132848845,3.40878802228189,31028.187745942894,1.8151605621368201E-2,423.33874389521162,0.821558528975186,5470.852579610756,0.52711690611765605,5575.8936538871203,309.15431109251602,604265.86637097795,0.62998222107972701,825.02020560174878,-3.5135060528698001E-3,31.666715070604557,3.5429017698576102E-3,16.6667458911725,2.05641178321035E-2,5.0000058859585836,891.66042347063399,30372.928421106015,8.7307212905958595E-2,283.33775041013894,1.6018115877475001,3703.7451637226568,1.22205121528662E-2,253.33601256207604,0.12600334690963499,1026.6985930994363,1.75218970037525,15655.435150830419,"<0.000",0,0,261447.90626037261,-4.3391155980295401E-3,21.666696146130121,-9.8832052049348892E-6,11.111166990466836,2.0227364529245401E-3,15.970847241915802,"<0.000",-28.333389893173329,-3.5899942458092401E-3,50.000096211829344,8.6922809414604493E-2,521.67538183417844,"<0.005",0,-0.37021292173272302,566.6868972286137,1.0292830465577199E-3,13.333367705344642,95.949891551133305,163863.53271449308,9.0528620204388393,405804.82468220359,4.9611619452437699E-3,40.000190536178295,2.1195839193823101E-2,400.00600854546815,3.14346610317986E-4,26.666749715803814,1.69558900981027E-3,43.333504348990751,9.6948085929422605E-3,433.34018061548204,-3.6614944370728198E-4,16.666702181100316,1.4722473922604701E-3,46.666840811567887,1.5439548041378801,21097.558556690881,8.9074055350800305E-2,566.67856007796684,2.2660612487443799E-3,136.66756108401873,5.4353034579836798E-2,2316.8393107726279,3.4859787184530799E-4,26.666771968204202,-4.8190821936929998E-4,33.33341787258658,5.4626945185020804E-3,453.3403710523192]
[null,false,"027SMPL.d","2019-10-22T12:03:06","Sample","19286427-003",false,false,false,63.9827978598424,14339.146235218437,14.894807372310799,776.68466046425726,126.05708626360899,59098.674249569456,7.1607279564380999,66.666906476017246,11542.941906977199,2610.1937477261758,290.69830034052097,65200.953605792391,211.254457091951,2376.8275606111551,0,26780.004483759334,0.119547924283051,28.666704813639324,5.6506435722617897E-2,696.68124114451939,0.11269709702557799,2916.9081189695776,0.21849468517179399,1033.3642932768023,20.511420342716701,136053.32335352874,2.2025354263837801E-3,125.00054491062149,6.7193216554276403E-2,980.02826501883681,0.37287919700609201,4175.5091011017112,314.76402872650903,609485.64187827404,0.58996678583025997,766.68455888800975,5.4843540621438303E-3,46.666814188160288,5.45103830204954E-3,20.000047385691889,-2.27527417351602E-2,1.666667833924276,749.70062867053502,25551.984482527012,5.5203116884592299E-2,216.67059548207956,1.53182158684368,3543.7028483991398,9.3124179053257795E-3,193.3345893025211,9.9419568241586306E-2,826.68766061433655,1.8455785497480699,16571.280677547842,"<0.000",0,0,263657.12497745268,-5.9537891006211504E-3,13.333352605501524,1.63955112920327E-3,55.556041581774025,1.36216563378081E-3,10.713752737138526,"<0.000",-41.333488881585666,1.50778725906179E-3,71.666927859183247,8.1966413715769904E-2,496.67454297332159,"<0.005",0,-0.350260984176753,600.02666860878469,6.0112616434124797E-4,10.000009785095704,94.919960079291599,163474.57113174736,8.8319038850515206,399245.54092577781,4.9195843185712903E-3,40.000130534170118,2.2769759819512399E-2,433.33991358667737,5.7941718854695999E-4,46.666850348310909,1.9400535734525699E-3,50.000147471823084,8.3015429130394395E-3,376.67196348301627,-5.6536905717973801E-5,43.333480507133203,2.08558612491557E-4,6.6666872302688072,0.40212067730648599,10186.733219227688,7.5531487883070098E-2,486.67552222795007,1.7268875591105899E-3,106.66729450224894,4.8294304641010297E-2,2100.1428418336877,9.7505511008169404E-4,56.666831125813268,3.0720756399051801E-4,80.000412066771403,5.8610768704855697E-3,490.00963116671193]
[null,false,"028SMPL.d","2019-10-22T12:04:34","Sample","19286427-004",false,false,false,64.144061974241296,13805.352328837342,14.980462611176099,750.01827393941664,213.408753966067,58630.458110713087,5.7814106197622097,56.666812052327209,11256.149313713,2453.5057719522847,342.32843028786601,70470.094348566854,191.516783773463,2096.7916049044584,0,25718.413516505359,2.32143605796203E-2,17.333347956339519,3.7141705622680599E-2,556.67675460363296,5.9396726726476001E-2,2303.4889386571108,5.1776965498117E-2,580.01087596002901,4.0929856024818303,33498.458593800977,4.6335232445378401E-3,163.33421781657808,5.5869318894424098E-2,880.02506216329732,0.32136625203824598,3592.0436390449427,254.494300480916,475230.9188253638,0.58371725834287802,731.68257576700182,1.0573172453844201E-2,53.33350181579339,9.6955966227617804E-3,26.666769584018443,2.31792555391531E-2,5.0000058859585836,715.04456372796506,23572.489883117527,8.3783727154277607E-2,266.67119463294932,1.57785065140402,3530.3687014624174,9.7890871015834692E-3,196.66799888012807,0.107506402413062,860.02330770060462,1.8680357940846599,16250.982016956079,"<0.000",0,0,255764.20842839964,-1.9668461499897502E-3,33.333391646543269,4.25363232612874E-4,22.222437295653595,1.0505342492072901E-3,7.9254171216674889,"<0.000",-26.000102688867308,-2.9088209875076801E-3,51.666848286983644,9.6711340623068395E-2,566.67642384752992,"<0.005",0,-0.36133835160071798,566.68688133404146,8.5293754379636797E-4,11.66668040057003,97.488844712174398,162872.07765114191,8.9759286034981898,393609.27798893134,3.3809359039061499E-3,26.666756470996791,2.4014098888464602E-2,443.33987529069373,3.6841763659091101E-4,30.000077039002225,1.73326699081293E-3,43.33345189690413,9.2842130570216404E-3,406.67267511279374,-1.6125106494367699E-4,33.333456814287253,6.4498463354536995E-4,20.000072817006615,0.420323017957279,10053.263062435804,8.0674411168897606E-2,503.34268455690125,2.3190630998677401E-3,136.66742756961642,4.2641864444097302E-2,1823.4417452895384,4.3346956522400499E-4,30.000141412017623,-3.5228408767120002E-4,40.000160733856355,4.1268826069966397E-3,336.67071158681239]
[null,false,"029SMPL.d","2019-10-22T12:06:03","Sample","19286427-005",false,false,false,63.957129155750799,14872.974219562529,12.4735318549826,680.01639306520838,-7.5361235531122501,58225.829210492528,6.7252729930738999,66.666857997573558,11851.001813765601,2770.215311096877,285.02617760894799,66723.121608565489,169.68423873056901,2043.4535421413154,0,27788.111565091542,0.10476775793160401,28.000027974446613,5.3831941854075797E-2,703.3484417694541,7.5796661041408003E-2,2636.8672295878246,0.101622225492196,753.35089514626725,8.6287165415228806,64520.51535816153,7.0360607956609803E-4,100.00035136937572,-1.6262620476379501E-3,590.01263057173344,0.28470451374449501,3483.7150080835695,218.117592522671,431113.4210944692,0.51462182134938295,681.68072002608494,1.2793311187244E-2,60.000188251333611,3.8208456554634E-2,80.00034848848459,-1.4994038273350599E-3,3.333335667848552,672.05543941269696,23222.072407870404,5.44315261297523E-2,216.67012857901841,1.5159689465991899,3527.0291629379622,7.9967707132374694E-3,166.66762421527071,0.107037425950034,886.69025371436908,1.77888224595105,15952.368740155764,"<0.000",0,0,261981.31771075749,-5.30159097233432E-3,16.666691849628709,"<0.000",0,1.39475788091429E-3,10.910006080220185,"<0.000",-22.333394274114653,-6.0295728820621702E-3,40.000103513398216,7.8826819432781997E-2,475.00717726835842,"<0.005",0,-0.27723039400175198,700.05571022754259,1.23428227756802E-3,15.000021631518678,95.965010989977401,164223.72151120918,8.9702426846582703,402921.72576683178,1.2377648646892999E-3,10.000045150517744,2.6264629660815199E-2,496.67503888395868,2.6864459386014E-4,23.333369145790201,1.5619724608350999E-3,40.000161528584933,7.8923330673012805E-3,356.67122627291241,-1.7076918219200899E-4,33.333447277544231,1.3643032768986001E-3,43.333491633333388,0.20663556583241299,8235.5691843234581,7.12063582455357E-2,456.67423839359134,2.1450936139706701E-3,130.00094220040825,3.9433202304981298E-2,1743.4323368951827,7.7093756402037399E-4,46.66686783233979,-2.5479158834783701E-4,46.666804254052977,4.59675386036177E-3,383.33969180773164]
[null,false,"030SMPL.d","2019-10-22T12:07:32","Sample","19286427-006",false,false,false,54.875119249938301,12437.776207774841,14.7924429382463,780.01873125604686,77.371123563995795,58647.14430838111,11.185961858426399,90.000526358676055,12044.896450369601,2736.877909249893,362.71816338720799,77448.947857420964,185.749422782268,2150.1360763589132,0,27073.778893444261,9.9126608283260906E-2,26.666689554850262,8.0092948920798407E-2,846.68794006077542,5.65384947033368E-2,2383.4956513321054,2.88658971331354E-2,546.67530516770523,2.1553622098460501,22948.565169515921,3.4520371172484599E-3,148.33414810398,9.6731062227653104E-3,643.34572151284829,0.34319494723718602,3917.1145996799855,417.39551057858199,806841.5870447316,0.64077697895722596,830.02076501616693,1.35113016536361E-2,60.000200966990974,2.0434508053648499E-2,46.666818559167503,"<0.013",0,970.86536304905405,32647.092612478747,4.81704330515655E-2,200.00334233053047,1.7801860889902299,4067.1537683334886,1.0573837845707999E-2,216.668354620511,0.111629564911262,906.69307629228024,1.8017301900087099,15907.324388613437,"<0.000",0,0,258285.32818249869,-5.9012126817211503E-3,13.33335538705157,-4.7806549619083499E-6,11.111166990466836,7.6793361889954603E-4,5.74375181858123,"<0.000",-31.666746780275105,-1.4005810604407101E-3,58.333470746872763,9.8044289025989301E-2,580.0102274615034,"<0.005",0,-0.246425035311678,733.3638831979905,2.3180481563130199E-3,23.333384245633322,97.237542130627105,164053.60107103584,8.9614494336536907,396847.9758573377,7.5328601686599104E-3,60.000259776906283,2.3422123313430299E-2,436.67339310039989,8.2048317787765997E-4,63.333623260255308,1.98040354358719E-3,50.000170121587757,7.2251323739871403E-3,323.33747381959148,-1.2544925730550099E-4,36.666844934222425,2.1289586725666199E-4,6.6666768987972,-3.52853721518194E-2,5817.7987445398876,5.9756788409716301E-2,380.00515689444802,1.9419429618965499E-3,116.66741197306794,4.5051042123841603E-2,1933.4528317036304,1.07145010943114E-3,60.000479519360077,1.03963220988186E-4,66.666892965631291,7.0580086175119697E-3,576.67871251684358]
[null,false,"031SMPL.d","2019-10-22T12:09:02","Sample","19286427-007",false,false,false,64.068500692308902,14472.846291104002,15.1021333090914,793.35305621197233,67.114353244794103,58242.51215295228,11.233547058553,90.000504901004263,11350.426802199299,2593.5215372812513,335.74419777819202,72913.044842624178,217.43296133464699,2456.838037120729,0,26993.632410263537,8.7991227915685602E-2,25.333368301391602,5.2794896760962003E-2,676.68092355110855,4.3144485670675699E-2,2240.1453828789981,8.0459430694323195E-2,676.68178821580932,3.4601686468015198,30827.012814688471,3.0540022254501998E-3,140.00074177979317,-3.18862478866458E-3,563.34422225737228,0.289058525586674,3417.0144526986533,380.54014810031703,728940.7226733031,0.545177075176982,700.0145362812101,1.8891253281115E-3,40.000094771383779,1.6954960953962501E-2,40.000113844869823,-2.2346155436415601E-2,1.666667833924276,783.78577243793495,26146.217058307677,0.13206852657176399,366.67381947255234,1.7183582434834399,3870.4434644613152,1.20547561391271E-2,243.33559309439468,0.119835052259715,953.36279402132641,1.74796089906184,15175.003842885237,"<0.000",0,0,253506.56523287448,-5.8525687509638803E-3,13.333353002865818,8.6272614060890296E-4,33.333604286120433,1.29112894788909E-3,9.7437549831905041,"<0.000",-31.666747813422266,-1.1312263804383201E-3,58.333477104701444,8.2396612973842695E-2,480.00735521305569,"<0.005",0,-0.45461256635144398,433.3476322395818,8.6624672639661103E-4,11.666693116227393,100.343381635691,166160.81783651086,9.2211883831718904,400794.88986102416,5.9693272981205404E-3,46.666807432967317,2.91465486414158E-2,533.3435869215391,6.0446412371036105E-4,46.66683048009628,2.0177397786928699E-3,50.000277807311051,1.10517983342937E-2,476.67465771226102,-7.6750004576004095E-5,40.000101526576756,6.5072828518225297E-4,20.000061690806422,-0.238008778753936,3817.175123976112,6.86768673902933E-2,426.67374292998898,2.6407675021586701E-3,153.33445638416524,4.5916267592408397E-2,1930.1227027656232,5.8514638283932202E-4,36.666797250507322,4.2067836209728397E-4,83.333653161918321,7.1505663573191202E-3,573.34485789122891]
[null,false,"032SMPL.d","2019-10-22T12:10:36","Sample","19286427-008",false,false,false,58.450414751316998,12504.411341100191,17.796398760975801,880.02287507023118,159.34964988054901,57112.310346876933,14.022929382925099,100.00048488377803,13255.411457689699,2806.8881824579607,347.557717019775,70821.538855120176,198.25841764348201,2146.7997930427423,0,25558.176702486828,8.7374447419496404E-2,24.000025431315105,5.0229962977707501E-2,630.01184160493051,8.0314550165689905E-2,2490.1757441434897,0.10910580927215401,720.01623988127346,5.0217383576821399,38909.616958564897,3.8007360974317101E-3,148.33411313592225,-1.0081240826155099E-2,506.67514046020597,0.35537300399028299,3875.4600119569404,290.51746641676198,542552.73022595688,0.61789212781553904,775.01776494059322,-6.9804828803798598E-4,35.000084514417978,9.6580410133061904E-3,26.666733026503525,-2.2049665034831599E-2,1.666667833924276,688.00083429181598,22791.463690328641,7.4792445764409493E-2,250.0032242388254,1.54159905220056,3470.3558849738656,9.3941801273503701E-3,190.0012560188583,0.11202683708515999,900.02643515626926,1.7517996887885301,15375.187352272063,"<0.000",0,0,258560.84250130985,-2.0373084414770401E-3,33.333410720029313,4.161482170186E-4,22.222437295653595,9.9614164262248005E-4,7.5766679750383341,"<0.000",-24.000065684317565,-5.9017828282609403E-3,40.000109473862608,9.56465053201144E-2,566.67658915107563,"<0.005",0,-0.34199925347501398,600.02267907113855,3.15953447056586E-3,30.00007227063071,95.514555141757796,161318.87129407816,8.9689410671023104,397603.40553358087,1.6721817700927899E-3,13.333363731701716,2.3933003164835501E-2,446.67400966077031,8.1955948360988801E-4,63.333604186769264,2.7696098796044298E-3,70.000219891465406,8.7085060905308792E-3,386.67252282292856,-1.2591574646687501E-4,36.666763871906738,3.1900383505443998E-4,10.000025282303115,-0.118207133402691,5034.2630875815148,6.2405243328435098E-2,396.67224292945502,2.4103435176756199E-3,143.33431666094586,4.7002199043130098E-2,2010.1341965277861,-2.6629681897535602E-6,10.000015348195801,-4.1769731961250402E-4,36.666780561207027,3.62774315845778E-3,300.00334610536675]
[null,false,"033SMPL.d","2019-10-22T12:12:06","Sample","19286427-009",false,false,false,62.204869761249803,13872.116907559464,17.111770426736999,883.35753078225946,73.287912316501703,57630.742956339178,1.7924861579253699,36.66677340864976,10414.121571826099,2380.1584524886043,318.87257586417098,69316.281660016713,188.54451019598201,2143.4641709407538,0,26646.479095124017,7.2497684118345798E-2,23.333354314168293,5.7897070403559001E-2,700.01507033881933,0.15472884628380801,3300.3089921627093,8.8069963366828402E-2,690.01463815548061,5.1435651787493804,40847.78126783677,3.9361517453107798E-3,155.00094182787919,4.9786422421463902E-2,866.69022858107746,0.47095658725037998,4959.0389088574802,294.05408904202301,560943.4370172295,0.49989215567300099,640.01199404380725,3.9576982819256104E-3,43.33341692884639,1.51048445595773E-2,36.666775792835516,-2.2327092173658299E-2,1.666667833924276,769.42392889263704,25678.858666834352,9.0546105713352795E-2,283.33727754661317,1.5788814372809601,3567.0378089889218,7.7341494407914E-3,156.66751628119474,0.113893904508163,913.35877646964627,1.79439277057172,15672.115651883947,"<0.000",0,0,255677.82171093687,-5.8748937410074299E-3,13.333364526430302,4.25651089814039E-4,22.222437295653595,4.0005366507700301E-4,2.7737486481604936,"<0.000",-25.333400428294134,-4.5541751981371201E-3,45.000113770364017,9.0661706805041195E-2,531.67563599831408,"<0.005",0,-0.38522424675620898,533.3491126699796,1.9211459428094501E-3,20.000029901663016,97.249711718179498,162417.7286995789,9.02311248878115,395544.70764759119,4.6503640638559697E-3,36.666841357943795,2.14935746447973E-2,396.6728151340364,3.2248195553474199E-4,26.666766405104102,2.26734724940892E-3,56.666859736042319,9.7637828260273007E-3,426.67275110871469,1.99976186170535E-4,63.333607365683605,8.6027093013494704E-4,26.666790246961657,-0.16359427796051901,4550.6853485012252,4.1722154506029403E-2,266.66940291718538,2.31988938931252E-3,136.66756744184741,5.4294641989416602E-2,2263.4990609164179,8.6762082017945196E-4,50.000186810888046,-2.35092838049771E-4,46.666837632653547,4.9623449608090101E-3,403.33977416148127]
[null,false,"034SMPL.d","2019-10-22T12:13:36","Sample","19286427-010",false,false,false,66.857371943972893,14806.254908582305,16.088222268114698,826.68626189202678,109.419063821162,58038.563588283003,12.7684664175284,96.667062590513694,11319.041249136901,2536.8455127351654,317.27421279048201,68596.902981827967,196.235862420177,2203.4695106227769,0,26466.120211092828,7.3658333237025503E-2,23.333360036214192,6.6150378606981503E-2,746.6835292177866,8.8816971919023199E-2,2653.5340994566563,6.88743103942666E-2,640.01270612061955,3.1788354439817099,28857.288242666906,2.6413969002538899E-3,131.66731956103465,2.64475759801178E-2,733.35043768062542,0.270304115696543,3250.3105890075244,281.62266798591003,540033.49188839749,0.52403589416738905,675.01390087087111,1.7674693996054701E-3,40.000086029369349,1.49184532798671E-2,36.666756719349472,8.7460197728880307E-2,10.000014156102923,722.70763405308696,24437.039837682649,7.3417781863249999E-3,116.66893425092202,1.5226497449109599,3493.6892179595052,7.1215847143281004E-3,146.6674973567203,9.4545216722237702E-2,783.35399731962889,1.7816201821519599,15878.977221067093,"<0.000",0,0,261859.47526465644,-4.9819493090849003E-3,18.33336206773874,"<0.000",0,6.3287860971198296E-4,4.728335449968819,"<0.000",-26.333397552370069,-4.41436427473757E-3,46.666775643823911,8.7066295775205499E-2,523.34217563258346,"<0.005",0,-0.25356818709506901,733.3599731332057,1.44365454111381E-3,16.666690260171539,96.074624975813407,164334.81851476207,8.8939417893521302,399308.70498733665,1.2383371025254299E-3,10.000015348195801,2.8040453213440501E-2,530.00998660907169,6.2869952668266095E-4,50.000158200658994,1.8231469423609201E-3,46.666799485681473,7.5861460456890196E-3,343.3375116685404,"<0.000",0,1.15494774077605E-3,36.666804800428878,-0.27869960927957399,3550.4681427136397,3.26446067358694E-2,216.66883781549078,1.7976419479327199E-3,110.00055149196756,4.4790649250225002E-2,1950.1225498797114,4.18322911082519E-4,30.000099291402609,-3.1151554015604701E-4,43.333521038291032,4.5989241887317601E-3,383.33908781400692]
[null,false,"035SMPL.d","2019-10-22T12:15:14","Sample","19286427-011",false,false,false,70.725195464700406,15807.117052535346,18.401192428189798,950.02597664754614,38.6047802452917,57002.007614145557,10.8006506799583,86.66715174912683,10902.1372842981,2480.1778296603084,348.16706834220997,74118.422066832078,210.49701375356901,2363.493210223915,0,26713.215747775404,5.9792930872762703E-2,22.000020345052082,6.2524253037442995E-2,730.0164940454091,0.102286607388701,2800.223753846331,4.5607166177020202E-2,583.34372604871203,3.6089003542958502,31594.750665507603,3.1724604001193701E-3,141.66729398070831,1.4811106796350801E-2,666.68031334856676,0.35923699589276398,4017.1504600702579,424.157863222245,812275.06698388758,0.60474546244534,776.68450787636868,-1.1532935156376E-3,35.000069811939156,1.31232322964355E-2,33.333418667315165,-3.21396827130775E-4,3.333335667848552,969.13071007413805,32356.500982694313,1.6017083348509001E-2,133.33524366224108,1.5770580440351001,3583.708849593369,1.1122954468890301E-2,226.66847527024436,0.116256765730642,936.69355904022507,1.76368772742391,15505.30046882281,"<0.000",0,0,257494.64460053077,-5.89328920596112E-3,13.333353002865815,4.1963777019966299E-4,22.222437295653595,8.7655497605373001E-4,6.5920894701469095,"<0.000",-31.000095859168532,-2.9947152878701901E-3,51.666826431947555,9.2598277004106194E-2,546.67581379399974,"<0.005",0,-0.48334064971665502,400.01577635641388,1.0549308348388399E-3,13.333360155423085,97.370994282005498,163775.82763767499,9.1734417463870894,404992.14240264683,4.6175445601394898E-3,36.666780958571323,2.6363610357147502E-2,490.00872835503918,2.7413010795100001E-4,23.333428353069795,1.4567543572628499E-3,36.666763474542449,8.7461917511704092E-3,386.67173445217213,-1.6394628181942001E-4,33.333467940487445,3.2032473435831101E-4,10.000025679667408,-0.27949523536918902,3483.7405665558204,5.2315048137355699E-2,333.33737194532097,1.9482707213708901E-3,116.66735157369546,5.4135033061711302E-2,2273.504452206128,"<0.000",0,-4.7312829970648301E-4,33.333428204058187,8.3221734086962491E-3,676.68255115525096]
[null,false,"036SMPL.d","2019-10-22T12:16:47","Sample","19286427-012",false,false,false,54.299929407721997,11870.647695749574,16.314427699094701,826.68595671625008,185.554196797641,58921.377702732221,5.00569377526491,53.33349148432179,11711.4644667096,2576.8581260712222,312.38376308351002,66924.391144314213,192.34309940337599,2136.7984199007587,0,26112.244077043073,7.0202705128986603E-2,22.666690826416016,3.8413741102069E-2,573.34389150126935,6.4031945181542399E-2,2386.8298492804684,7.4148029673349006E-2,646.6792201000568,5.3034807113766202,41415.142730713596,3.0769996253284301E-3,138.33406242230441,9.5516260313095493E-3,630.012490103456,0.27899731708871101,3300.3176473152112,215.75779274737201,411121.75398095394,0.638715062466928,816.68608402183543,-2.1059736828014002E-3,33.333402772743462,1.12501990961438E-2,30.000080615280854,"<0.013",0,738.12940090994596,24867.622513558134,9.6960849786132203E-2,300.00429824004243,1.5431304910118699,3537.0383180622507,5.3425205607059998E-3,110.00059917568267,8.8772574720244801E-2,740.01819779448408,1.7739615844199399,15840.603605102448,"<0.000",0,0,262909.08462401602,-7.5310543878721803E-3,5.0000062833228762,4.0220611135031499E-4,22.222333980933673,1.1387176017906701E-3,8.8645896238399473,"<0.000",-21.666726027924923,-5.6625894011899398E-3,41.66677412887254,8.0232684734673099E-2,485.00743143249406,"<0.005",0,-0.37243173799194901,566.68786679752372,-2.26007266275869E-4,3.333335667848552,92.344249850109406,158587.83617852375,8.6115700637676902,388180.98379898589,2.8779247119166602E-3,23.333425968884043,1.9848449694216602E-2,376.67163287592479,4.0196093597831502E-4,33.333435356615453,7.78229906430355E-4,20.000072022278029,8.0174881843363098E-3,363.3382997412736,-2.5018609345317598E-4,26.666780710218635,8.3660902245721403E-4,26.666779915490054,-0.32664853263091898,3100.3572521190231,4.7988932205609201E-2,313.33716879282639,1.8478411774795801E-3,113.33390067020207,4.9819220673739999E-2,2153.4831425525276,-7.6543755765266601E-5,6.6666880249973923,8.3603154494653906E-5,66.666948596632253,4.9452290533102596E-3,413.34046065796718]
[null,false,"037SMPL.d","2019-10-22T12:18:17","Sample","19286427-013",false,false,false,68.303715744688304,14706.141991305545,15.456940141789,773.35058967246721,127.427414042857,56814.798958081024,15.1448666505791,106.66720867156175,12847.491941205601,2750.2190243668506,347.038681279971,71223.53064470434,215.11492039474001,2320.1537740207646,0,25731.75580146279,9.8459285203896907E-2,25.333353042602539,5.9246874663452499E-2,686.68096154906891,9.8862512620969498E-2,2680.2042243847613,6.0558565879990603E-2,603.34438696487166,4.2667486148368301,34634.236410528734,3.1918969342189002E-3,138.33394162355947,2.9986796445696798E-5,566.67634755358574,0.365964756763349,3983.799451227474,351.72421260578801,659795.49000799516,0.59248316360884101,746.68377081527649,1.2415855588389299E-3,38.333431308466722,5.7768338510524198E-3,20.000047385691889,-2.2134950484808399E-2,1.666667833924276,839.74705356869799,27785.423576850342,0.13224469050398499,366.67621955296886,1.5920184733034299,3593.7119012021235,1.0519701664969299E-2,213.33517392473655,0.123802281000383,990.03159637205408,1.7640420919038999,15513.641940052679,"<0.000",0,0,258908.33598571515,-2.6895220559379198E-3,30.000136246281823,8.3582600656752703E-4,33.333604286120433,1.28040510998874E-3,9.8654180042136161,"<0.000",-45.333470702169279,3.45227526713165E-3,78.333564226822602,8.0355184606448499E-2,478.34072671819632,"<0.005",0,-0.29512846956957201,666.69330497642295,1.4674203475463501E-3,16.666691849628709,96.490957886081304,163186.81071936869,8.9679347088788592,398093.09302628395,3.33987480749985E-3,26.666726668674844,2.4257598417161401E-2,453.34054907152239,4.08884836885742E-4,33.333425422508135,2.6341926542619099E-3,66.667108337077877,1.0656136587113001E-2,470.00773683178812,-4.7259989117928299E-5,43.333472559847344,1.0619249190902201E-3,33.333605031168382,-0.323449168421726,3083.6898857186461,9.6475276569859894E-2,606.67932242136214,1.99573022116217E-3,120.0008374452466,4.0489486373854099E-2,1763.4324891850476,1.3988823832937499E-4,16.666703770557486,-5.34045994941432E-4,30.00010247031695,6.2171854376920303E-3,510.00919853633849]
[null,false,"038SMPL.d","2019-10-22T12:19:45","Sample","19286427-014",false,false,false,56.487214783800198,12871.582064639128,13.9240879972408,740.01593440747354,-88.406565954217996,55199.831534229248,5.1951799936415801,56.666982124244441,11850.508083508301,2713.5399986276029,312.82109968654902,69835.393813546849,187.92511499188001,2183.4729695796027,0,27220.572511048533,8.5863872171037497E-2,25.333372433980305,4.0279608787009098E-2,606.67811443391281,0.10111801268904599,2840.2334171498801,0.105326295410994,750.01690064842148,5.7789780936964998,45692.269901510153,2.676334462902E-3,135.00060198206799,1.6198617224039299E-2,686.68112685261474,0.28083615315953703,3400.3440345403956,274.81878578115698,535677.6975202047,0.55981444033149297,731.68211800333677,1.4408466203047999E-3,40.000094771383779,1.09633561954833E-2,30.000080615280854,-2.2800671514793602E-2,1.666667833924276,715.23075313335596,24433.694297367914,5.5299253168971801E-2,216.66917490468072,1.5885195077746399,3667.0586611768799,9.9843753947245796E-3,206.66813860334744,0.103803981879735,856.6889317330382,1.7794838433956699,15882.307299142471,"<0.000",0,0,261287.79558567508,-5.9308665494295096E-3,13.333349823951474,"<0.000",0,1.1454752353233599E-3,8.8645890019648128,"<0.000",-21.666714504360439,-5.6003790120762002E-3,41.666750287014985,0.102850066585568,615.01163837793013,"<0.005",0,-0.39346996429560799,533.35801363046494,4.03438595407296E-4,8.3333538721002061,94.660604108147396,161562.91100398335,9.18817682183907,411618.18397057778,3.3094597956356002E-3,26.666727463403429,2.3506461416042901E-2,443.34019953995647,6.3017661526502901E-4,50.000203500188341,3.3932606544710702E-3,86.667126317812105,8.8465021151036202E-3,396.67230650774178,-5.1950904979081598E-5,43.33344156543253,7.3657513822569895E-4,23.333431929348432,-0.28136256364519002,3517.0876335446751,7.8922416050916905E-2,503.34222679323619,1.5689934487376799E-3,96.667124579343337,3.8431232107017099E-2,1700.0941722577688,7.0279056841255499E-4,43.333495606976307,-5.3877789958971896E-4,30.00016565123947,4.8948255120185701E-3,406.67408655076093]
[null,false,"039SMPL.d","2019-10-22T12:21:14","Sample","19286427-015",false,false,false,63.5778400523578,14038.956510897995,15.463809759882199,793.35191180280981,15.646297550764301,55798.25925968698,6.7308575351604301,63.33353425065377,12252.514132517899,2706.8733320602773,327.93060391586698,70052.263935581781,193.82065469075599,2173.4693584819233,0,26385.975355516246,0.111537309019748,27.333358764648438,4.57099904898636E-2,620.01217236103355,0.10495618810681701,2786.8905224369387,6.0805176776318498E-2,613.34544221542126,4.0197539876833099,33598.822772476582,3.8529172107351498E-3,151.66759582855906,5.3384645435586901E-2,876.690584470472,0.29330559697439701,3392.0051061313497,385.43817485738998,725553.40163404902,0.54520784377707798,688.34776806313141,4.3160124126236802E-3,43.333424876132234,9.6486328527090309E-3,26.666752099989569,-2.20064829124345E-2,1.666667833924276,995.33055514510897,32560.294111167972,0.159642847147559,416.6744245838866,1.6460330835327499,3660.3911808074831,1.0700945551892899E-2,213.33525657650944,9.9955382088793704E-2,796.68664380878238,1.7585091948324501,15101.598896062389,"<0.000",0,0,251158.60953869508,-6.4914251842484701E-3,10.000012566645752,4.4098208371763599E-4,22.222333980933673,6.2247742713692697E-3,48.016351964628996,"<0.000",-25.666740273434815,-3.5141406535439601E-3,48.333464538055694,9.26192477809019E-2,533.34258238460745,"<0.005",0,-0.35384382076486498,566.6987386849737,1.7498730484774301E-3,18.333376372853273,97.538174278575298,160020.12782384761,9.3163951058076897,401182.52006358124,6.8858801191549503E-3,53.333614269888187,2.61093193850696E-2,473.34124813465405,7.9816714295288501E-4,60.00019619861947,2.9870155471765302E-3,73.333700994645042,1.08355993388261E-2,463.34187135087637,-3.1366123160499503E-5,43.33355203270586,1.3136225033944899E-3,40.00010112921246,-0.36032480678668299,2650.2410845264044,4.5865823774178298E-2,286.67000025505814,1.6976327241228499E-3,100.00049283106387,5.2140886600366497E-2,2143.4818202731735,1.10801239200574E-3,60.000302692249875,-3.3942135465668002E-4,40.000117023784163,5.0531338360874098E-3,403.33951984833402]
[null,false,"040SMPL.d","2019-10-22T12:22:47","Sample","19286427-016",false,false,false,64.338055460304801,13872.169296069749,16.4040874430389,820.02015481362503,76.868068087858802,55801.577944533354,7.6335130268620004,66.666934291517734,11841.883263007199,2566.8496830237455,385.41584639639501,77164.35922801304,224.33958250982201,2410.1659288480987,3.5527136788005001E-15,25765.14854185062,6.6777352045271895E-2,22.000019709269207,6.8221634618199398E-2,740.01695166006266,0.106653330369191,2756.8843176431806,6.9643892330284701E-2,626.67876263441497,3.9969849262681598,33064.432501742973,1.49445672905838E-3,108.33373226224781,4.9220869625525202E-2,846.68972025280618,0.29727130601920398,3408.6881127908655,507.14327913143399,951896.60330779711,0.61855460539460305,780.01824806106708,1.2216987226493399E-3,38.333405877151996,1.14923561221535E-2,30.00010763605275,4.5063187438641801E-2,6.6666737198828594,1040.1174486495199,34297.033994142701,9.9096502541444295E-2,300.00621154918247,1.6889125896286601,3813.7564535872675,1.8070425715449999E-2,366.67106082035508,0.11613449529183301,933.36035291313601,1.7146579957332699,15086.597675582801,"<0.000",0,0,258996.23963023862,-5.5866152125861601E-3,15.00003275771887,4.14731470018987E-4,22.222437295653595,1.2099927773813299E-3,9.3041725417351202,"<0.000",-28.333402290939254,-3.47545603927825E-3,50.000116874772552,7.8897535798586493E-2,470.00701203931857,"<0.005",0,-0.53243144205895299,333.34122449177261,1.4667040783514101E-3,16.66668747862149,94.690574526026793,160196.66427955034,8.7824801896587807,389993.06710531708,7.0948306285494797E-3,56.666837881006238,2.44276517109254E-2,456.67360261072309,1.31844969544787E-3,100.00047375757784,4.0816099717369296E-3,103.33394532401445,1.07307653647256E-2,473.34169318266169,-4.7434813304462399E-5,43.333471765118759,3.18467243124394E-4,10.000015348195801,-0.42315995222007902,2133.4986209906328,4.4953370632290698E-2,290.00382309152945,2.3473642426840502E-3,140.00087847310979,5.9271551246526902E-2,2483.5314062704929,8.5372148481862502E-4,50.000194758173905,-4.76494339540468E-4,33.333451251187157,8.8909719953279807E-3,726.68575475584839]
[null,false,"041SMPL.d","2019-10-22T12:24:17","Sample","19286427-017",false,false,false,61.645970349747898,14205.923779441204,14.3407395227092,770.01763785852529,-138.958506138053,54674.958495175466,7.9992419686737097,73.333602448300482,11424.4994922927,2660.2003591229463,349.28878290366401,76579.760419373328,191.04958457604101,2240.1438570001142,0,27534.473059756463,0.124747112431222,30.000040054321289,5.0080777269534402E-2,673.3467637497173,6.7016751322083595E-2,2523.5139089299155,5.2535187326002403E-2,616.67843217633515,3.9052331585470599,34234.063468641463,1.4739341326274601E-3,113.33384186028677,1.11850390017698E-2,660.01405368227677,0.39884147029231998,4450.5717954415732,367.60516944843903,718177.21331813652,0.51141523492097396,670.01323973119395,5.3443387071382504E-3,46.66676928599523,9.11614179423249E-3,26.666723489760511,8.6123784053575195E-2,10.000017335017263,1000.43041909909,33862.722477176947,0.104269684675025,316.67419187741416,1.63832982038819,3763.7510628936034,1.5708481324510901E-2,323.33695883546829,0.107188033235614,876.69002498154805,1.78157578636596,15770.538095078795,"<0.000",0,0,258476.06385673987,-6.8700973788234703E-3,8.3333419511714286,4.1642071568717502E-4,22.222333980933673,1.3357529619286699E-3,10.291857160153313,"<0.000",-107.67190748942036,2.9595116449265602E-2,185.00874323136495,0.10054947560118099,595.01106647137203,"<0.005",0,-0.484423917930277,400.02070367382549,2.1046339725258999E-3,21.666702901323095,96.037420250377707,162148.68573483793,9.0296124763380803,400161.76487045846,4.1818298522242102E-3,33.333443303901305,2.1260875259550699E-2,396.67259896786118,5.4637883027044596E-4,43.333521038291046,1.9789412442350802E-3,50.000147074458788,8.5544363586212892E-3,380.00544935456736,-4.6398728190715699E-5,43.333461433647152,4.25477862286116E-4,13.333364129066007,-0.34392265348295897,2883.6171476854743,3.4729826422427897E-2,226.66854520635988,2.2346392628174402E-3,133.33410064378231,7.87616563113364E-2,3223.6675545026724,7.1279779939564E-4,43.333466202018663,-1.8611252329355199E-4,50.000132769344255,6.8464854892321602E-3,560.01139760000399]
[null,false,"042SMPL.d","2019-10-22T12:25:45","Sample","19286427-018",false,false,false,58.307881652263497,13572.183530901006,17.076267817689299,920.02455294095637,-154.321951244579,54862.14924759444,8.4507575660661605,76.666921426848745,11392.3883940342,2680.2100227245187,340.17320235418401,75841.945054574579,213.42422983313699,2490.1766088081904,0,27808.268628593258,0.18630623365305601,37.333386103312172,6.4954133874775499E-2,773.35158149374149,6.8509588528921306E-2,2563.5189945968141,2.0153140559928302E-2,536.67571221775245,3.71752612484448,33381.658664432485,1.5918779315000401E-3,116.66715448100634,-4.8835732299261801E-3,570.01015131657095,0.57883737183634099,6092.7809487656459,449.088718298103,885756.79063154047,0.51786985303475397,685.01374813397103,7.0220104826644999E-3,50.000131179887092,1.4469937136235801E-2,36.666775792835516,"<0.013",0,1065.3321867919101,36367.685004954263,1.46081451668502E-2,133.33500325683514,1.6399898443665,3803.7542157805824,1.1869903029104499E-2,246.66878650582649,0.13190147990241399,1073.3696841194776,1.76958534445178,15813.89950393787,"<0.000",0,0,260927.85418478865,-8.1624313672053993E-3,1.666667833924276,4.0850317004660798E-4,22.222437295653595,1.64214614872212E-3,12.86459021631012,"<0.000",-25.000065430997825,-5.5864554019876998E-3,41.66677571832971,8.9937919854554405E-2,538.34278576061945,"<0.005",0.66666690508524573,-0.55772527777499803,300.00641221815772,1.24185547174332E-3,15.000026399890189,96.454347173466999,164397.29572737272,8.8912798207507802,397768.9914686135,4.9710398522365704E-3,40.000109076498312,2.4246810269796401E-2,456.67347545414947,5.4081087637279396E-4,43.333483686047543,3.3979394895640399E-3,86.667073865725499,9.5593446140983702E-3,426.67319615672233,-5.1246256573700501E-5,43.33349242806198,8.4296045693354105E-4,26.666749715803817,-0.424816300568153,2133.4910551742205,4.5680154041588102E-2,296.67005732650472,2.1542327683060202E-3,130.00081186492028,6.0346635598647398E-2,2543.5450874237454,2.7885775596816603E-4,23.333380271990393,-2.23610752180716E-5,60.000311434264312,8.3341770712614897E-3,686.68406416946561]
[null,false,"043SMPL.d","2019-10-22T12:27:15","Sample","19286427-019",false,false,false,60.6357205059766,13905.862483264909,14.413630744665999,770.01781587772848,-65.500410782789501,56089.087705874139,9.2349848941454695,80.000377098713656,10427.4251213999,2450.1700989876667,391.96183283406998,83123.98053635117,231.35403758936499,2633.5296238432879,0,27400.83292502911,0.137519159324024,31.333377838134766,5.5705271361547098E-2,706.68160974957129,7.99584578519595E-2,2650.2061576117153,0.14362458698201899,856.68867741989072,8.4607420944681699,62911.679466081456,1.7245714452641899E-3,118.33378297586565,-5.4538513284960602E-3,563.3437772093647,0.33303500867557001,3892.1237543948337,534.12191887267898,1049910.5377902458,0.61433417851324601,810.01941745450961,5.1012363930814601E-3,46.666759749252201,5.30512975124235E-3,20.000047385691889,-2.3021616647548201E-2,1.666667833924276,1040.90657753255,35649.54048440545,-1.6777943696420299E-3,100.0011903544533,1.72560509798412,4027.1430877773505,1.19395344123343E-2,250.00213886297536,0.14151945464963001,1156.7071615153477,1.68787937607236,15251.741716090713,"<0.000",0,0,264719.6249616202,-4.3903096670689203E-3,21.666723961630606,"<0.000",0,1.68959704335846E-3,13.440425530268538,"<0.000",-27.00005681812679,-4.9345154445322196E-3,45.000094696877987,8.72258687215232E-2,530.00900750345477,"<0.005",0,-0.44474879344177398,466.6873718306083,3.8989287137818302E-4,8.3333443353571841,93.922775148059301,162409.22284033065,8.8970496043797809,403811.19189941109,5.7164781182879997E-3,46.666816174981754,2.73885616142522E-2,523.34260796493379,4.8790094103824E-4,40.000112255412652,2.3187219080574402E-3,60.000221629934195,9.4933991995828707E-3,430.00654215604254,-1.9815581639204701E-5,46.666872998075597,5.19303638827556E-4,16.666712512571923,-0.40918802838791202,2316.8826938965526,4.28732857131182E-2,283.33631093298925,1.6620216064773699E-3,103.33389128247065,5.5315579178030302E-2,2383.5197093558354,2.01901082697429E-4,20.000064074992178,-2.0657624650521799E-4,50.000166147944832,8.6566481241237593E-3,723.35120076907867]
[null,false,"044SMPL.d","2019-10-22T12:28:50","Sample","19286427-020",false,false,false,55.7427625042709,12337.872335912909,15.8356088020572,813.35386971602054,-83.585249433385798,53722.111829686481,1.2365670569463501,33.333418667315165,11297.521582478899,2530.1808298103883,361.90665717157998,75507.348286028035,221.26724725292701,2443.5048057113368,0,26439.416923730325,8.6274076438011799E-2,24.666693687438965,5.5313150332931701E-2,680.01354475795904,9.1050640022922305E-2,2666.8681446681198,0.116993639367174,760.01948177785459,3.4309026291137998,30259.383814370769,2.31005451538503E-3,125.00053855279278,-1.03397518036099E-2,516.67502587027809,0.34876619829265099,3900.4425920644985,524.35399851627596,997620.07438187965,0.63201110299658103,806.68654193451187,9.6881314738865197E-4,38.333408658702048,2.07791931214295E-2,46.666809022424481,-1.15190524097274E-4,3.3333380520343074,922.44362454497605,30696.859080341572,9.0236201921961801E-2,283.34034917270924,1.6232006097860601,3677.0604920825294,1.8203588212530701E-2,370.00462934367914,0.14109660346219599,1120.0383083019906,1.7567825869491001,15423.56991991048,"<0.000",0,0,257450.94668451525,-4.2747974372973402E-3,21.666714424887584,8.4297242744752203E-4,33.333604286120433,4.8128141803936399E-4,3.4404159287549989,"<0.000",-23.666725044448299,-4.6308679069931996E-3,45.000105823078165,8.4271923935281504E-2,498.3412350464676,"<0.005",0,-0.69799903048675604,100.00148043039771,1.2672875747036299E-3,15.000020836790092,96.166682185520798,161722.95982972343,9.1679436083911607,404680.72313645406,1.3435205732925601E-2,106.66754563648186,2.1166136249820701E-2,393.33882063619063,1.23512278721636E-3,93.333804011337904,3.4438286487464098E-3,86.667062739525306,8.5112433133011697E-3,376.67206520827511,-1.2403395095031199E-4,36.666795661050145,7.4755243093332501E-4,23.333430339891262,-0.39193519679530697,2416.8751938936034,3.9783989675278297E-2,256.66934584573886,1.9486232013948999E-3,116.66744058329699,4.4339657424279103E-2,1900.11787126286,2.13733453336687E-4,20.000041822591793,4.9642342770150103E-5,63.333631604905456,8.8620057400208996E-3,720.01852869959873]
[null,false,"045SMPL.d","2019-10-22T12:30:20","Sample","Acid rinses ",false,false,false,3.8687109054585198,100.00226721172633,31.672225913667202,180.00141193467263,15705.737498603799,45319.454077291673,151.549038421809,96.66711027422879,15396.389928730699,373.33892265947276,1105.18903560786,21613.258336284074,158.234422953744,206.66822761294898,2.8421709430404001E-14,2980.2581343016304,4.1217516894230402,23.333354949951172,0.92040993978909003,190.00137999651761,9.7236532988859601,2290.1513330615858,10.081650074075901,496.67439038543324,103.14850235662,10036.391673305443,0.45313026433106901,111.66713548202613,7.6325677112876997,526.67603025819813,5.2821716886485897,483.3422927808777,1680.5548599188801,30159.471295861509,1.9855699836870899,21.666706080237436,8.8931647229512993,116.66717355449241,"<0.078",0,3.23758989975516,1.666667833924276,17679.0260020112,3653.7203435045431,7.2869017179681999,83.334408079608465,23.971175638781901,286.66933268304666,0.130358248149998,13.333380818366296,8.4266794268407094E-2,3.3333380520343074,9.8205281977543901,376.67135313146287,"<0.028",0,0,996.69846638989748,1.3289527794734799,26.666710774103148,1.31132364617351,133.33485159612448,"<0.032",0.62602107307342436,"<0.022",-530.68172665414625,57.5380453494191,906.69180472654398,2.5988977053801201,58.333477104701444,"<1.235",0,387.45883748674498,2100.1602006491071,5.4137065842728899E-2,1.666667833924276,107.777797049825,701.68272880192501,14.8647977603397,2540.1976143290954,1.08448273184408,33.333423435686676,0.13899736265078599,10.000024884938822,5.9052244498857598E-2,16.666702181100316,"<0.010",0,0.10142106418312,16.666692644357294,"<0.003",0,"<0.008",0,159.637816340744,5884.496820371528,1.6877996017904699,40.0001368919988,7.6181022196138107E-2,16.666714896757679,0.92150088246152795,136.66769459842104,0.12964932326592399,23.333391795554878,0.103952121851438,23.333391795554878,3.2008177470295301E-2,10.000015348195801]
[null,false,"046SMPL.d","2019-10-22T12:31:50","Sample","Base rinse",false,false,false,1.3355903437646399,33.333787073701828,11.3295360967316,60.000175535676249,20208.894008033101,50954.146376140729,168.62469449056201,96.667081663999738,49619.240925621198,1000.0296125804939,1165.55428867363,20445.026600373883,141.63795684123801,170.00097657241955,0,2686.8880316582122,0.64071992107780196,8.0000033378601074,0.15137575590032401,110.00049427150942,2.4221246080421599,2063.4554746232111,1.7544608852361501,380.00527133536434,19.032090910164399,9102.5990295494394,8.8992199010887604E-2,116.66715765992069,0.86071412873568798,363.33829974127366,2.3433592479697598,1308.390455880548,66.980602089816998,7635.0915708040802,0.157284768439235,11.66668040057003,0.36300088684939202,35.000094051161,"<0.011",0,"<0.263",0,800.43775659474898,1303.3830067011265,1.57539444384755,150.00298358506674,0.46204909647995002,46.666818559167503,"<0.001",0,0.37895003429902402,120.00057836372785,0.82382421916640602,283.33610112464277,"<0.003",0,0,9305.8946367219887,1.2091043828375001,226.66838626064282,"<0.001",0,0.17649112877161399,50.833923148605777,"<0.002",-407.0088655698965,4.7660932331370702,703.34814930933487,0.117633873328406,25.000071153043638,"<0.132",0,22.990256992175102,1200.0591546310295,5.2256447280410803E-3,1.666667833924276,2.1288556990763499,130.00059887765946,0.19422081050876699,310.00347947075744,"<0.003",0,"<0.001",0,7.54948528825229E-3,20.000061690806422,"<0.001",0,9.3387433685423904E-2,143.33419904111526,2.0381864294931301E-2,63.333658228313062,1.18178689309376E-2,13.333364129066007,43.5237032570911,15140.637545450872,0.11847174905912,26.666729847589185,9.6873475533201005E-3,20.000053743520571,4.1412026142237197,5707.7029585683895,9.1186904597052004E-2,153.33479334908534,0.13551258921200299,283.33633636430392,2.2079397016571702E-3,6.6666880249973923]
[null,false,"047SMPL.d","2019-10-22T12:33:19","Sample","2% Nitric acid rinses ",false,false,false,"<0.168",0,363.97680721863298,486.67465756324941,81143.830222186894,49797.753164205649,220.45967723072999,33.333426614601017,79638.635610576704,420.00551233680773,4526.9893569680999,19610.714142760702,645.36209683667005,176.66799917815129,2.48689957516035E-14,713.35008194024238,4.4973107795672904,6.000002384185791,1.54595411146047,73.333586553728779,37.551870201168803,2050.1224466643366,38.244685076026897,436.67278910667511,419.39433197928003,9469.3239421246344,1.1313041891186699,65.000173871713272,23.9018586391722,383.33857918771236,24.4683533564748,516.67693021419734,134.359440149899,566.68495809079366,8.4862455083365607,21.666699325044465,8.1316680625832305,25.000036184985891,"<0.332",0,13.9670271783367,1.666667833924276,19352.069932225899,936.69279610078354,25.012052403413101,66.667858262883442,1.1832128517448,3.3333380520343074,"<0.042",0,2.38049750424508,20.000066459177933,14.6626342993228,131.66739903389316,"<0.120",0,0,233.33551059163347,1.4197393347314899,6.6666768987972,0.46650159238074901,11.111166990466836,"<0.138",1.6247349338036177E-2,"<0.094",-29.000075861810512,13.0895408928568,48.33345976968419,8.2504282190425595,43.333437591789597,"<5.277",0,236.12288688711499,300.00838314512231,1.6372707899152199,11.66668040057003,38.266289339581398,58.333496178187488,3.5827550725515902,143.3341481784858,"<0.139",0,"<0.059",0,"<0.015",0,"<0.044",0,0.173570883652684,6.6666872302688072,"<0.013",0,0.47132100912875002,13.333374460537614,59.468920457143398,516.68111048671108,4.2088061922448698,23.333400537569315,"<0.020",0,2.60221338536613,90.00045483310339,0.71300626019811397,30.000141412017623,"<0.019",0,"<0.014",0]
[null,false,"048SMPL.d","2019-10-22T12:34:53","Sample","Blank control2",false,false,false,"<0.004",0,0.69262247415840095,66.666965285932534,-358.70099849266001,49530.279209857988,0.42131315942116598,30.000113993881438,236.21960379947501,436.67318329205335,-4.0835241863831397,18993.496111386681,-12.320416542132699,190.00123058754357,0,27481.137285811641,1.8689508567091999E-2,18.000020027160645,1.1304574497763301E-2,433.33927144598056,3.3072382505052997E-2,2193.4713926890781,-5.5540944025919397E-3,466.6732082763333,0.35692561536431,12070.869461292588,1.7291371634993799E-4,90.000284761186165,1.0939054802140901E-2,666.67976657530016,-3.9899789419343898E-2,600.0143821044004,0.218347688796206,1233.3864169828041,2.3485399390022698E-3,8.3333439379928915,-7.6941632106343298E-3,25.000045324364621,-3.87260188280653E-3,3.3333380520343074,4.0500945146277501E-2,6.6666761040686149,15.2680514187578,1153.3732178801315,-1.0240800869242801E-2,83.33416568738096,-3.6346593792729501E-4,16.66669979691456,3.1325165965002201E-4,6.6666856408116368,-4.7190520436239903E-3,33.333428204058187,9.3452316447226597E-3,100.00036090611873,"<0.000",0,0,270187.26159889606,1.07729225926167E-3,51.66678351660395,-2.3362286689054201E-5,11.111166990466836,"<0.000",0.65208299042658724,"<0.000",-15.000025048851592,-9.8353686608873697E-3,25.000041748085987,-3.6938344109737101E-4,8.3333447327214767,"<0.005",0,-0.47405699542373397,433.34763223958186,3.6902404067675398E-4,8.3333419511714286,2.77862403014527E-2,66.6668389240875,2.4423698390810301E-3,116.66713858643465,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.5859503765949599E-4,26.66672706603914,1.01758829446069E-4,3.3333384493986,-0.45754624908168601,1883.4665083666605,1.43817138539697E-3,23.333432326712721,4.9770105088035699E-5,10.000026474395993,-3.14044240674632E-3,96.667084842914065,"<0.000",0,-8.8776360046398405E-4,10.000026474395993,-4.5724610687465703E-6,6.6666880249973923]
[null,false,"049SMPL.d","2019-10-22T12:36:23","Sample","Reagent blank10",false,false,false,0.15732618104365201,66.667574147403656,8.2618561995537604E-2,33.333406746386387,-163.98176864622599,50991.029513611953,4.44716803467783,50.000171711044935,686.41989155757801,496.67487358041291,-1.4381629913793299,18342.780968261333,-6.6460736332382204,233.33500832316767,3.5527136788005001E-15,25952.12119531451,-9.7815203176868908E-3,14.000008900960287,1.5498771390883801E-2,436.6729671258783,3.1551852430543502E-2,2076.7904353620047,-2.5166522866135601E-2,396.6717215875031,2.5317466964465098E-2,9569.3516644938918,2.6677699326031301E-4,88.333699181670454,3.0367089418104901E-2,756.6839995480974,-1.26557800499591E-2,816.69141687508125,0.39208836369484001,1533.4348067663941,2.51097313247544E-3,8.333341553807136,-1.52297042860546E-2,11.666677221655689,-1.9931137751869801E-3,6.6666856408116368,"<0.013",0,5.80450386312486,820.02096861569623,5.3301002498170102E-2,216.67037295806745,-3.1735082254698002E-3,10.000014156102923,"<0.000",0,-8.1327836439606099E-3,6.6666856408116368,8.0496554800625196E-4,21.666694954037244,"<0.000",0,0,275212.42428443191,-4.5463030397720101E-3,21.666706477601725,2.3445622655732198E-3,77.778478877427631,4.8007642027341499E-4,3.6675022488708571,"<0.000",-22.000039726495149,-7.2602675442524798E-3,36.666732877491917,-9.3260008857573302E-4,5.0000035017728282,"<0.004",0,-0.412498227252797,533.35898319937496,-2.4457439833640499E-4,3.3333384493986,8.5524015037946304E-3,33.333414693672239,1.4780219894864001E-3,73.333568274971313,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.7611558387883801E-4,10.000025679667408,3.6950878745832099E-4,83.333715150747963,4.9950427540060001E-4,16.666702578464609,-0.50046495497280197,1483.4175123541493,-6.6837870300724797E-4,10.000026474395993,4.6686509344759403E-5,10.000026474395993,-4.1672573252548297E-3,56.667090604696313,5.1544689525962098E-5,13.333365321158887,-9.9944781811860703E-4,3.3333384493986,-6.0135236013692001E-6,6.6666880249973923]
[null,false,"050SMPL.d","2019-10-22T12:37:52","Sample","QC1-2",false,false,false,180.71063948691099,38340.639081263915,8.0252794516439199,410.00524545701478,-137.37555664339999,50452.754782049538,449.659182698216,2393.4923959748089,290.11538547639702,413.33894749474103,509.50512352205499,94680.515625389016,378.75559869518997,3810.4137615797858,3.5527136788005001E-15,25391.269152890611,-5.1402861051569E-2,9.3333427111307774,0.44575373577775101,2920.2411470774641,0.44711397777342698,5927.6651134938229,0.46308885528959698,1600.0743881850285,17.709543709225699,114536.36323717055,0.20697248682051,3758.732752421507,0.88745615928769706,5624.2292456195364,4.5006272265695602,39043.015399046584,42.7624616138785,81589.971713000894,0.44235442630143601,565.00924857940402,-2.0786755951357899E-3,33.33339482545761,2.2798906623553399,4053.8062953878493,7.9895749025331702,608.34431059642168,7.2606901628067,853.35651397670551,21.224836651076298,44039.24455531736,9.8802994028837493E-3,40.000113844869823,"<0.000",0,-4.6586212441304797E-3,33.333409130572143,3.0144324505319499E-3,41.484127997217662,1.7972641457068199E-4,1.666667833924276,0,270965.47773313307,-4.7920896373623203E-3,20.000046193599015,0.81312160603667305,22481.536689567478,0.210932860609923,1769.0934425526732,0.211447995002015,2566.5336614270955,-7.1281577971100503E-3,36.666720161834554,7.1993426376812397E-4,15.000018452604337,"<0.005",0,-0.384230682616858,566.70067782279386,0.45216308043174602,3745.4196626133767,2.1326937795210901,3792.1002649451257,6.4137704908775897E-4,33.333432575065409,"<0.000",0,"<0.000",0,4.0825343919059203E-5,6.6666872302688072,"<0.000",0,-9.7743487724956493E-5,13.333374460537614,-2.2160953596818799E-4,30.000088959931002,3.0440004505513901E-4,10.000025679667408,-0.51487172990338004,1316.7342162395714,4.0178733120950797E-3,40.000137686727385,0.18995279692808301,11287.37165797693,1.9855488339795799,79800.010204116348,0.237213785892985,11597.602366244757,0.20103154901252401,12208.129700897191,4.0606714830781203E-2,3447.0566299499574]
[null,false,"051SMPL.d","2019-10-22T12:39:20","Sample","QC2-2",false,false,false,0.704938386543706,200.00453442345267,0.60461058794330003,63.33354855576831,150.581516610791,62383.028041774611,3.1913002086678199,46.666808227695896,376.19701182890401,473.34084123361845,-8.9023254577374509,18559.732979132725,56.251809267874897,896.68992295826592,0,28008.562408160728,4.6648107431519001,554.00889078776038,2.7052281861173798E-3,386.6713021198218,2.7334788338699199E-2,2173.469205894035,-4.9341544908616003E-2,356.67069857313186,-0.12016771202649799,9202.4914573193164,1.84603823941218E-4,91.666957760846273,-1.8567023310443499E-2,496.6743649541184,-1.2365606943680401E-2,858.35881562339591,-0.105304081824557,600.01972268069164,4.7448088556275303E-3,11.6666819900272,0.58946258788505801,1075.0335024540598,"<0.001",0,-2.5520864190259501E-3,3.3333380520343074,9.9416565037832196,983.36396341476893,2.8197510515443901E-2,166.66976585991156,4.5993333721576004,10986.822752950249,"<0.000",0,2.59789010989418,20572.226662850975,0.69068760481784297,6419.5364440549938,"<0.000",0,0,272113.58969519153,0.56196853229407795,3105.2827499516393,2.9202925226477602E-2,822.24970642124788,"<0.000",-20.567260135299136,"<0.000",-2565.3412636914659,1.00164076808796,4375.568985905712,0.47512606879184399,2920.2550834379331,0.41932739529072899,92.666966756184891,-0.36329439199221703,600.02269496571091,-3.9413495726622698E-5,5.0000062833228762,1.26148491966328E-3,20.000033080577357,4.2400741980595802E-4,23.333418021598192,"<0.000",0,"<0.000",0,-2.83477971472861E-6,3.3333384493986,"<0.000",0,0.433219167351866,19378.284216449396,-3.4557417662253999E-5,46.666828890639117,0.41878369763303802,13816.003618995295,9.1307112455599704,98001.373447011807,0.53795003954929499,3490.3975920319917,3.2804794880078798E-4,26.666801770526146,-5.9010837187002804E-4,200.00213960803342,2.4329933978373699E-3,130.00084047514932,4.54339877296005E-2,2820.2607815425476,1.12414314160483E-4,16.666703770557486]
[null,false,"052SMPL.d","2019-10-22T12:40:53","Sample","QC3-2",false,false,false,"<0.005",0,0.62855404319181296,60.000214477376915,-296.41904436941002,48246.949150856097,1.9525249220734899,36.666791687407219,200.17686173347499,406.67181044809308,2.4833936787439201,18993.465593807872,-6.3590943021714796,236.66847512123275,0,26018.877379215595,-5.9698872312229499E-2,8.6666730244954433,9.7401640411938304E-3,403.33837543917139,4.1312705881027402E-2,2173.4693076192939,-3.1916488555316203E-2,380.00455925855204,-3.3590169774331299E-3,9402.6829595966792,5.4035511134293104E-4,93.333668907478412,-3.2727181055459402E-2,390.0050931671496,-1.64867834007262E-2,783.35550787597788,4.6464299743600203E-2,866.70988748567549,-2.6532315536437999E-3,1.666667833924276,3.4749487726388798E-3,43.333447923261211,"<0.001",0,2.0402157222646201E-2,5.0000035017728282,318.62301679265801,11323.694492722136,-9.8551044963291392E-3,83.33416568738096,1.6753134107673599E-2,56.666937619443672,"<0.000",0,5.9632157250182304E-3,116.66721329092165,2.2840376104703301E-2,225.00161153572381,"<0.000",0,0,273532.55451779731,-5.7404707630923403E-3,15.000028386711653,8.7346395854242791E-3,255.55960908541206,8.5488571819327601E-4,6.8191690994426972,"<0.000",-26.00005492567934,-5.6665202442374597E-3,43.333424876132234,-9.2764177251645401E-4,5.0000062833228762,"<0.005",0,33.841728693253103,51384.033334395041,3.5666764259081001E-4,8.3333447327214767,-1.5955814053952599E-3,15.000018452604337,2.7925894127242199E-4,16.666722446679238,"<0.000",0,"<0.000",0,4.0009296133950602E-5,6.6666872302688072,"<0.000",0,5.68720106045061E-3,273.33647638554658,-2.24807267400656E-4,30.000078231095102,3.4175038210154402E-3,113.33406597374778,2.6605452750054499E-2,6784.8133332116404,0.123530768207625,816.68923750486135,4.7704537063917199E-5,10.000015348195801,-3.33480607164631E-3,90.000415096674146,3.2360013374495798E-4,26.666751305260988,1.07797854691435E-3,130.00086908537841,"<0.000",0]
[null,false,"053SMPL.d","2019-10-22T12:42:21","Sample","NMS 59",false,false,false,1066.0979902619099,222408.19182016884,41.963181520291499,1990.1112577799274,-160.76501073614901,49155.961181581602,2685.1220200623402,13938.670039693467,197.44593418628301,390.00491514794658,8.0472091744805105,19060.215168053746,456.16334717629002,4457.2275133347384,0,24984.024692813186,3.3811079688186599E-2,18.00001335144043,2.5513202250309801,14906.224143114327,2.4540452764267902,24409.959759961173,2.4118927509558499,6357.8097604044815,-0.18464568662683201,8035.3893804513027,1.28653678533005,22704.443834854505,8.6202690506137006E-2,1050.033299376071,25.760848364571402,217397.994768078,-6.5999333842614194E-2,633.36201335057149,2.67738425212299,3373.6574814668675,5.94293313973056E-2,135.00058926641063,12.850193590936801,22708.046536623737,51.758871147536098,3910.4385810528415,4.28264701215919,753.35092057758186,128.61641917862099,266349.63160441985,25.0298938809695,57932.865347670537,4.7715397644139201E-4,10.000014156102923,-6.3966246420843103E-3,20.000047385691889,8.2774975791640604E-3,90.000297476843528,"<0.000",0,0,273268.53108423157,-6.0427004652145802E-3,13.333360155423085,4.9690419437856699,138493.15359469515,2.55704864500621E-3,21.213351102868781,1.26884163717301,15735.297843426997,-3.3416223834855101E-3,53.333497842150457,-9.2685700658181398E-4,5.0000062833228762,"<0.005",0.66666690508524573,4.8731796900789899,8369.1509845911442,9.0067545671894594E-2,756.68534740777795,12.3625403674039,22082.87646130408,9.910593523627261E-4,50.000213036931363,"<0.000",0,3.2108011845518899E-3,63.333540608482451,"<0.000",0,"<0.000",0,2.72211949359577E-3,140.00099609294045,2.6352157873009202E-4,73.333672384415976,"<0.000",0,-0.18475898239088701,4650.73659483748,3.4271838654837398E-2,236.66864042477846,1.18525671048891,70991.543538497732,12.1633655047833,491853.71662914177,1.3742392270927699,67707.603157744757,1.27737775909294,77887.451638865692,0.246158444793316,21037.598970369512]
[null,false,"054SMPL.d","2019-10-22T12:43:48","Sample","Reagent blank11",false,false,false,0.60929411412998402,166.67074734975083,0.27742890027916201,43.333518654105283,-276.97987812893001,49155.896077415899,5.5547202769597304,56.666856557127993,512.75713399519498,470.00707561760532,-0.34898946579707502,18743.243318520956,-4.1795597859799098,260.00238666928232,0,26279.188247732498,-2.95788282283268E-2,12.000010172526041,3.3343821018902403E-2,546.67573750005556,5.9037977699753899E-2,2353.4927017466321,-2.8347508976790701E-2,390.00552549949998,-0.21818680598010001,8102.1102173104673,-7.4102289880909099E-4,70.000197639065021,3.0199386895454601E-2,753.35079342100835,5.3019770883525903E-3,966.6981913156402,-0.107247553124661,566.68498987993826,1.3007795671271E-2,21.666704490780266,-9.1790426191992302E-3,21.666705682873147,7.4151228821325698E-3,23.333431531984136,2.1346417853512E-2,5.0000035017728282,6.9719080777642102,843.35455591448317,7.1402487152841601E-3,116.66845741375316,5.6585715556822704E-3,30.000128298995964,"<0.000",0,-6.3544481705795604E-3,20.000066459177933,2.9111457829189302E-3,39.817427182057095,1.8391031578468301E-4,1.666667833924276,0,264140.25648067013,-5.0121690543593802E-3,18.333361670374448,8.6611418697088202E-2,2344.6193795276376,2.16457531289913E-4,1.3641663845752861,"<0.000",-0.66667102773977938,-1.01005521397051E-2,23.333378682533226,-1.4595116257818399E-3,1.666667833924276,"<0.005",0,1.55580544403532,3333.7129066472658,3.9215449753577801E-4,8.3333419511714286,2.5672414399672901E-3,21.666706874966025,7.1082786378447895E-5,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.36995609148834E-3,76.666965931649514,-9.6298287296932004E-5,40.00010112921246,"<0.000",0,-0.384681923819434,2550.2428624833269,2.27781631156573E-2,156.66797404485979,8.0219860726490498E-4,53.333499431607635,6.1296864155832098E-3,456.67528107749496,3.0726365810407299E-2,1473.4042647470396,9.1112981142287094E-2,5427.6182256000502,2.7978682689768397E-4,30.000101675588365]
[null,false,"055SMPL.d","2019-10-22T12:45:17","CalStd","Blank control3",true,false,false,"<0.004",0,0.78067445416023296,40.000085234640757,2113.8075826409399,48293.620699639803,5.8697407802854897,33.333414693672239,2092.28685812514,413.33907465131466,122.018874927563,19760.912510375711,25.275822426967299,253.33549122012423,16.789927684734,27380.946043815027,0.13026997098266299,14.666678110758463,7.0066080000694003E-2,433.33909978460616,0.22574452766891301,2246.8105744300697,0.12965835835675099,340.0039749343598,1.3597532227772999,8502.1511706693982,5.6493818431130804E-3,105.00046891470051,5.5085837279854999E-2,326.6700505216412,7.5960260606831698E-2,666.68190856808712,0.22244254279942799,433.34320560119767,8.9858031289419306E-3,11.666684374212954,1.86456929372627E-2,31.666721825797527,3.7015002829832201E-3,6.6666856408116368,8.7131655300839397E-2,6.6666737198828594,26.975524187702799,896.6917285816113,4.0430893553099902E-2,83.334408079608465,1.1643545156889999E-2,26.666752099989576,"<0.000",0,2.64435675611399E-3,20.000037848948867,5.82437757862373E-3,51.666793053346971,"<0.000",0,3.7056140145440501,259536.07686177772,4.1731315588822001E-3,21.666708861787484,3.4003987120378297E-2,900.0286091539366,1.4716630472016E-4,1.1824993414307836,"<0.000",-13.000023966033901,7.3131058499069299E-3,30.000050415594625,1.7120978106384701E-3,10.000012566645752,"<0.005",0,0.66263577333549195,933.41320852617139,1.05183047319981E-3,8.3333447327214767,9.8320505104675607E-3,16.666686683892905,5.9927980729213503E-4,26.666736205417866,4.1647428209728402E-4,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,1.56406076390156E-3,66.666954954460934,3.9524787614735599E-4,33.333437740801209,"<0.000",0,0.22491686541486899,2150.1732988942654,3.4073310040638197E-2,210.00205044942027,1.7580955767890601E-4,10.000026474395993,5.6442292293266199E-3,216.6688632468055,4.7735695160255201E-3,223.33553299304546,3.2550438073356801E-2,1883.4537988386483,"<0.000",0]
[null,false,"056SMPL.d","2019-10-22T12:46:51","Sample","19286426-001",false,false,false,"<0.005",0,-0.57294383616963296,10.000014156102923,40.931436531421198,46251.747357670254,1.62733854132069,40.00019808609985,-80.844830297805203,373.3377528189954,-6.5100114602574504,17575.368638282813,-1.2050022898977599,226.66829725104128,0,25725.131452603167,2.6894468618009398E-2,16.666683197021484,2.9931103964862299E-3,426.67296727488991,2.1923000960916102E-2,2333.4892433858527,5.38488840290565E-2,456.67374248295408,0.23776311292553001,9502.9807625383819,6.5029074597298801E-4,111.66709415613968,1.7818699055409198E-2,413.33925267051774,-6.6097789649473401E-3,583.34432952130567,4.4715525326841803E-2,500.01269504456286,1.67770593172858E-3,13.333354592322985,-3.3973587764041498E-3,25.000040953357402,"<0.001",0,-1.97613077750061E-2,5.0000058859585836,-3.4601439211053302,760.01751085096339,2.57898369601091E-2,133.33619733657881,-5.69793584796643E-3,13.33335220813723,"<0.000",0,1.3878076489210901E-3,30.000090152023876,-2.4022953328407901E-3,30.000071873266421,"<0.000",0,0,258241.94927689512,-6.2432482713328395E-4,18.333368822931714,-1.16430346228544E-2,588.90579268374074,3.39957918179954E-4,3.8945860515845503,"<0.000",-13.666691973805051,-3.7167188677014198E-4,28.333379402756005,1.1556990611150599E-3,16.666692246993001,"<0.005",0,0.36007186909311001,1433.4403358423785,1.0623725972995601E-3,16.666689862807246,-9.3884738980283197E-4,15.000040705004722,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-4.6373139868205801E-4,46.666878958539975,-7.7465286641488395E-5,26.66672706603914,"<0.000",0,2.7415549145091402E-2,2400.2337674352125,-1.2874697473503401E-2,130.0007037818327,2.36468803726893E-4,23.333412855862388,-2.8516222490281998E-3,106.66716098784663,-2.69705155421897E-3,96.667240609716771,-1.20542352858952E-2,1180.050514438989,"<0.000",0]
[null,false,"057SMPL.d","2019-10-22T12:48:21","Sample","19286426-002",false,false,false,0.156829179004355,33.333787073701828,-8.0976355392201799E-2,33.333504498002362,83.501196534003199,46676.119096138289,-0.18798500070090099,30.000105251866998,49.137132311690003,393.33920210591151,-4.0807461533227203,17758.797088834577,-5.2449367247137104,186.66815161702803,0,25458.205187049432,-3.5074037146459403E-2,10.000008583068848,3.6176120880817E-3,426.67282740265881,1.7740591384102E-2,2276.8168300864572,4.5625476506253398E-2,433.33981186141847,0.24111093314125401,9469.2857951511251,1.8651842334056399E-5,100.00035136937572,1.6911652126535101E-2,406.67206476124034,-1.4481976750854301E-3,625.01344488342852,-6.1852101256193101E-2,300.00704005376338,3.53146926644907E-4,11.666684771577247,-5.4310921048946297E-3,21.666709656516066,-1.79084168705515E-3,3.3333380520343074,2.5219657376365202E-3,6.6666737198828594,-2.2935067509470102,800.01929044694771,8.9986436088454105E-3,100.00094796222579,1.80260514346303E-4,26.66677117347562,"<0.000",0,-4.2112953301473398E-4,16.666709333657582,-2.0577118017894298E-3,33.333402772743462,"<0.000",0,0,261412.54167756625,-1.9421906326296001E-3,11.666689142584465,-7.3295263004024697E-3,711.13295555163575,"<0.000",-0.24166869945820779,"<0.000",-16.666711558897624,7.5424643187473799E-4,33.333416680493706,8.3761477310425995E-4,15.000021631518678,"<0.005",0,-0.169250135344166,700.02613042850101,-6.3411865180481501E-4,3.333335667848552,-5.9274594723134396E-3,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-7.8764535778008899E-4,33.333416283129409,-8.1319424340593193E-5,26.666739781696503,"<0.000",0,-4.4892819626929897E-2,1733.4473043705666,-1.31317940817651E-2,130.00079597034858,-5.9444643140747902E-5,6.6666768987972,-2.1958049798510601E-3,133.33410064378231,-3.2173876675242398E-3,73.333700994645042,-2.23696536858316E-2,593.34605286499777,"<0.000",0]
[null,false,"058SMPL.d","2019-10-22T12:49:51","Sample","19286426-003",false,false,false,1.01775032914504,233.33922748777533,-0.32660887626479201,23.333375900983174,-69.898320304579102,46833.215447964038,-2.9434270314094801,16.666707744200412,-241.64099825780301,366.67103538904036,-6.8269236767893302,18709.843863021819,1.5843288584170501,270.0021639962668,0,27461.009519184488,5.4908382569827403E-3,15.333345095316568,-8.8327992314239708E-3,380.00459740552407,-2.3516849611446702E-2,2020.1170556726495,2.6086776345952901E-2,410.00519459438533,0.55216726910575098,12004.218051925209,-1.13925896188029E-4,103.333754589154,1.0377747991117899E-2,390.00492786360388,1.19171546311938E-2,775.01864576133039,-8.6247868384080301E-2,266.67229930572381,-5.1551215216098503E-3,5.0000058859585836,6.7286683389814602E-3,43.333440373339648,"<0.001",0,"<0.013",0,-5.5488055443014899,716.68270314709287,7.7643618598775598E-3,100.00143274668081,-2.9707286289431399E-3,20.000066459177933,"<0.000",0,-1.3318108901493699E-3,10.000014156102923,-2.2821267330364599E-3,31.666717057426023,"<0.000",0,0,261718.45313134114,-3.5364656894546799E-3,3.333335667848552,-2.0682520310518399E-2,355.56182596432609,"<0.000",-0.49792637400934092,"<0.000",-51.666896288590181,1.48462207435896E-2,91.667057896647975,1.1176054501460899E-3,16.666694631178757,6.2731138717565197E-3,1.3333341280619304,-0.16982196515511,700.03304456744945,-2.1738262489110399E-4,6.6666768987972,6.7430702525778097E-3,28.333386157948979,-4.9967187916331504E-6,26.66675686836108,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-7.8855266375631096E-4,33.333425422508135,-4.2490855386862898E-5,30.00010922550992,"<0.000",0,-5.3748918135840998E-2,1650.1008720470165,-1.8519630807241401E-2,96.66715954740107,-1.17695107384018E-4,3.3333384493986,-1.42483388297394E-3,163.33478684224508,1.35259474592843E-2,863.35778539342994,-2.5238018807107999E-2,426.67365392038738,"<0.000",0]
[null,false,"059SMPL.d","2019-10-22T12:51:21","Sample","19286426-004",false,false,false,2.0424807238025098,466.68387502470335,1.1068874639424999,96.667064179970851,-110.617407709334,45743.866505863269,-1.1716394990351899,26.666721900303333,-437.89653671064002,326.670253972159,-5.5059748262398802,18859.942539879386,-1.9840302647682599,233.33507190145443,0,27367.480874744841,7.1235351654538107E-2,22.666693051656086,1.7899057366056199E-2,543.34293827400199,-2.92974912495858E-4,2240.1462984063282,5.7587895346219303E-2,490.00713889786897,0.13700399442049699,9335.9618684243578,-1.24299602668996E-3,81.666943604743324,5.8158198707270597E-3,360.00433067474268,-1.3068436505789901E-2,550.01269616215006,1.8062742920763899E-2,466.68449491302289,1.32814042946602E-3,13.33335538705157,-1.07580070605872E-2,13.333345055579963,"<0.001",0,-4.3329478170536201E-2,3.333335667848552,-3.8784705711633398,763.35112387908805,4.0956392946013602E-2,166.66881417239537,-5.7818915642756201E-3,13.333361744880252,"<0.000",0,9.0807018952387204E-4,26.666723489760511,-3.5512096023245601E-3,20.000036656855993,"<0.000",0,0,257157.05345972799,-3.5251726393568401E-3,3.333335667848552,-1.6633085426632301E-2,455.56414880705546,1.8411683047797099E-4,2.6375005289291855,"<0.000",-30.00007584690935,4.9881732412727503E-3,50.000126411515588,1.5838558985527801E-5,10.000009785095704,"<0.005",0,-0.13719418307745601,733.37181458956582,4.3435525034208998E-4,11.666682784755784,3.0678765687231898E-3,21.666718001166213,-3.7247086552804601E-4,10.000025282303115,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-8.5372363181579101E-4,30.000066707530618,8.3439073865856301E-5,40.00026365120813,1.0691497317163099E-4,3.3333384493986,-0.13165732733739299,883.37561674510573,-1.6606183185292601E-2,106.66732152302085,-5.7519012417714199E-5,6.6666768987972,-3.10275139350303E-3,96.667145242286551,-2.40067049658015E-3,110.00046884019473,-2.6561867273503799E-2,343.33812202009375,"<0.000",0]
[null,false,"060SMPL.d","2019-10-22T12:52:55","Sample","19286426-005",false,false,false,"<0.004",0,0.758024478345102,76.666942089791959,-76.525089045549706,45262.559546889468,-2.8516995484908598,16.666699002185975,336.91746432616799,466.67315741370385,-13.386157681103301,17108.193843498193,5.1623464074139003,296.66943425929395,0,26626.371673548641,-2.5581558814900302E-3,14.000011761983236,5.6751805315864304E-3,456.67303040614183,3.12643629094331E-3,2223.4747491755875,9.51488627293187E-3,356.67085751884889,5.7133004210687698E-2,8668.9177861319495,-1.16435230757484E-3,81.666937246914657,1.54002521789946E-2,410.00544890753258,-1.5130845751356601E-3,641.68257558133007,3.8009951259526699E-3,433.36102341674069,7.9664295561296707E-3,21.66673747201655,-1.71851905972076E-3,28.333386555313272,"<0.001",0,-6.5081637156821595E-2,1.666667833924276,-6.2915723765259601,680.01382450242102,-1.5959757454041001E-2,50.00059517722665,-5.7774209288806097E-3,13.333409428595361,"<0.000",0,-8.7203489119341403E-4,13.333361744880252,-3.1858623664500802E-3,23.333370735247371,"<0.000",0,0,259589.23441306897,-2.5684105645734801E-3,8.3333447327214767,-1.80549792462441E-2,422.22999880728622,2.37524238011501E-4,3.0916688780238495,"<0.000",-16.666703134774618,8.10905485824674E-4,33.333398004371951,-3.5059592278733502E-7,10.000012566645752,"<0.005",0,-0.35506262231453101,433.34763223958186,2.10107487125978E-4,10.000012566645752,-6.8830392025981202E-3,5.0000062833228762,-3.7459592555642298E-4,10.000024884938822,"<0.000",0,"<0.000",0,4.5381868129514501E-5,3.3333384493986,"<0.000",0,-9.3856560740613302E-4,26.666737794875036,-2.7669786133657998E-4,10.000025679667408,"<0.000",0,-5.0601460994735802E-2,1666.7718306207323,-1.7851332277234899E-2,100.00044991572027,-5.8627317281723999E-5,6.6666768987972,-4.1683565989633501E-3,56.666929672157814,-4.2749444000563797E-3,23.333380669354685,-2.7136355397538101E-2,313.3372005819698,8.2144582851602599E-5,6.6666880249973923]
[null,false,"061SMPL.d","2019-10-22T12:54:24","Sample","19286426-006",false,false,false,0.760236281691402,166.67074734975083,0.91521662874446097,83.333810518178197,-68.3183416208756,44818.142235803469,1.47485627517565,40.000110665955482,212.538722994247,436.67229319603803,-7.3559174801228302,17808.832693794106,-4.1220054569305598,203.33484982448545,0,26259.157323551815,-5.01463943678317E-2,8.6666707992553711,9.4549932567033999E-3,473.34146430082927,1.81266439490935E-2,2340.1584530846508,4.58915834169957E-3,340.00416566922019,0.15043216189403799,9135.8314013046092,-1.39321685610857E-3,76.666940500334789,8.4452036297394006E-2,803.35379357108786,1.5817713927441E-2,783.35363231644703,-0.117050101393971,200.0084603828096,-5.0392172138836603E-3,5.0000035017728282,1.3482600960083E-2,53.333459695178369,"<0.001",0,"<0.013",0,-6.5584408337568298,666.68098727840697,2.5208343102980101E-2,133.33619733657881,-1.3227340966136601E-3,23.333375900983174,"<0.000",0,3.1410657256581102E-3,43.33349958061924,-1.2856177567077401E-3,40.000083645183594,"<0.000",0,0,259064.01957714526,-3.5299417179133802E-3,3.3333384493986,-2.13871824222495E-2,333.3382998904703,2.5529175549324603E-4,3.2279181674930131,"<0.000",-17.000026449560721,-3.9369915405294403E-4,28.333377415934539,7.4357405820243798E-3,53.333467642464228,"<0.005",0,-9.3648646377815697E-2,800.03846685177632,6.3416564096460595E-4,13.333357771237326,-8.8470543911524802E-3,1.666667833924276,-5.2423341359052796E-4,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-5.4556899954364899E-4,43.333511501548024,-1.5766716522354999E-4,20.000050961970526,"<0.000",0,-0.106143926992859,1133.3872094762446,-2.4320399125142701E-2,60.00021845101984,1.1774017139406599E-4,16.666703770557486,-2.59950435967317E-3,116.66733567912377,-3.9884266637212002E-3,36.666767050821079,-2.8510512285248001E-2,233.33553602294819,"<0.000",0]
[null,false,"062SMPL.d","2019-10-22T12:55:52","Sample","19286426-007",false,false,false,0.74899526561985297,166.66984135912998,1.8258125133629299,130.00079279143424,-103.20958654142601,44714.569646200987,2.5723047731001101,46.666805048781555,-306.88266991675101,343.33753074202644,-7.4634995119967602,18059.051917323141,7.1817849800448998,316.66986629362106,0,26653.113209608495,-3.3138679060120399E-2,10.666674772898356,-1.2715552717453799E-2,346.67137900981237,0.10976497124102801,3270.3015666658448,0.129420405484486,666.68009718239148,1.28737284007,16274.559384760845,9.9505544323867895E-4,121.66716394324357,0.16699504542369401,1300.0489867719662,5.1720099920434401E-2,1108.3710600319989,0.20922609833151101,833.37137177671582,1.3569124606410601E-3,13.333359758058791,9.6832743520774297E-5,31.666726594169045,1.8689053982021901E-3,10.000023692845945,-6.5319091741921104E-2,1.666667833924276,-7.1306696230444198,660.01376122215743,4.0084554613012503E-2,166.66881417239537,-7.1440550747193396E-5,26.666809320447701,"<0.000",0,4.3381278296259797E-3,53.333639303838616,-1.0253156790188201E-3,43.150783512848861,1.83567181576315E-4,1.666667833924276,0,264497.80855867552,-3.85814368163647E-3,1.666667833924276,-2.5765592634744101E-2,222.22524715801219,1.3683816885506901E-3,12.410422271732308,"<0.000",-35.000100567935391,6.6400862750498196E-3,58.333500946558992,1.0878698604444401E-3,16.666689068078664,"<0.005",0,-0.47687390527718498,266.67357087150748,5.9952868413603796E-4,13.333356181780156,1.62165267060361E-2,45.000104233620995,-8.4746786561950099E-5,23.333388616640537,-7.8126683078602196E-6,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,-1.4105891474400999E-3,6.6666768987972,2.2528660872432501E-4,53.333590825394928,"<0.000",0,-0.105162127388673,1166.7249622457341,-2.7174251015931301E-2,43.3334646125615,-1.1830577676428201E-4,3.3333384493986,2.79112276931747E-3,330.00415946040312,-3.86473584403366E-3,43.333455870547063,-2.9441427131799501E-2,183.33472296593504,"<0.000",0]
[null,false,"063SMPL.d","2019-10-22T12:57:21","Sample","19286426-008",false,false,false,0.46082358635628301,100.00136122110548,0.52144339087492697,63.333573987083035,-14.789031327780799,45523.259412796404,-4.0147260855427902,10.00002289811736,698.37702405646098,523.34192131943621,-7.1540635054894999,17658.741137704925,0.64995578515212304,246.66941275195163,0,25992.090270239758,-6.1804741050557201E-2,7.3333383401234942,1.16085853002375E-3,420.00541061154883,2.6288091986637299E-2,2396.8332060650009,4.6898611475156997E-2,443.33950653663015,0.151741492396407,9069.3201184865866,-1.8262637327984899E-3,68.333511923747579,7.4400901612495698E-2,740.01721868886716,1.0559756196997901E-2,733.35216889820867,-8.0997837394790301E-2,266.67136152595839,-7.66214461864702E-3,1.666667833924276,4.5799545059044197E-3,38.333389982580293,"<0.001",0,2.1579627024928402E-3,6.6666784882543704,-4.92646086507607,716.68286845063858,1.7220836707189401E-2,116.66748983166468,-7.2052664729299601E-3,10.000014156102923,"<0.000",0,-2.19815477903161E-3,3.3333380520343074,-3.1718489612472698E-3,23.333371927340252,"<0.000",0,0,258879.87895488917,-1.59854242777433E-3,13.333352605501522,-2.8111974740462502E-2,155.55740280287952,"<0.000",-0.37791892481964096,"<0.000",-23.000054970382823,2.0551110479311301E-3,38.33342495063804,-8.5387902830085698E-4,5.0000062833228762,"<0.005",0,-0.37793879258756102,400.02071956839774,-8.4093129215789602E-4,1.666667833924276,"<0.001",0,-3.7398027037874798E-4,10.000024884938822,"<0.000",0,1.7838173522191801E-4,3.3333384493986,"<0.000",0,"<0.000",0,-1.40725843031943E-3,6.6666872302688072,3.9725217750341801E-4,66.666924754774698,1.0620346232599E-4,3.3333384493986,-0.12179158781199401,983.37040556056706,-2.4855675516584502E-2,56.666874041156852,1.17949161696387E-4,16.666714499393382,-3.9902052038234097E-3,63.333608955140768,-4.2021499017629504E-3,26.666761636732591,-2.9142940309322199E-2,196.66825955110403,"<0.000",0]
[null,false,"064SMPL.d","2019-10-22T12:58:55","Sample","19286426-009",false,false,false,0.46561168732029201,100.00226721172633,1.1581586814774001,93.333789706223357,-65.699115471445594,43962.74348552907,2.25211443064756,43.333487659690462,242.435200399211,433.33916336289303,-2.6714055283828202,18159.367267872654,-1.91285326688747,220.00183095531915,0,25725.034610156697,-4.4559112967322801E-3,13.333345254262289,-8.3420505805847697E-3,360.00388562673498,1.24808468170487E-2,2240.1437044122263,4.4925699463072399E-2,433.33913793157831,0.17542021075971501,9102.5430806549575,-1.2130009859051701E-3,78.333575353022795,6.8352091007377394E-2,696.6820549465906,-5.25463044588108E-3,591.67917302527178,2.8355914886362198E-2,466.67886823443035,-9.3653715115524596E-4,10.000016540288678,5.9246915214927996E-3,40.000083645183587,1.5316953547642401E-4,6.6666856408116368,"<0.014",0,-7.6194926854637002,620.01198162617311,4.3193307897595699E-2,166.66976983355462,-7.1358314078349E-3,10.000014156102923,"<0.000",0,1.8901574316485001E-3,33.333428204058187,-4.1326917579419401E-3,14.634683366457374,3.81698025089019E-4,3.3333384493986,0,254369.47537626614,-3.1905413780560698E-3,5.0000062833228762,-2.1582620097078301E-2,322.22788789218743,2.33883908944823E-4,3.000834130235047,"<0.000",-22.000045448540963,1.80668637162259E-3,36.666742414234939,-8.3866134913015001E-4,5.0000062833228762,"<0.005",0,-0.22802483474938001,600.02172539680089,-1.93274985613688E-4,6.666674117247152,-7.8257008605326101E-3,3.333335667848552,-3.69985326268206E-4,10.000024884938822,"<0.000",0,"<0.000",0,4.6313121440964803E-5,3.3333384493986,"<0.000",0,-1.4842695346867199E-3,3.3333384493986,3.7097764478513998E-4,63.333637565369848,1.0808663040592001E-4,3.3333384493986,-0.116404601861646,1016.7110034584975,-2.35885744185385E-2,63.333567629254354,6.3364371585016105E-5,13.333364923794592,-3.9608767021449302E-3,63.33360299467639,-4.4101000998318998E-3,16.66670337319319,-2.9670300293798899E-2,163.33477412658769,8.3830221146173099E-5,6.6666880249973923]
[null,false,"065SMPL.d","2019-10-22T13:00:23","Sample","19286426-010",false,false,false,"<0.004",0,10.2138216125671,550.00923608726407,-140.24662886139001,44022.89159661196,0.74340012927380394,36.66674559314928,171.66428053306501,436.67264923444418,-5.8016203597834597,18376.140242281712,-6.8903679597884293E-2,246.6687928636552,0,26733.274341226508,-3.04198032815289E-3,14.00001049041748,-6.7153493220310197E-3,383.33768909169703,7.1431657746650898E-3,2270.1484850523598,1.9447536971464702E-2,383.33859826119846,0.27038052447655803,10003.2038063527,-2.3633262693556001E-3,60.000161230561709,3.1797252235001502E-2,506.67451739299526,4.1719434085151802E-2,1016.6995911555895,9.0138647022067406E-2,600.02253601998791,6.6320177471237696E-3,20.00002433856292,-5.7239164357813302E-3,21.666690980394321,-1.8287588486283399E-3,3.3333380520343074,-4.3090787823456697E-2,3.3333380520343074,-4.2896663011826703,746.68429215722847,3.2364998052594702E-4,83.334406092786935,-1.38360005612561E-3,23.333385437726196,"<0.000",0,3.1147030601966601E-3,43.333490043876218,-4.5048439061731197E-3,11.66668040057003,"<0.000",0,0,259387.8411869033,-2.24597146022117E-3,10.000012566645752,-2.2663037988181201E-2,300.00448897490998,1.7379552802921001E-4,2.5774947704251718,"<0.000",-46.000162392852268,1.1386657491701899E-2,76.666937321420448,2.2850824995313602E-3,23.33337669571176,"<0.005",0,-0.49689248921856799,233.33875859789259,-8.4134429848171596E-4,1.666667833924276,4.9244806163176403E-3,25.000075921415149,-5.2432710203754096E-4,3.3333384493986,4.1695149004875802E-4,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,-1.2510702324114E-3,13.333374460537614,7.9320341402283904E-5,40.00010112921246,"<0.000",0,-0.11327092825150301,1066.7097012211564,-3.0826403366021399E-2,20.000041425227504,1.0047212819483201E-7,10.000026474395993,2.6966438386480801E-3,320.00388622278143,-4.7022820859217397E-3,3.3333384493986,-3.01871570323912E-2,136.66754201053269,1.23312473399421E-4,10.000026474395993]
[null,false,"066SMPL.d","2019-10-22T13:01:53","Sample","19286426-011",false,false,false,0.60282224154328401,133.33605428542816,1.16920563161325,96.66710868477162,-119.914427325017,44076.395827585518,4.4433468856183298,56.66690185665734,209.65166012568801,440.00623683125423,0.68488175681207497,19227.273989333375,-2.58991605464356,220.00206301606599,0,26492.828381267766,-1.4170938953389599E-2,12.666675567626953,-1.1755985788824599E-2,350.0040892262644,1.85637197616196E-2,2363.4989577010429,7.4047949742315697E-2,520.00894407417957,-8.7941538459466208E-3,8235.3995574479213,1.36635885334298E-4,105.00037672618464,6.5632803826223399E-2,700.01555353379911,-9.0579966267413797E-3,575.01136157687063,0.161416852211101,733.37371399095503,2.77752740731156E-3,15.000018055240043,-2.6815317711616298E-3,26.666733821232114,5.6476526811472198E-5,6.6666856408116368,-4.2964287688720099E-2,3.3333380520343074,-7.0775172754493099,653.34831535793865,3.8482471663735902E-4,83.334408079608465,1.5616487586259901E-3,30.000071078537832,"<0.000",0,2.6724956583619001E-3,40.000094771383779,-4.3380951191996702E-3,13.150684618810544,1.8671439245234001E-4,1.666667833924276,0,259921.22659562191,-3.2115294110476702E-3,5.0000035017728282,-2.68780736819757E-2,188.89089582699387,9.0544610847182599E-4,8.4704214997527725,"<0.000",-17.333365738391393,3.9484577898933398E-4,31.666722620526116,3.4165884962744199E-3,30.000059952337647,"<0.005",0,-0.47360378818073001,266.67258540802516,2.0849629006687901E-4,10.000017732381556,-2.95981062606192E-3,11.666685566305832,-5.2448091278955403E-4,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.4859738243281999E-3,3.3333384493986,7.8346614465408498E-5,40.000112652776949,7.40446754546393E-4,23.333399742840729,-0.11698127106152,1033.3802136292604,-2.7052773463793499E-2,43.333476136125974,-5.8776991812948103E-5,6.6666768987972,-1.91590295206914E-3,143.33418632545792,-4.1332983406930602E-3,30.000088562566713,-3.0134483342374899E-2,140.00093569356795,4.1019762150678202E-5,3.3333384493986]
[null,false,"067SMPL.d","2019-10-22T13:03:23","Sample","19286426-012",false,false,false,0.30650675582356002,66.667574147403656,0.92868399816966296,83.333618193860573,-137.994185808997,42950.318370407047,-1.55136221961427,23.333384642997615,372.72708314184803,463.34103211749044,-7.3655836208841698,17666.993599640682,1.6387204889060301,256.66928226745205,0,26052.211525854294,-3.7082617181093999E-2,10.000006357828775,1.0091096720182401E-2,473.34044704824026,1.51817196016197E-2,2293.4824283895528,5.60884524371753E-2,466.67389492183088,-5.9650714707744697E-2,7801.8444533561669,-6.1246509982186396E-4,90.000327676529764,9.5511707858797895E-2,860.02305338745737,-4.1036977825444597E-3,608.34668247679906,0.238011531201237,866.71924938875748,-1.0276717413783901E-3,10.000023692845945,4.6371624734710204E-3,38.333415413895018,"<0.001",0,-6.4735340818618106E-2,1.666667833924276,-8.2410636601880896,606.67910625518709,-7.3361568266781E-3,66.667379438893036,-2.7226367767162899E-3,20.000047385691889,"<0.000",0,3.1900915077375299E-3,43.333461433647152,-3.7263704282824799E-3,18.333376372853273,"<0.000",0,0,256852.9519017594,-1.2538626892640301E-3,15.000018452604337,-2.5520384489605499E-2,222.2253584200183,3.7114857279018298E-4,4.1216681567579663,"<0.000",-12.000023186206471,-2.3867715003163701E-3,20.000038643677453,1.4595377479449201E-3,18.333360478281566,"<0.005",0,-0.30396242801794998,500.01725678681163,1.0738059391357899E-3,16.666691849628709,-4.8646725375306602E-3,8.3333419511714286,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.1689610818715801E-3,16.666733970243722,6.0319574796144804E-4,83.333751708262881,"<0.000",0,-0.142118678402917,783.35585755656848,-2.5329406558501699E-2,53.333660364146134,-1.16594259886363E-4,3.3333384493986,-2.8365181767424901E-3,106.66724999744817,-4.2696324923113997E-3,23.333380669354685,-3.0862355259846801E-2,96.667103916400109,"<0.000",0]
[null,false,"068SMPL.d","2019-10-22T13:04:56","Sample","19286426-013",false,false,false,0.43469831968887801,100.00226721172633,0.124377366545111,46.666778822738252,-293.66738732201901,41847.795600377583,-1.7867934039626301,23.333403716483655,305.40119640381999,476.67386934150454,-13.872003444314,17625.334561437397,-5.1164152459869596,203.3349006871149,0,27554.458411021158,-2.4363384793711099E-2,12.000008900960287,3.2986134611166398E-3,456.67352631677903,1.01833432745887E-2,2363.4937697128389,3.7036675935696498E-2,440.00664373228983,-3.0181144145649399E-2,8368.7504413693005,-7.5028166897307897E-4,91.667069022848182,2.1962189110827399E-2,460.00665614992403,1.3630372598170201E-2,791.68858708241896,-0.171461784153807,100.00140890482237,-1.3369141989615699E-3,10.00001216928146,-3.05349037226412E-3,26.666709184645971,-2.6013694010272898E-5,6.6666856408116368,"<0.013",0,-5.5692928610012604,716.68285573498122,1.5761346297465399E-2,116.66773222389219,-4.4195631586534702E-3,16.66669979691456,"<0.000",0,8.5527346599234197E-4,26.666733026503525,-5.2649764935650597E-3,5.0000035017728282,"<0.000",0,0,261549.39063386037,-3.53605362439056E-3,3.3333384493986,-2.77555238602476E-2,166.669322136435,"<0.000",-0.37791918469588814,"<0.000",-19.666722243030037,1.9594963688157399E-3,38.333434487381062,-2.9633207534335602E-4,8.3333447327214767,"<0.005",0,-0.192993489986514,666.68737630095677,-4.2559022808995903E-4,5.0000062833228762,-6.9051402782163603E-3,5.0000062833228762,-3.7627979684141602E-4,10.000024884938822,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.33125799534978E-3,10.000025679667408,-4.22632300702026E-5,30.00007584690935,1.0511949351100099E-4,3.3333384493986,-0.105539196124048,1150.0865239669351,-2.76331461374042E-2,40.000194907185509,-5.9505333970677597E-5,6.6666880249973923,-4.0932527085190702E-3,60.000250240163247,-4.0665853382798097E-3,33.333448469637112,-3.0892666194180601E-2,96.667142063372197,8.1528958343013002E-5,6.6666880249973923]
[null,false,"069SMPL.d","2019-10-22T13:06:26","Sample","19286426-014",false,false,false,0.59091761551120903,133.33786626666981,0.40567842540740301,60.000234345591544,-197.24843853780001,43220.984056477981,7.6889489317300899E-2,33.333424230415261,27.403493771698798,413.33841343713181,-9.5696283910300703,17975.651443453786,-0.67969646774262704,243.33554858959394,0,27026.920170184178,7.5711078451230797E-3,15.333346048990885,-7.9393383566158694E-3,380.00508060050385,1.6167176141092501E-2,2383.4940237279629,3.32647515543349E-2,423.33930338413558,-1.5655002084477E-2,8335.5435009298017,3.1219012410213802E-4,110.00043069322261,5.5221185135233199E-2,650.01332903881882,2.61875165560186E-2,891.69167286461845,3.5378648125250399E-2,500.02020522997236,3.8968579773675802E-3,16.666699399550268,-3.8863521380381998E-3,25.000064795214957,5.5683980392107297E-3,16.666728407143626,7.3399429065673595E-5,6.6666737198828594,-10.6308025157174,543.34258223559584,-1.6216628963355701E-2,50.000597164048195,-8.7405197784197908E-3,6.6666856408116368,"<0.000",0,2.1821726454720699E-3,36.667017390325405,-3.2306397667641001E-3,23.150702748556391,1.8527873146247301E-4,1.666667833924276,0,261803.90234881788,-2.5819853207174298E-3,8.3333447327214767,-2.2351531975500501E-2,311.11596061134583,9.4186509979829494E-5,1.9562451158527021,"<0.000",-45.000144287941687,1.08113155306303E-2,75.000240479902814,-1.48286242197377E-5,10.000024884938822,"<0.005",0,-0.24037063448790899,600.01578082676235,-2.17654678946264E-4,6.6666792829829555,-6.9079856633922597E-3,5.0000062833228762,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.02138221994895E-3,23.333400140205026,2.7085362540903998E-4,56.666886756814222,"<0.000",0,-0.131585624671298,900.02974296676041,-2.8711733310004699E-2,33.333438535529787,-1.17714075146749E-4,3.3333384493986,-2.2009585021870999E-3,133.3341546853261,-4.2791621734742302E-3,23.333391398190585,-3.0494511517217101E-2,120.00079135098868,"<0.000",0]
[null,false,"070SMPL.d","2019-10-22T13:07:55","Sample","19286426-015",false,false,false,0.30080649938221699,66.668480138024492,0.15875044381682801,46.66678677002411,-204.93133536206099,42282.141297030328,-1.6317325486530201,23.333375106254593,100.59176846186401,420.00524530800311,-5.6208351338937499,18275.973919242035,4.2251566134132101,286.66987309848463,0,26546.259370054915,9.7936771384070608E-3,15.333345095316568,-9.8048729080406306E-4,416.67249694457905,1.65081460275637E-3,2213.4794795491375,2.4565700201501599E-2,396.67235101254255,-2.62044895718423E-2,8202.0008648882504,-7.4062404098164901E-4,90.000364234044696,7.8441641903820802E-2,783.35183565787736,-1.30750133674061E-3,650.01565478777127,-0.20524360227247401,33.333802968274121,-7.6991196980278499E-3,1.666667833924276,9.7236601682690002E-4,33.333405554293506,-1.85673434102852E-3,3.3333380520343074,-6.5479132611223997E-2,1.666667833924276,-10.6730429892161,546.67642414555314,-1.6515108930045201E-2,50.00059517722665,-7.35058928865306E-3,10.000023692845945,"<0.000",0,-2.21336983983548E-3,3.3333380520343074,-2.9003628553673998E-3,26.666763226189762,"<0.000",0,0,268952.17698188359,-3.5535889047850302E-3,3.3333384493986,-2.9142811851389401E-2,133.33528074957644,"<0.000",0.47041582837082396,"<0.000",-12.333356673518473,1.3602698777993901E-4,31.666726594169045,-5.9941475554245899E-5,10.000010182459997,"<0.005",0,-0.50278588958024995,233.33974406137494,-4.4282720601333998E-4,5.0000062833228762,-5.0881331097740396E-3,8.3333498984572802,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.33766578567205E-3,10.000025679667408,1.00585197709172E-4,43.333461433647152,"<0.000",0,-0.10546181462802,1183.4046946233568,-2.88542361832979E-2,33.333428204058187,"<0.000",0,-3.8007687558300699E-3,73.333602448300482,-4.4985596809511003E-3,13.333376049994785,-3.1383025500431101E-2,70.000288238123744,"<0.000",0]
[null,false,"071SMPL.d","2019-10-22T13:09:24","Sample","19286426-016",false,false,false,"<0.004",0,0.635029569377772,73.333630263800956,-281.47544405334799,42322.204772995836,-1.80552343609646,23.333375900983174,210.951678610931,460.00645269940622,-14.877706407946301,17541.879664552504,-6.5247564934269304,190.0012560188583,0,27681.410720327334,-4.2541666244788903E-2,10.000009218851725,-1.10616304029982E-2,370.00502988688612,1.4064323829496701E-2,2423.4999231969327,1.7855277804864601E-2,393.33826114726668,-0.233899305234994,7168.3555933713933,-2.2198022850435399E-3,65.000165924427407,5.9960150254132601E-2,696.68255085722785,7.60773507220698E-3,750.02107707209814,-2.1925525194049299E-2,400.00938673835117,-1.4746745869969799E-3,10.000011771917167,-4.8729080631059298E-4,31.666742091376449,"<0.001",0,-6.5975713857224505E-2,1.666667833924276,-7.9120929694571203,653.34676404774029,-3.2609038406376599E-2,16.666784261666383,-3.20748706830699E-3,20.000056922434911,"<0.000",0,3.2812152043393402E-4,23.333385437726196,-4.0188229122295697E-3,16.666688670714368,"<0.000",0,8.8817841970012504E-16,271108.38137683697,"<0.000",0,-2.5162721550133198E-2,244.44779041748109,4.5402357792331903E-6,1.2733331212401577,"<0.000",-16.000027418136185,-1.0900436350563499E-3,26.666712363560311,-6.1941935431882402E-4,6.6666792829829555,"<0.005",0,-0.39078558534653601,400.01281996596703,"<0.000",0,-8.8908141951543809E-3,1.666667833924276,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.41433109146785E-3,6.6666872302688072,2.0966242312382801E-5,36.666763474542449,"<0.000",0,-0.13311886267510201,916.7055970687428,-2.6306951503706499E-2,50.000213036931363,-7.5044590203584304E-6,10.000026474395993,-4.3141937560472396E-3,53.333525657650945,-4.3643356385051303E-3,20.0000628828993,-3.0895967843570599E-2,100.0005293885788,"<0.000",0]
[null,false,"072SMPL.d","2019-10-22T13:10:59","Sample","19286426-017",false,false,false,0.15100878618063801,33.333787073701828,3.8007039587851601,226.66846255458699,-216.149891022719,41864.525743357452,3.8562853550188301,53.333526452379523,39.553394273198997,406.67163242889001,2.3176845352248798,19443.874539422173,-0.82201875963400906,236.6688756644397,0,26439.451103417308,-1.39906348965352E-2,12.666675567626953,1.38810936746625E-2,503.34138755985015,9.0615640025609508E-3,2270.1479255634358,6.7551678483896593E-2,503.34108238407345,0.57605679374248497,11804.047911953012,-1.4299502882570699E-4,100.00034183263269,2.4147871379164899E-2,460.00668158123864,1.9872115481993999E-2,825.02103106476181,-1.35098821531237E-2,400.01126229788207,-7.6820959344048902E-3,1.666667833924276,-7.7031870752383696E-3,18.333358888824399,-1.8289706628246299E-3,3.3333380520343074,"<0.013",0,-6.4558656718680698,676.68154661831932,5.7018372520199201E-2,200.00262707477722,-7.2665642285394396E-3,10.000042766331989,"<0.000",0,-3.1708037328257798E-6,20.000085532663977,-2.6456653660070699E-3,28.333388144770442,"<0.000",0,0,262178.68098906655,"<0.000",0,-2.94328134863951E-2,122.22325015401491,-1.82692996761087E-5,1.046248835530178,"<0.000",-21.000067636369646,1.1328669675700299E-3,35.000112727282747,2.6521633631717602E-4,11.666680003205736,"<0.005",0,-0.24096525023062201,600.02859185203249,-1.06018213960042E-5,8.3333447327214767,-3.0189859111316802E-3,11.666680797934321,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.3318170073350899E-3,10.000015745560093,-3.9841324650426304E-6,33.333414693672239,2.0973468914928101E-4,6.6666872302688072,-0.114458656104776,1066.7166630438217,-3.03255737462029E-2,23.333391795554878,-1.17797121357317E-4,3.3333384493986,-3.0654718219375E-3,100.00045786300613,-3.7861658646661102E-3,46.666898826754597,-3.0953670825216099E-2,93.333894610396598,4.0666566986465102E-5,3.3333384493986]
[null,false,"073SMPL.d","2019-10-22T13:12:28","Sample","19286426-018",false,false,false,2.9386822975193398,666.70563121723796,2.75914586945107,180.00111629563892,-322.54507994164101,40615.011601820413,0.63618521778449699,36.666802018878826,628.01729213575197,533.34193388608185,-15.263643987612101,17158.21775005251,2.8735437542697202,280.00281234578074,0,27173.869223984046,-1.7027802606029398E-2,12.666677157084147,-9.9127491224868195E-3,370.00522697957513,3.7332276609676203E-2,2606.8615842829909,2.2118738175292401E-2,396.6715181369853,0.345349552024107,10636.843855509149,2.7661583879939498E-4,110.00048155585206,2.0817397414501199E-2,450.00660543630619,-1.5238660115122499E-2,533.34505955404427,-0.102878566737324,233.339426169929,2.5351845452156502E-3,15.000022426247263,-2.0258479395680899E-3,28.333384171127516,-1.7680636049779201E-5,6.6666856408116368,"<0.013",0,-6.4570358989267298,686.68098698038375,-1.6382874220407399E-2,50.000352784999144,8.5364821169174009E-3,46.666866242882612,1.5993468005425101E-4,3.3333380520343074,-9.0176074144511997E-4,13.333380818366296,-2.8584952436674599E-3,26.666705608367337,"<0.000",0,0,263793.21221500816,-2.9098141699986902E-3,6.666674117247152,-2.90477505429944E-2,133.33506617285045,3.0388719823007702E-4,3.683718737705834,"<0.000",-51.000789836037313,1.3073252113581599E-2,85.001316393395527,2.5311342833586901E-4,11.666674837469932,"<0.005",0,-1.0723303480283299E-2,933.37084949100426,-2.2394598320507001E-4,6.666674117247152,-2.09334904399994E-3,13.333345850308548,-2.30773996752691E-4,16.666731586057967,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.2562974508641201E-3,13.333363731701716,7.1396093048285106E-5,40.000194112456931,"<0.000",0,-0.123712070589236,983.37329837272489,-2.8752161059000299E-2,33.33346873521603,2.27792455824306E-4,23.333391795554878,1.2758544103096501E-3,270.00330840542955,-4.4230844225651798E-3,16.666714896757679,-3.1133480396335001E-2,83.333743760977029,"<0.000",0]
[null,false,"074SMPL.d","2019-10-22T13:13:57","Sample","19286426-019",false,false,false,"<0.005",0,3.2698977133990499,200.00166277088229,-312.154857909246,39666.299050853173,0.83042628447925904,36.666782150664197,551.64806084084603,503.3409679431573,-0.85159460733400705,18910.004084780903,0.26279384743744,246.6688945889141,0,26386.071384160645,-3.8337158296575399E-2,10.000006675720215,1.17295631586209E-2,490.0077746807371,3.2591362076641602E-2,2496.845309880694,3.8777390310254997E-2,430.00633234769606,-0.310808538607662,6401.3624986317327,3.0771056752912899E-4,108.33389756579352,6.5828673146180905E-2,703.34883595483245,-1.6093259283981899E-3,641.68162190699229,-0.18772675058887001,66.667605936548242,1.07373166636126E-4,11.666679605841443,-7.7485136326430103E-3,18.33336047828157,"<0.001",0,-4.3365991379207998E-2,3.3333380520343074,-7.1774967349878702,656.68150876937045,1.6024298238263401E-2,116.66845741375316,-1.0195360742298499E-2,3.3333380520343074,"<0.000",0,-4.62033521352003E-4,16.666709333657582,-3.4167980652586201E-3,21.666698132951584,"<0.000",0,0,265129.10412218515,-3.54465583170854E-3,3.333335667848552,-2.4963302450600099E-2,244.44854805876048,1.2237867821963101E-4,2.2125008382524167,"<0.000",-10.000016540288678,-3.33597688357486E-3,16.666694233814464,-3.1544742563929503E-4,8.3333447327214767,"<0.005",0,-0.407822158641582,366.67505130190517,-2.2188874848016301E-5,8.3333447327214767,-8.8695871295327595E-3,1.666667833924276,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-9.5163554995427802E-4,26.666727860767725,-4.7029020166671501E-5,30.000086973109536,1.0370019377001999E-4,3.3333384493986,-0.10032722846916101,1216.7266799771294,-2.6131815203176201E-2,50.000130385158506,-1.1844269829123599E-4,3.3333384493986,-3.7741865112580798E-3,73.333608806129163,-4.3551062143987497E-3,20.000074803828081,-3.0971440750570602E-2,93.33385328451017,"<0.000",0]
[null,false,"075SMPL.d","2019-10-22T13:15:26","Sample","19286426-020",false,false,false,0.59830338771091396,133.33696027604898,1.75517665699724,126.66753898062996,-337.378546130887,39566.060608336418,-0.45084237246483,30.000087767838124,227.00693756269399,446.67498876638729,-10.4874675869941,17608.720282908485,1.6747822344444501,263.33578988906061,0,26693.104257189669,2.7460097336352501E-2,17.333344141642254,1.98312818676614E-2,543.34273482348419,1.93549446665294E-2,2386.8305104946517,1.1783136551051799E-2,363.33740964525822,0.57598143015000403,11870.825460646127,1.26486757429352E-5,103.33375776806834,3.7542097904223699E-2,540.00924895193305,2.7265135865975301E-2,891.69351663500481,-3.14218014998872E-2,366.68684507454833,1.4241210033365799E-3,13.333356976508741,-6.7220779212731402E-3,20.000072419642326,"<0.001",0,-2.1115159003561101E-2,5.0000106543300946,-9.9755234699290405,560.01030405347092,5.7266408923935301E-2,200.00262310113416,-4.3246741370461398E-3,16.66674748062967,"<0.000",0,3.54755475143148E-3,46.666828095910525,-1.51783359453523E-3,38.150761334954282,1.8659414536072801E-4,1.666667833924276,0,260015.12633380605,"<0.000",0,-2.98137948572527E-2,111.11208581264343,7.9021004375504501E-4,7.5458365194127044,"<0.000",-8.3333425074814382,-3.25775722058702E-3,16.666683902342857,1.4209661679568299E-3,18.333362067738744,"<0.005",0,-0.426431472874088,333.34023902829034,-8.4185209528912005E-4,1.666667833924276,2.2554012260785E-2,55.000191504753758,"<0.000",0,-7.6730837842082296E-7,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,-1.4860020240075201E-3,3.3333384493986,7.8175843087171802E-5,40.000134507813037,"<0.000",0,-0.13268316324768101,883.3686072387236,-3.08342364472192E-2,20.000041822591793,5.8171436180499003E-5,13.333376049994785,1.6294584125523201E-4,223.33544716235826,-4.2046449701532299E-3,26.666752099989569,-3.1745413799778903E-2,46.666796306767118,4.1004948608940702E-5,3.3333384493986]
[null,false,"076SMPL.d","2019-10-22T13:17:00","Sample","Acid rinses ",false,false,false,17.233853730411798,133.33605428542816,116.475007902665,203.33474174139783,43427.172986949401,35213.843876314502,445.05743097929201,86.66705320278227,50255.8412883727,350.00410194192182,3382.83367384756,19210.402854516171,819.81668921268601,286.66966329013809,-7.1054273576010003E-15,926.69152468405866,4.6706234257749397,13.333341916402182,2.1482252092128702,266.66918675101016,16.752729919690701,2713.5412193307097,11.3801546491756,413.33857874067758,145.912494867698,11003.577414731675,0.55229758820255204,110.00051652390982,9.0741150657480496,516.6744790970115,4.95617569323257,383.34065340475712,1483.23879086582,23015.234889430587,0.165396870045177,1.666667833924276,9.4408598909078005,110.00046884019473,"<0.087",0,21.6393257885012,10.000016540288678,12855.992526432099,2430.1793562842504,11.052195864358,116.6684534401101,21.548564595905699,240.00225345290323,0.10447574547437399,10.000023692845945,0.39414249714639799,13.333361744880252,0.68221455115740604,25.000040953357402,"<0.029",0,0,953.36182763136685,0.69494189646901905,13.333357771237326,8.5372818944151003,833.35966277512489,"<0.034",-6.042242722676356,"<0.023",-485.01295336564948,55.295838018337498,833.35496296453039,1.9403251996505499,41.666743929186303,"<1.292",0,366.55762857018902,1900.1405504874017,0.285290612265656,8.3333447327214767,0.52548842602181201,3.333335667848552,2.1610767759464702,353.33841433120148,1.47350211574464,43.333597332235222,0.19375424403893399,13.333363334337422,2.4713899932923899E-2,6.6666768987972,"<0.011",0,"<0.006",0,3.18845282482947E-2,10.000025282303115,"<0.009",0,200.581469967278,7051.6214039539109,13.217170488912,300.00319351747839,7.9592670071809798E-2,16.666714896757679,0.70352540505800798,100.00046422083481,0.48011963389302198,83.333705614004941,0.31247045854595601,73.333693047359191,5.5917458101580601E-2,16.66670337319319]
[null,false,"077SMPL.d","2019-10-22T13:18:31","Sample","Base rinse",false,false,false,2.29447018916848,33.333787073701828,26.858786368691401,90.000370591873363,24683.478791399801,38908.000722308941,520.59801501121603,190.00124330320094,38264.809670021699,506.6743648051069,1648.8494909719,18225.94441779827,372.44801599731198,253.33555479841098,0,1740.0920862939147,0.59468357231837199,5.3333364327748614,0.314767804301745,160.00104347863231,3.59782011121627,2633.5328281889433,1.97456863735299,393.33834379903959,19.686809344137199,9669.4533919917576,7.7170671720803594E-2,116.6671671966637,0.91481899121501498,450.00659272064877,1.5804134664964,1175.0450079028931,35.613245490559599,5834.5096620683544,9.4906221309018796E-2,11.666689539948758,0.18946219938594,31.666740501919282,"<0.006",0,0.13872723953464,1.666667833924276,156.10810958857499,610.01177832466692,0.55478271189651895,133.33500524365664,0.180856398791855,50.000147074458795,"<0.000",0,0.565610954885857,533.34295113867108,0.15075259539018901,190.00921443091011,"<0.001",0,0,41814.287300095646,0.43815949295874002,370.00433052573106,-2.8792805188038499E-2,22.222333980933673,3.1665087808000598E-2,41.182615837517055,"<0.001",-415.67606338347895,1.07454049045922,715.01570589817004,1.5999195783714101E-2,16.666684697071442,"<0.029",0,4.4784484797936202,1166.7578957995302,5.4767484901668299E-3,8.3333423485357212,1.45784729619875E-2,6.666674117247152,7.9559071689342602E-4,10.000024884938822,"<0.001",0,"<0.000",0,1.4086914705116499E-3,16.666751056908303,"<0.000",0,2.7074944551889499E-2,196.67170867316364,4.0206156332712002E-3,60.00026454527778,1.31505204821905E-3,6.6666872302688072,12.974470486867901,20329.756618956253,3.6422574354642898E-2,70.000305722152603,2.3703898005241601E-3,23.333400537569318,0.916690058554814,5704.3633183186748,1.64530674048761E-2,160.00101168948891,0.35652585451211899,3627.1028105352902,1.0199321418556001E-3,13.333387573559271]
[null,false,"078SMPL.d","2019-10-22T13:20:00","Sample","2% Nitric acid rinses ",false,false,false,"<0.243",0,558.882567875419,516.6746825475293,92704.581841829204,39031.598539477614,157.01851703379199,16.666718870400604,103730.340090731,376.6718999047294,6364.2336699233101,18926.629553088711,916.09578688639704,170.0009257097901,-7.1054273576010003E-15,493.34178189424011,9.8744370368817798,11.333340644836426,1.7789387607233,80.000300804769481,44.420953152054103,2386.8283742642147,42.385206895276298,486.67500088599826,398.464774861298,9202.4756899036001,2.19015975768527,130.00074192880479,25.551014835127798,426.67311986277815,17.664731575243199,391.67238951983921,67.364676157157007,300.00704005376343,11.0802245094239,30.000066707530618,2.53610757328592,8.333341553807136,"<0.312",0,78.495726945744707,10.000042766331989,4794.8138462704901,250.0022151569195,23.247225615078801,66.667379438893036,2.19770859778165,6.6666856408116368,"<0.039",0,2.5770770592599899,23.33340451121224,3.0794777681040499,30.00007584690935,"<0.111",0,0,253.33565652366988,2.9556270646040499,15.000020439425802,"<0.039",0,"<0.127",0.19791522605348,"<0.087",-23.333381384610412,10.3983773735358,41.666748697557814,5.2602969088845501,30.00006154179481,"<4.861",0,459.96490495897802,633.3555045232165,0.21446285549329799,1.666667833924276,3.0119737715272099,5.0000035017728282,"<0.023",0,"<0.128",0,"<0.055",0,"<0.014",0,"<0.040",0,"<0.024",0,"<0.012",0,0.325583475939347,10.000025679667408,85.509174840706606,800.02503593818676,2.7363230712527602,16.666703373193194,0.11989906852672801,6.6666768987972,0.79498623992488804,30.000090152023876,0.36017955475052799,16.666724036136408,0.14450283034650299,10.000015348195801,"<0.013",0]
[null,false,"079SMPL.d","2019-10-22T13:21:28","Sample","Blank control4",false,false,false,0.15882922732186999,33.333787073701828,0.35311748788039199,53.333501021064798,-174.89496021349501,40668.516646596043,8.8354840455750896,76.666975468392536,94.823267182356602,396.67210941505272,5.6156938629900699,18976.847754894767,-3.1782012929093399,203.33495472865866,3.5527136788005001E-15,25137.624137138962,2.3736835251514099E-2,16.000017166137695,8.9651817322323902E-3,453.33986242602469,7.8579740419963601E-2,2823.5610541170681,3.5860825352282501E-2,406.67197575163874,0.70163607809190598,12137.714657285391,-1.2910611994337001E-3,76.666942089791959,8.6146350249957704E-2,796.68557569356392,5.5142158234920399E-2,1100.0360098985932,0.50775711061140105,1366.767595958979,-3.6682738333033899E-3,6.6666737198828594,5.6773969821747499E-4,31.666721825797527,"<0.001",0,"<0.013",0,-7.6322849844029399,633.34542920174056,4.0855988207369E-2,166.66905259097982,-2.9080569540014E-3,20.000047385691889,"<0.000",0,1.7213432370428001E-3,33.333456814287253,7.3345983860963901E-5,53.333513736722161,"<0.000",0,0,268456.82542676484,2.4068185286491899E-2,151.66750046112884,-3.27864587514336E-2,33.333604286120433,8.9960281973894103E-5,1.9708314637716153,"<0.000",-25.000073378283684,4.4703944655938299E-3,50.000132769344269,2.1897444219477299E-4,11.666678016384273,"<0.005",0,-0.182196509753662,700.02616221764549,1.6842413458975199E-4,10.000017732381556,-6.0299171879105804E-3,6.666674117247152,-5.2685914755960404E-4,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.41285244142556E-3,6.6666768987972,-1.31336442158325E-5,33.33346714575886,2.04829506032006E-4,6.6666768987972,0.38365706977439901,6017.85202761657,-3.0413218427988602E-2,23.333392192919174,"<0.000",0,-4.3010562961238903E-3,53.333551883694248,-3.6715004363840498E-3,53.333502610521975,-3.2216276840783903E-2,20.000042219956086,"<0.000",0]
[null,false,"080SMPL.d","2019-10-22T13:23:01","Sample","Reagent blank12",false,false,false,0.31105490555389198,66.668480138024492,-8.6795675984169499E-2,33.333398799100536,-238.15488465138699,40177.39604844556,-3.3655136948542199,13.333360950151667,-130.63977677222101,363.33828066778756,2.9596935662088302,18976.694658374417,0.61880390465107404,243.33547865347842,0,25671.632918244111,8.1563696261187594E-3,14.666675806045532,-5.5894143955219199E-3,376.67117511225979,3.2145320678656102E-2,2433.4998213226622,3.33128896609425E-2,406.67200118295347,0.38784644281178798,10436.535877806005,-5.9790813882741904E-4,90.000299066300713,7.1112522309077705E-2,720.01657048836489,6.6394881796540501E-3,700.01653010683447,0.41321354396378202,1200.0676423326352,5.5325121118085302E-3,18.333370809753177,-1.55428402087031E-3,28.333378210663124,"<0.001",0,-4.2817095654948699E-2,3.333335667848552,-12.034275836854899,490.0074822206177,8.3984738394850007E-3,100.00095193586888,-4.3451549310671404E-3,16.666728407143626,"<0.000",0,-4.49713108103763E-4,16.666690260171539,-4.3377964126229598E-3,13.333359758058791,"<0.000",0,0,265135.66011167079,5.8822470145388998E-3,53.333453337349688,-3.11274990253993E-2,77.778481526523009,"<0.000",0.72833114412928324,"<0.000",-23.000062917668679,3.8225890004373801E-3,46.666782001652592,2.19842782172648E-3,23.333372722068834,"<0.005",0,-0.38466160565230201,400.01281996596691,-8.4590738297188599E-4,1.666667833924276,-4.05740819868606E-3,10.000015348195801,-2.3263944011164199E-4,16.666749467451133,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.41095812534857E-3,6.6666872302688072,1.0772250730527E-4,43.333461433647152,1.0369762958314399E-4,3.3333384493986,-3.2060535787936897E-2,1883.4571623581512,"<0.000",0,5.36527597648513E-5,13.333385984102101,-4.8792304851299302E-3,30.000119954345823,-4.6340855055337403E-3,6.6666768987972,-3.1366219175852597E-2,70.000401089582823,"<0.000",0]
[null,false,"081SMPL.d","2019-10-22T13:24:30","Sample","QC1-3",false,false,false,197.81832298116299,40612.301526245254,6.3909094061313496,330.00358725582174,-93.138838817223601,41460.264681673078,480.43485257375602,2480.1724890842165,462.919005812107,453.33970983813634,576.77539152750001,101634.68448199869,495.46042329935898,4687.2830518988312,0,24590.144490349125,-5.6667022342811501E-3,12.666675408681234,0.52727764224069595,3353.6530058534991,0.52375590459027599,6807.972717563719,0.46748201089008901,1436.7278075705797,21.2276529828763,130275.45457203958,0.245412345030427,4327.1935390354165,1.1093954834330699,6437.8447527014505,5.1467899294425701,42960.404402622684,51.364981848890203,94686.564209508593,0.55212452854668503,690.01405323524193,1.9678021940767901E-3,33.333382904528833,2.7210905669524599,4697.2983105386547,9.30825986968709,691.68046556392608,-9.3503011528544597,566.67693247382442,24.964970910000801,50373.195502556257,4.1688376086970097E-2,120.00065465767203,"<0.000",0,-1.31243665713158E-3,10.000023692845945,1.1065925910642501E-3,61.666813567278581,"<0.000",0,0,264354.59241497627,-3.9122716261944201E-4,20.000025133291505,0.91214823426579605,25508.011773490507,0.234491333152404,1920.3523400163153,0.255479950718906,3059.6158337849006,2.2599717969611802E-3,40.000077287354905,5.2909141247443501E-4,13.333360552787376,"<0.005",0,-8.1808831343047603E-2,833.36439405947806,0.52629006440069603,4255.5431494780905,2.4451946963097102,4238.8741433788227,-5.2573532966646095E-4,3.3333384493986,"<0.000",0,3.4937499364128401E-4,6.6666768987972,"<0.000",0,"<0.000",0,-1.2569510132956899E-3,13.333363731701716,-7.2042718878817003E-6,33.333446085451357,7.2802878712691001E-4,23.333390603462004,-2.6353808023005298E-2,1933.4770316911074,-3.3011342913611803E-2,6.6666768987972,0.23708614773220299,13745.986123293847,2.3893493702582398,93645.161364995016,0.27422665197048202,13295.609339640398,0.20509947686583099,14006.348261341229,4.7516913150104402E-2,3927.1691309337452]
[null,false,"082SMPL.d","2019-10-22T13:26:00","Sample","QC2-3",false,false,false,0.628812112409744,133.33696027604898,0.341494073482354,53.333517710365093,339.31771029999697,51987.058144603303,2.3566869356857798,43.333459844189981,436.21890683952103,463.34062521645484,11.415445687568001,20044.661397740128,64.719726488485193,836.68689752588341,3.5527136788005001E-15,25398.007434040202,4.7694081948412803,514.00758870442712,-3.8459491928068E-3,383.33808963490395,4.7845010901305003E-2,2560.1889165214361,1.8397001297722801E-2,366.67088915898074,2.2989123845781299E-2,8201.9759421988911,9.1088650283184498E-5,101.66704026360749,0.12051383204180401,996.69795776360297,3.36700815424344E-2,925.03022036272284,9.6521990490367401E-2,600.02067635502942,1.6025377095565099E-3,13.333349823951474,0.59688824504159399,1018.3637848740251,"<0.001",0,"<0.013",0,-9.7031292257478601,566.67728851223058,2.4546343304715602E-2,133.33500524365664,4.9944486697752399,11463.800691806266,"<0.000",0,2.7001291034620798,20615.572609470695,0.66721199691865796,6072.7409983762946,"<0.000",0,0,267366.77987009118,0.58295027693464896,3140.2915995276235,-5.0711051121000101E-3,788.91075626064537,"<0.000",-25.456568789627312,"<0.000",-2754.7241309077917,1.0994816319977301,4677.3182489874225,0.50079650603003201,3023.6059119760107,0.42677206738677897,92.66695912679036,-0.24915296839933801,600.02069224960167,-4.39216016839656E-4,5.0000062833228762,"<0.001",0,-4.53847974537492E-4,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,0.40638711024189,17913.270810545422,2.18627872046952E-4,53.333581288651907,0.430540626568853,13956.09842485043,9.26817753715512,93490.816298666949,0.50778669339083604,3440.376840824234,1.6551144174156999E-4,20.000064074992178,-2.01972863267527E-3,143.33417360980056,-1.3847104987963601E-3,163.3344466984106,2.18659355341938E-2,3243.6757939506206,3.98774555988047E-5,3.3333384493986]
[null,false,"083SMPL.d","2019-10-22T13:27:30","Sample","QC3-3",false,false,false,0.62340573953096201,133.33696027604898,-0.36347979230731903,20.000045796234719,-251.597045216048,39806.536678971795,2.28595752219104,43.333456665275641,180.028757840168,420.00570307166822,-6.6895039659280302,17475.161625138062,3.51644514250034,270.00257725513114,0,25618.267033622658,-4.1980701647880397E-2,9.3333382606506348,8.6074580040670901E-4,413.33898564171312,3.3534530144258098E-2,2440.1721336418564,4.9848683556103598E-2,446.67399694511295,0.13543235247594301,8902.3648026617448,1.6897499830455801E-4,103.33372280001061,2.9916176200969601E-2,483.34127341695711,1.9994787516733302E-3,658.34854325934475,0.16745977183076799,733.35681011331906,-2.3606318026776199E-3,8.3333439379928915,1.3658815064521899E-2,53.333469231921391,-1.804094104184E-3,3.3333380520343074,-4.2618341914710797E-2,3.333335667848552,326.15001875227199,11527.142569117606,1.6824203017732998E-2,116.6686958323376,-4.3070494203753902E-3,16.666690260171539,6.4969009662943597E-4,13.333409428595361,3.0936337657803798E-3,43.333442360161108,1.1186342610217E-2,151.66744641958505,"<0.000",0,0,263419.58591849229,2.7849817391527401E-3,36.66677340864976,-2.61454858020146E-2,211.11375432881331,"<0.000",-1.1177118027550805,"<0.000",-93.004716758856432,3.1917526009620903E-2,163.34120507029886,-3.0638356180765999E-4,8.3333447327214767,"<0.005",0,29.5926558903009,43256.436106830653,-1.5507139584898201E-5,8.3333419511714286,-4.0197831498678498E-3,10.000024884938822,-4.5166854937430398E-4,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,7.9902623941116198E-3,413.33982487509911,3.83595006320479E-4,66.666974027946978,3.5487052654667099E-3,113.33405643700475,0.51043310887365201,7135.030778783751,0.116733141337615,943.36434548053637,5.5147615813231902E-5,13.333385984102101,-2.90651144510853E-3,106.66719277699003,-4.0716033705805298E-3,33.333511650559622,-2.9542216127766498E-2,176.66818037626874,1.21425400475665E-4,10.00004912416067]
[null,false,"084SMPL.d","2019-10-22T13:29:03","Sample","NMS 59",false,false,false,1004.5734810858301,209599.23971094974,41.632360625809198,1983.4435230973831,-219.24541445748699,39505.968649596442,2595.2052213636498,13481.657513951397,-465.42507930091199,293.33604375517302,-3.7927195578988799,17475.237410458765,399.64115920755899,3887.0879547731879,0,24990.675083339087,5.3507534096621798E-3,14.000008265177408,2.4037355889184702,14092.207228160791,2.3350088963965998,23582.208307451878,2.3693860213934701,6091.0541898351776,0.20513476884495899,9135.8451342150729,1.2414898028394701,21739.919923445781,0.14598583436295401,1123.3709803714708,25.073917137869799,208889.01834090837,0.40760645897833497,1166.7422237512471,2.4811198997189301,3086.939549541381,4.2748775330670702E-2,100.000325938061,12.430366753469199,21573.237797023288,49.603752087061402,3678.7196310552013,-6.7512622656319099,653.347310821007,126.26736918844,255238.4497154406,25.371474096690299,57244.027467310021,1.6344469648963799E-4,3.3333380520343074,7.56085770695892E-3,76.666977057849707,8.2501528838539603E-2,785.01855316233798,"<0.000",0,0,263403.18617915333,5.7132412580748705E-4,25.00003141661438,4.8613857469534301,131503.41548273317,3.4797219931425299E-3,29.576718718161224,1.2789665874485501,15314.574339271967,2.2945541139963201E-3,40.000096360840949,-8.6861684250294697E-4,5.0000062833228762,"<0.005",0,5.5412016199912397,8869.1654916169482,8.9737674675670204E-2,730.01626516357658,12.3769733082837,21310.212084796844,-8.2608097325087204E-5,23.333408882219459,"<0.000",0,1.92850580039275E-3,36.666774203378345,"<0.000",0,"<0.000",0,9.7875151819251096E-4,110.00048155585209,1.11031997909634E-4,43.333574285106245,6.2627869607840604E-4,20.000050961970526,0.21318762539885899,4250.6215091933655,8.5592187246599297E-3,266.66965087250395,1.18186362627246,68236.056925911849,12.1697687497121,474349.94084830181,1.36899841948917,65230.772139964916,1.26643465034848,76282.617678404582,0.25813996292958602,21257.921232711455]
[null,false,"085SMPL.d","2019-10-22T13:30:31","Sample","Reagent blank13",false,false,false,0.94572442767334097,200.00906437655684,-0.147789626637396,30.00006869435208,-270.946060821836,38951.45042869274,0.47498270286028799,33.333444893358475,-49.561347833383302,373.33771467202331,5.0855182651429098,19043.516966183117,-2.26852157376269,213.33545049028422,0,25331.159290505064,4.17290500288829E-2,18.00001335144043,1.5923479569075701E-2,496.67482271778346,3.8401416074890199E-2,2466.838443872753,4.8330357175563297E-2,440.0063131251984,0.14348748784103901,8902.4583899034205,1.51266882307546E-3,126.66729102531139,2.1351093961480699E-2,433.33909978460616,4.26751905571503E-2,1000.0343228387557,8.0968013333974909E-3,433.34601099320776,-5.0129508905742796E-3,5.0000058859585836,-3.5295148438006302E-3,25.000036979714476,1.9739932245686699E-3,10.000023692845945,2.3692789567180701E-2,8.3333534747359135,-9.0985053910223694,586.6780893006212,4.0733804509421601E-2,166.66905060415829,2.03510354972638E-2,73.333588143185935,"<0.000",0,-8.9864467846848602E-4,13.33335220813723,6.2928250150319205E-4,58.333516841130695,"<0.000",0,0,268010.48428758257,-1.3195296235691799E-4,21.666709656516066,6.8040398423379903E-2,2789.1356349874386,7.0218734915267203E-5,1.8037494977750148,2.1396539502241802E-3,12.666695485511983,1.6220178766382899E-4,31.666731759904845,-5.4135995607583399E-5,10.000012566645752,"<0.005",0,0.43736875747727399,1600.1016974471888,-6.4440163574471996E-4,3.3333384493986,-3.1088181232958299E-4,16.666694631178757,"<0.000",0,"<0.000",0,1.7230461019003101E-4,3.3333384493986,"<0.000",0,"<0.000",0,-2.7664650751074298E-4,56.6668875515428,-5.0772601740355997E-5,30.000097304581146,1.02585313200064E-4,3.3333384493986,4.1856149873341603E-2,2633.5852893746583,-1.1028595681384701E-2,146.66773577529588,6.1869374592051396E-4,46.666843195753643,4.8669022685369897E-3,416.67301828653086,2.1169228287494599E-2,1253.3877122394088,4.78501302276036E-2,4804.0918270137136,4.3759912154841299E-4,36.666778177021278]
[null,false,"086SMPL.d","2019-10-22T13:32:00","CalStd","Blank control5",true,false,false,0.87619866229130405,200.00634640469431,0.91212575218985803,46.666840811567887,1637.4178708966799,37354.813896494896,8.2297101347621098,46.666814585524577,1791.1895872206701,353.33733350032571,107.85756440185099,17441.881188196265,24.313789336799001,243.33586648102801,16.789927684734,27340.797932434114,9.4850906347703101E-2,10.666672070821127,6.4191173547548996E-2,396.6721157728814,0.27780916360322999,2763.5504755842117,0.139857891637583,366.67085736983728,1.5409540970059701,9636.0211300953979,6.2746054985686897E-3,116.6671195129486,6.9090364242712199E-2,410.00586852422549,9.5804681126961597E-2,841.69109954242879,0.27342486087787998,533.35400025096044,1.6661352361936999E-2,21.666708464423195,1.66515168885807E-2,28.333377415934539,3.6934386431786499E-3,6.6666856408116368,4.3458313119352603E-2,3.333335667848552,20.100763185252799,670.01410439589461,9.6713139170295101E-2,200.00571062180248,1.8852552854408601E-2,43.333480507133196,"<0.000",0,2.19446218935923E-3,16.666718870400604,7.1059286784700398E-3,63.333502461510363,"<0.000",0,3.7056140145440501,260976.59767364562,3.5116188924587399E-3,18.333360478281573,3.0893933932825999E-2,822.24930375874965,"<0.000",-2.3147969406624949,"<0.000",-93.338056445051379,3.8385747039204401E-2,158.34120713162611,1.9864251613338798E-3,11.666699474056074,"<0.005",0.66666690508524573,1.0355091333684501,1466.752466561328,2.5104610780678199E-3,20.000042617320378,8.8000029288708894E-3,15.00001924733292,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,5.4439859377283697E-4,23.333389014004833,3.53759266382376E-4,30.000066310166329,1.0535021038907101E-4,3.3333384493986,0.31904884281034901,3066.9906030171005,3.11958508161315E-2,193.33472917475208,5.82797215220839E-4,33.333427011965313,4.4040968648245997E-3,170.00142162042724,5.4557597140578798E-3,256.6691296795637,2.8360009986096901E-2,1650.0926979865726,"<0.000",0]
[null,false,"087SMPL.d","2019-10-22T13:33:29","Sample","19286424-001",false,false,false,-0.11499585145772,166.67074734975083,-0.232909265368933,33.333407541114973,112.964426194016,38303.439370640524,-3.3270832943264801,26.66676084200401,340.39964676246802,403.3381465573388,13.8350140730922,18876.636672739511,-0.35525727598773599,230.00193571048075,0,26225.8150388924,4.7145065258997303E-2,15.333346684773764,1.1568189757301599E-2,450.00685974945344,1.5022185815487001E-3,2673.5374052295474,5.2860995256321798E-2,486.67405992735331,-0.15297831166977099,8368.7295876904482,-4.23786899548978E-4,105.00036401052728,1.5209851519857299E-2,483.34068849671849,-9.4952334245034103E-3,733.35429877089643,52.689281508778798,100016.60990156105,-1.26937237736702E-2,5.0000058859585836,6.57298137187271E-3,38.333426142730922,7.7179344272655503E-3,20.000056922434911,6.8324193900649394E-2,8.3333443353571841,-0.46771231647584,636.67989417890851,-3.0569403158917E-2,133.33548208082553,6.4148297545301904E-3,56.666880398985541,"<0.000",0,-3.9857981997355702E-4,13.333361744880252,-4.6237945974462302E-3,21.666693364580073,"<0.000",0,0,256307.86076404862,2.0142846147415401E-3,28.333378210663128,-2.8380753999235299E-3,733.35494515016978,"<0.000",-3.3609195621671351,"<0.000",-70.668562948675515,-7.9413556540595693E-3,123.33649769420678,1.19195628735153E-3,18.333364849288788,"<0.005",0,-0.48436263555099301,766.70855010642845,-1.2323668903842799E-3,10.000014950831508,-4.8176489012563897E-3,6.666674117247152,7.5853391110017494E-5,3.3333384493986,4.2171980588944798E-4,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,1.6829441643680099E-4,30.000168035425229,2.4657997355874E-4,50.000170121587757,1.91899075670585E-6,3.3333384493986,-9.8363136411364002E-2,2083.4824074093049,-1.3123206798936599E-2,110.00061030188287,-4.0477332341054101E-4,10.000026474395993,2.560771708229E-4,176.66811997689624,-3.7242949835092401E-3,80.000447034829151,-1.19094769355762E-2,940.0312665100206,8.3196097545209995E-5,6.6666768987972]
[null,false,"088SMPL.d","2019-10-22T13:35:04","Sample","19286424-002",false,false,false,-0.101378451448013,166.6725593309925,-0.63558191860037405,13.333360950151667,143.63425941518099,38290.025471100824,-2.6156514883370399,30.000088562566713,252.99043218767699,380.00480085604187,15.354846649750399,18776.694653904069,4.3143575534070902,270.00306045011092,0,25765.178652527258,4.9618809968932601E-2,15.333344141642252,-3.69698319240292E-3,353.33749880387148,1.6736670336472501E-2,2773.5565280881042,7.6732308623999601E-3,366.67094002161019,7.65512189554753E-2,9602.7111928274262,-1.56065459277985E-3,83.333664288118527,1.70259932179154E-3,400.00547448785892,-3.0426793833779298E-4,800.02033114478729,57.256171544946298,107157.31847588187,-1.39808047438395E-2,3.333335667848552,8.9199798906808708E-3,41.666797970730094,"<0.001",0,-2.0831157470965998E-2,1.666667833924276,0.48949526818019601,660.01381208478688,4.5365160599544603E-2,283.34062335408134,6.8250897079884507E-2,193.33497077224197,"<0.000",0,1.43164357802159E-3,26.666733026503525,-4.4095302089607099E-3,23.3333747088903,"<0.000",0,0,254401.78657370131,-1.8741769510122301E-3,8.3333419511714286,-1.2478222716295499E-2,477.79036636181991,4.4239741256570401E-4,1.2279161087485888,"<0.000",-17.000035032629448,-3.1339529947328498E-2,28.333391721049079,5.1333269501246601E-5,11.666674837469932,"<0.005",0,-0.431948090248576,833.3792554845744,-1.4374024477927E-3,8.3333471169072322,-5.7908590152390797E-3,5.0000062833228762,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.2527421358099501E-4,13.3333537975944,-1.92469645453475E-4,13.3333537975944,"<0.000",0,-0.11449777587047499,1916.7947561819176,-1.57466155197401E-2,93.333729306850884,-4.47237026465035E-5,30.000100086131194,-5.0629326967945704E-4,146.66761656600809,-3.2025287954616198E-3,103.33387061952745,-1.9132887439822999E-2,523.34351077660642,"<0.000",0]
[null,false,"089SMPL.d","2019-10-22T13:36:33","Sample","19286424-003",false,false,false,"<0.005",0,0.12977723952023101,50.000161379573321,141.532248890402,38066.351971829652,-4.4694514026023899,20.000082353749637,244.561876790387,376.67184904209995,2.9528492404728399,16807.910531936071,-1.5880458954619801,213.33495140073273,0,25644.936141298171,6.0678459678491104E-3,10.666672388712565,2.4016353619010899E-2,513.34165443964321,-6.1426611467152404E-3,2550.1869838905282,3.4948446479090199E-2,433.33926508815193,6.3023229494429497E-2,9502.762053223596,6.1842335630566899E-4,121.66724977393078,5.1973285855846499E-2,683.34765369672095,-1.12033503441906E-2,708.35056298761344,59.723492183071798,111752.47238718532,-1.26425594596237E-2,5.0000058859585836,-8.4769452287546294E-3,13.333349823951474,2.07501641696592E-3,10.000042766331989,"<0.014",0,-5.5788707604231798,466.67342444250852,-1.34327965899304E-2,166.67000825213907,-1.28696102895618E-2,13.333380818366296,"<0.000",0,-1.7434703082958301E-3,3.3333380520343074,-4.6171279247009499E-3,21.666709259151776,"<0.000",0,0,256310.4388890102,-3.1865693017547E-3,1.666667833924276,-1.51660139503286E-2,411.11783840605085,4.4900512280168699E-4,1.2895624524917113,"<0.000",-41.667161559057803,-1.9872721292815299E-2,75.000833347427346,1.48086612462051E-3,20.000039835770334,"<0.005",0,-0.55625816857838295,666.700250904516,-1.2323775624149299E-3,10.000032037496089,-5.8132670354615404E-3,5.0000062833228762,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.27649184036425E-4,13.333414196966872,-1.13626057400623E-4,20.000062088170719,4.3099073291900101E-4,16.666702578464609,-0.12661445175602101,1816.7814660843044,-1.31233726550345E-2,110.00070725877026,-4.0477531215263101E-4,10.000015348195801,-1.67837241481194E-3,103.33379273612609,-4.3736083503913298E-3,50.000131179887092,-2.0776643615659798E-2,433.33995173364951,4.1597630354344902E-5,3.3333384493986]
[null,false,"090SMPL.d","2019-10-22T13:38:03","Sample","19286424-004",false,false,false,-0.41607272702869802,100.00407919296799,-0.843699009876686,3.3333380520343074,106.29691087435999,37875.843348366361,-3.2906232317461201,26.666750510532399,444.99434730808599,420.00581751258443,16.148937880267699,19093.563252295229,3.6708252872743099,266.66905323660791,0,26032.207457141965,5.4606417523409403E-2,16.000022252400715,1.06205697703569E-2,440.00617325296736,5.32857425041017E-3,2680.2081916698585,3.1198339391764002E-2,426.67395909616408,-0.262987133703456,7601.7564565335915,-2.0333412841291401E-3,75.000246837731495,5.7917081321837399E-2,716.68230896171463,-1.7981017833861801E-2,650.01404943596924,56.761338036477298,105750.63306032542,-8.5698198295035104E-3,10.000011771917167,2.9346658963756399E-3,31.666727388897627,4.0820861896633696E-3,13.333371281623274,2.2946406202412201E-3,3.3333380520343074,1.0650504340212701,670.01387551406208,-1.18192112259068E-2,166.67072152107085,-9.6855596247497795E-3,20.000047385691889,"<0.000",0,-3.4427413040941801E-4,13.333371281623274,-4.3459254074252E-3,23.333383053540441,"<0.000",0,0,247423.34917768813,-8.1781758106908996E-4,13.333353002865818,-1.6362469691052699E-2,366.67288434188225,7.3583517005669605E-4,3.4420013137360699,"<0.000",-69.335209846468615,-8.12668144680527E-3,118.33646280065483,2.5033624836059199E-3,25.000040953357409,"<0.005",0,-0.63834468199051697,533.35009813346187,-5.2447935034213801E-4,15.000018452604339,-1.5806331276084001E-3,11.666683182120076,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.93883919546846E-4,30.000087767838124,1.4375848834157501E-4,40.000193317728339,"<0.000",0,-0.114214714876695,1866.7953272689331,-1.6445509106595299E-2,86.667094528668699,-5.2132515089746098E-4,3.3333384493986,-2.0358924234250599E-3,86.667046844953589,-4.3347360591524404E-3,50.000263899560814,-2.2921438346455499E-2,300.00323166445042,4.3091757218143797E-5,3.3333384493986]
[null,false,"091SMPL.d","2019-10-22T13:39:32","Sample","19286424-005",false,false,false,-0.398188018513979,100.00136122110548,-0.55668355742216402,16.666707744200412,141.21508749780801,37187.804621639778,-2.45710615819773,30.000088562566706,71.023608233560495,336.670082161773,20.746761604381501,19060.111916912119,6.5764943349017697,283.33650166784969,3.5527136788005001E-15,25057.534620103223,6.0072000509229298E-2,16.000017166137695,1.20112208204914E-2,433.33940496038286,1.87402217306861E-2,2720.2111411063202,2.3891593622340002E-2,396.67184874407684,0.175307495437032,9936.3428429441792,-1.0569152283892199E-3,90.000349928930163,5.6602432452371898E-2,693.34831476189231,5.0212068560227703E-3,825.02484576211282,51.851422039889698,94887.211186649642,-5.6960836267530403E-3,13.333359360694496,-7.2421242179551001E-3,15.000024810433018,"<0.001",0,4.8954622396369399E-2,6.6666737198828594,-0.25573375557355099,623.34666276951634,-6.2633337853126803E-2,66.668575505458264,-9.6716220358399503E-3,20.000085532663977,"<0.000",0,1.12276152712819E-4,16.66669979691456,-3.7767722612142901E-3,28.333387747406146,"<0.000",0,0,250587.15772429373,-5.1935570527720804E-4,15.000016465782872,-1.82851144000909E-2,322.22703488347429,"<0.000",0.10707936723412725,"<0.000",-27.000097826121781,-2.7024319902994101E-2,45.000163043536304,1.5600541651643401E-3,20.00002235174146,"<0.005",0.66666690508524573,-0.76590685553673099,366.67702222886982,-1.8568252094687901E-3,5.0000062833228762,-4.7267347225851703E-3,6.6666745146114446,2.3275518885726001E-4,10.000015348195801,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.3942043373982299E-4,16.666702181100316,1.4667649176246E-5,30.000119556981531,4.3678602464260196E-6,3.3333384493986,-0.14208371237432299,1633.428832642385,-2.0552858069553401E-2,63.333581934368873,"<0.000",0,-7.1676406524622904E-4,136.66750704247494,-4.2013080638443299E-3,56.666915764407577,-2.44220825581518E-2,220.00214566783885,"<0.000",0]
[null,false,"092SMPL.d","2019-10-22T13:41:05","Sample","19286424-006",false,false,false,6.9500092405132802E-2,200.00362843283179,0.213004749564632,53.333544731136989,89.6049469134596,36503.036565437367,3.1907994904518602,60.000237524505884,470.40682058822898,413.33953241497971,12.899783748150201,18092.565305116495,0.13142791518696501,226.66853884853117,0,25331.158476702996,1.02403723754362E-3,10.000010172526041,1.42408669001236E-2,450.0067325928797,2.6855012797463499E-2,2816.8904202646449,3.6421586631579198E-2,430.00630055855271,0.41701996706279798,11403.860474215806,-6.9853064187544499E-4,96.666994243855356,0.32994505853800099,2210.1393815357578,0.169192505652733,2175.1457700538435,56.884917027458201,104275.24036784262,-1.1185369272718E-2,6.6666737198828594,2.1622382384070202E-3,30.000073860087884,"<0.001",0,2.8226688169287598E-3,3.3333380520343074,-4.46099007527389,490.00773653376496,-2.8253836671989E-2,133.33500524365664,-8.0842609813342806E-3,23.333414047955262,"<0.000",0,1.2886740427191801E-4,16.666728407143626,-5.8509167542042395E-4,55.000127926466938,"<0.000",0,0,247672.28797606201,-1.4932972679640899E-3,10.000021308660189,-1.9896335250851702E-2,277.78194836453645,"<0.000",-1.4584330003690305E-2,"<0.000",-13.333359360694498,-3.1999569465367098E-2,25.000045324364624,1.60179524190696E-3,20.000035464763112,"<0.005",0,-0.68834090895709599,466.67948812274977,-1.84913289832703E-3,5.0000035017728282,-5.5758822111163196E-4,13.333351016044352,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.9314184217407199E-4,30.000086973109543,6.0421543596519397E-5,33.333446085451349,"<0.000",0,-0.169234295893369,1366.7374916390515,-2.1561099322557401E-2,56.666893909371481,-2.7574541570272201E-4,16.666714499393386,-1.21932784986812E-3,116.66738972066757,-4.6345028818968103E-3,36.666820297636285,-2.3953160921326801E-2,243.33579018708383,"<0.000",0]
[null,false,"093SMPL.d","2019-10-22T13:42:34","Sample","19286424-007",false,false,false,-0.55503872355013495,66.668480138024492,9.0863390604191102E-2,46.666788359481266,109.183683192537,36235.886504314047,-4.9977369538224004,16.666690260171539,-25.958976005226798,316.6703622042582,18.912750230140201,18643.106495807468,4.6195573221071697,263.33594565586333,3.5527136788005001E-15,24863.896113874074,2.2278548292052398E-2,12.000008583068848,7.3275567674456303E-3,403.33892857026666,3.00951137369176E-2,2800.2229909068897,3.78437835909212E-2,426.67300542186194,2.24643394801218E-2,8969.1604075889863,-1.2026799659497699E-3,86.667024592553219,1.32819925733242E-2,450.00743195403464,-5.28608097294585E-3,733.35167616646743,50.211688930238203,90967.958576261066,-4.1953841227122396E-3,15.000012889504241,-6.0835738251059698E-3,16.66669979691456,"<0.001",0,3.2727730714858902E-3,3.333335667848552,0.72639477148988596,646.67960156977767,-3.6348773869319798E-2,116.66917266950645,-1.42050575637515E-2,10.000023692845945,"<0.000",0,-1.2598123122222801E-3,6.6666761040686149,-2.3425638245980701E-3,40.000078876812076,"<0.000",0,0,247061.12913100226,-2.1627440077983001E-3,6.666674117247152,-2.0310152656825099E-2,266.6699098216888,5.6196975854247097E-4,2.107082494093766,"<0.000",-18.66670485337518,-2.6862187583302201E-2,45.000097875792314,2.5099471730034599E-3,25.000052874286183,"<0.005",0,-0.76205838729831299,366.67800769235208,-1.8474965920112899E-3,5.0000062833228762,-6.7343014745678701E-3,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.1579191023369699E-4,13.333374460537614,1.99257718484664E-5,30.000117570160068,"<0.000",0,-0.16153652254162701,1433.4173522908393,-2.4378031791754198E-2,40.000126560527185,"<0.000",0,-1.21144711590118E-3,116.66748190918342,-4.7821614041324897E-3,30.00006869435208,-2.49710569139064E-2,186.6682151953149,"<0.000",0]
[null,false,"094SMPL.d","2019-10-22T13:44:03","Sample","19286424-008",false,false,false,-0.38962642121885499,100.00226721172633,0.24565103128416199,53.333470821378569,116.142624602943,36018.790967184439,2.2164200210383602,53.333535989122545,592.33189736793304,423.33908721796041,14.5706016190589,17825.650422848397,1.21071396462454,230.0022599597435,0,24616.851846721987,7.5716469920202797E-2,17.333353042602539,2.5620870397969099E-2,503.34146385379444,-3.88972175921326E-3,2480.1782874239739,1.8813243297412598E-2,380.00459740552407,0.110842927754086,9469.4208354373459,-2.6698266983526198E-3,61.66681038836424,7.7523662083417003E-2,803.35224226088974,-1.6981772893254601E-2,641.67990529318433,47.549788172596998,86746.78742071263,-1.8054094716993599E-4,20.000025530655797,-6.1920747202801303E-3,16.66669979691456,-1.7283537394026901E-3,3.3333380520343074,-2.04170466968248E-2,1.666667833924276,-2.03057816837743,570.01106684390106,-4.5968908790201997E-2,100.00094994904732,-6.7198566672288899E-3,26.666733026503532,"<0.000",0,-8.2693222800870204E-4,10.000023692845945,-4.0195436958682098E-3,26.66671673456753,"<0.000",0,0,255576.64983744445,-2.5336698154769998E-3,5.0000062833228762,-1.9810157703569999E-2,288.89342529916308,"<0.000",6.166404393816638E-2,"<0.000",-26.333400413392972,-2.6833558268181099E-2,46.666780412195415,6.2150810442962803E-4,15.000021631518678,"<0.005",0,-0.55488506170901997,666.69626136686986,-2.0832111356764399E-3,3.333335667848552,-8.1250122558604299E-4,13.333358963330205,2.2821144608408899E-4,10.000024884938822,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-3.8556989695787602E-4,6.6666768987972,-1.12936742160919E-4,20.000050961970526,2.1737841292329501E-4,10.000025282303115,-0.17740812935547601,1333.4077815231433,-2.4605195728048999E-2,40.000114242234119,-5.23286207288658E-4,3.3333384493986,-1.7587234810265901E-3,100.00052462020729,-5.0216566002337397E-3,20.000042219956086,-2.4615965736673798E-2,213.33521207170864,"<0.000",0]
[null,false,"095SMPL.d","2019-10-22T13:45:33","Sample","19286424-009",false,false,false,"<0.005",0,-0.34896931149115601,26.666740973789377,66.747631660152805,35982.060010700523,-4.4188765958207297,20.000037054220282,271.94564953147602,376.6715947289527,15.0327143877544,18392.783003884175,4.8354488772082096,270.00256453947378,0,25304.536569547607,-1.17121756409924E-2,8.6666712760925293,8.8376274236588704E-4,373.33803256345738,2.3491264962729699E-2,2786.8896069096086,3.3536996617385503E-2,423.33909993361777,-0.15748513197682401,8068.6354851169372,4.4289701223735201E-4,116.6672530273509,1.077018044692E-2,443.3400024472673,-6.7075486019014304E-3,733.35003902552114,58.330366461816702,107258.31335118617,-7.1137490365825303E-3,11.666684771577247,-1.03851437682934E-3,25.000041350721698,"<0.001",0,-2.0435180693809899E-2,1.666667833924276,-2.3254886262471799,560.00974456454696,-6.2703450030859295E-2,66.667860249704972,-8.1583676521545904E-3,23.333385437726196,"<0.000",0,4.7213489916372196E-3,50.000193168716727,-5.7346968460542797E-3,11.666684771577247,"<0.000",0,0,250166.10304306788,-1.84645119871276E-3,8.3333491037286951,-2.17480532325003E-2,233.3365280595805,7.2085949551079105E-4,3.3641697293892641,"<0.000",-12.333352223038389,-3.2484734268198499E-2,23.333366761604442,-2.1020078115338599E-4,10.00001216928146,"<0.005",0,-0.44628789341641001,800.03352363979263,-1.4192346485162701E-3,8.3333419511714286,-6.7599402363653603E-3,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-3.82134778906494E-4,6.6666768987972,5.6292471270417997E-5,33.333425422508135,"<0.000",0,-0.15987452376101,1466.7468716718797,-1.9412723085362502E-2,70.000400294854231,-4.6120080999644903E-4,6.6666880249973923,-3.5021246582514498E-4,150.00092918672769,-5.01226790616987E-3,20.000041822591793,-2.5252185938097398E-2,173.33460867403039,4.2619310782601998E-5,3.3333384493986]
[null,false,"096SMPL.d","2019-10-22T13:47:06","Sample","19286424-010",false,false,false,"<0.005",0,-0.182379813364907,33.333426614601017,120.023029508836,35795.000280415006,-1.64552927024242,33.333426614601017,536.83169475176896,410.00533446661638,17.805310490846299,18142.799274338267,3.2923611313080201,246.66862756010948,0,24409.887331576832,-1.5363517245603199E-2,8.0000044504801426,7.9627781403368292E-3,400.00574151666359,2.55280407484226E-2,2713.5433555611467,3.3631905040949699E-2,410.00527088832951,0.164475764021427,9636.0531735531458,-1.7219768163612499E-3,76.6669452687063,7.7551960513912804E-2,790.01954490910657,-1.1369383247547601E-2,675.01503545824517,60.6177812647473,108334.93487590157,-8.2498648970415998E-3,10.000009387731412,-1.6862526895532299E-3,23.333394974469218,"<0.001",0,"<0.014",0,-1.89160247599678,560.01063466056235,-2.71492189469556E-2,133.33523571495493,-7.9245087568573895E-3,23.333394974469218,"<0.000",0,1.0980688329104099E-3,23.333375900983174,-5.3539923364308798E-3,14.634693697928977,3.9465630328992601E-4,3.3333384493986,0,246214.68383111534,-8.0459325412828801E-4,13.33335538705157,-1.89462308035716E-2,300.00404392688569,6.6229106769906897E-4,2.8645825940681431,"<0.000",-25.000063523649221,-2.76790990986228E-2,41.666772539415369,1.02145983873297E-3,16.666692246993001,"<0.005",0,-0.78606326453042796,333.34319541873725,-2.28871331871296E-3,1.666667833924276,-6.7272016648107398E-3,3.333335667848552,1.5792598488680699E-4,6.666686435540222,4.3900753443078998E-4,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,-2.1466247038513101E-4,13.333364129066007,-6.2116733271350201E-5,23.333379079897515,"<0.000",0,-0.17937442343872601,1266.7313858881159,-2.3784484158616202E-2,43.333495606976307,-5.2102338516694599E-4,3.3333384493986,-2.5734591803161502E-3,66.666932702060549,-4.8549457812457897E-3,26.666792631147416,-2.5870285224132698E-2,136.6676373779629,"<0.000",0]
[null,false,"097SMPL.d","2019-10-22T13:48:36","Sample","19286424-011",false,false,false,-0.54443931149138702,66.668480138024492,-2.40475254697445E-2,40.000101923941052,101.30666937318099,34919.983203610034,1.1184995165073299,46.666824916996184,339.415271408867,370.00455940756365,6.4562709685143496,16274.127052394409,-0.10054277274019099,213.33528518673847,0,24069.517056700828,1.9256056715541099E-2,11.333342870076498,1.19084796945957E-2,416.67216633748762,3.5857898431718903E-2,2773.5568841265108,3.6942268081687699E-2,413.33876947553796,0.15179187201196001,9469.2494283697142,-4.8133388313523998E-4,96.667005370055563,7.65513356581418E-3,410.00523274135747,-9.6761120671011395E-3,683.35083263654872,51.319098943439997,91170.945234444007,-5.3787900891108902E-3,13.333349823951474,-1.1279047323278999E-2,8.333341553807136,-1.6738978893274701E-3,3.3333380520343074,3.9203780544130899E-3,3.333335667848552,-6.9493934185892998,403.33838815482869,-1.8360693714192699E-2,150.00202991072902,-1.5729027064768902E-2,6.6666761040686149,"<0.000",0,2.0337589421650501E-3,30.000118762252942,-4.7183334707881202E-3,20.000029901663016,"<0.000",0,0,247483.28407262836,-3.1749755366610398E-3,1.666667833924276,-2.1648895588655199E-2,233.33684595102636,6.8407494037985201E-4,3.0462495722435832,"<0.000",-21.00006191432384,-2.9438260270990401E-2,35.000103190539733,1.60453729145942E-3,20.000045001506134,"<0.005",0,-0.78734265864626096,333.34220995525493,-2.0692389190488502E-3,3.333335667848552,-5.70673709253807E-3,5.0000062833228762,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.34341298617075E-4,16.666702181100316,-2.21619299663652E-5,26.666727860767725,1.16838388589894E-4,6.6666872302688072,-0.207516667245323,1016.7162168782103,-2.3255286645130699E-2,46.666873395439886,-5.2134003805127797E-4,3.3333384493986,-3.4934710250776602E-3,33.333427011965313,-4.6338752396586202E-3,36.666840165850921,-2.5762222081979699E-2,143.33424354591602,"<0.000",0]
[null,false,"098SMPL.d","2019-10-22T13:50:05","Sample","19286424-012",false,false,false,"<0.005",0,0.23569023592000299,53.333492279050368,70.425805054992196,35384.188112057425,-3.0515919578806399,26.666722695031922,292.72563461401398,373.33790540688375,18.174079045692999,18509.572252100086,-0.109139369938706,220.00185956554819,0,24830.487911246895,1.5823129150755198E-2,11.33333937327067,-1.5639639334433699E-3,353.33755602432956,2.2746160769816299E-2,2736.8843179412038,4.02328932043352E-2,433.33925237249451,0.23799537597873899,10236.462376833364,2.3506349199791299E-4,111.66717362899819,4.3863566464275701E-2,620.01198162617311,4.4376861718892898E-3,816.69578788246235,68.195409545047696,124172.56543823353,-7.0363225267994003E-3,11.666684374212954,-7.2178270557066403E-3,15.000015273689996,"<0.001",0,2.5949238272744202E-2,5.0000058859585836,-3.8505330830877802,510.00888064490437,-7.1162224719527004E-2,50.00059517722665,-1.57950354888515E-2,6.6666856408116368,"<0.000",0,2.8698954468399499E-3,36.666756719349472,-3.5922484634989399E-3,30.000060747066225,"<0.000",0,0,251824.76187251866,-2.8499391421690501E-3,3.3333384493986,-2.5702151149098001E-2,133.33485159612448,"<0.000",0.47041603023187972,"<0.000",-15.66668884456125,-3.0429979418488601E-2,31.666707520683001,9.54450569815945E-4,16.666689465442953,"<0.005",0,-0.74283970483998896,400.01577635641388,-1.4264217147102901E-3,8.3333447327214767,2.34644167040297E-3,18.333366836110251,1.54407535323705E-4,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.22008244450563E-4,13.333363731701716,9.4326943964631103E-5,36.666785329578538,1.1300784046316101E-4,6.6666872302688072,-0.200455369764658,1100.0512686879968,-2.3392194528117499E-2,46.666801869867221,-3.4120620873807997E-4,13.333376049994785,-1.2718390365180901E-3,116.66754548747025,-4.7214746044793596E-3,33.333428998786779,-2.6163240925194699E-2,123.33411827682279,"<0.000",0]
[null,false,"099SMPL.d","2019-10-22T13:51:47","Sample","19286424-013",false,false,false,-0.37212943367554702,100.00226721172633,6.23954844299629E-2,43.333469380933003,103.39914517572301,34515.855605465338,-6.8769921385627004,6.6666761040686149,464.16743505983601,386.67112410061873,20.872363626989902,18092.58870192691,7.4936719118016297,276.66928832725745,0,23762.36455476734,8.1687245021134003E-2,17.333347956339519,1.5763151828133801E-2,433.33920150986506,3.3704874277666698E-2,2730.2138875392989,3.10961894394613E-2,396.6715181369853,5.4356954680691003E-2,8869.0040027624273,6.5105728042999403E-4,115.00045605003156,7.1797467257777595E-2,750.01931662332038,-1.0538490871706101E-2,675.0173560658003,54.988360971931002,97569.883365752685,-2.5646820793255101E-3,16.666689862807246,-6.9984574790696004E-3,15.000027591983065,"<0.001",0,-1.9874796774363102E-2,1.666667833924276,-5.5128168554714101,450.0070250529991,-4.4919961410820899E-2,100.00214601561258,-1.26667032581077E-2,13.33335220813723,"<0.000",0,-8.0261283076962699E-4,10.000042766331989,-2.5988765823299498E-3,38.333412234980678,"<0.000",0,0,252301.903792492,-2.8511910320128901E-3,3.333335667848552,-2.0961808096018698E-2,255.55917463376932,6.02762406262417E-4,2.4704166374045333,"<0.000",-15.666696473955668,-3.0445021816893798E-2,31.666720236340364,9.4889031865033602E-4,16.666697412728805,"<0.005",0,-0.62167976998439101,566.68688133404146,-7.7927840985483399E-4,13.333350221315769,-6.7772115816638898E-3,3.333335667848552,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-3.8350838398908197E-4,6.6666768987972,1.21635347189292E-5,30.000108430781339,"<0.000",0,-0.18991735718693301,1200.069454313877,-2.5075994453163501E-2,36.666818708179115,-5.22513785697986E-4,3.3333384493986,1.0446575492652E-3,203.33488161362882,-4.5762839621025899E-3,40.000114639598408,-2.6937773555841302E-2,80.000504255287282,"<0.000",0]
[null,false,"100SMPL.d","2019-10-22T13:53:21","Sample","19286424-014",false,false,false,-0.71431680789830898,33.333787073701828,-0.18988686713866701,33.33341787258658,47.578054212232701,34676.176241097208,-4.3198457299978603,20.000028312205846,475.42125750757799,403.33901122203952,11.7073112093908,17441.813032270267,-0.31457743384321502,216.66848813491333,7.1054273576010003E-15,24663.600706702462,3.6185461267816903E-2,13.333340962727865,6.5392186482708097E-3,396.671823312762,3.9901453496161102E-2,2876.9057798846693,2.0289660723890601E-2,383.33800698313115,0.14185459962062399,9636.2708656153045,7.5233250274761305E-4,120.00060538449975,1.50592893234277E-2,460.00683416912699,-1.3479179752680199E-3,766.68610697034728,49.851859640658397,90599.794505531972,-5.6283354445825296E-3,13.333347439765719,-8.8739771701807496E-4,25.000060821572035,2.5789461132830801E-4,6.6666761040686149,2.9003839222464002E-3,3.333335667848552,-3.1878032200884898,530.00866418070598,-2.00008112973563E-2,150.00322001682963,-1.120572359478E-2,16.66669979691456,"<0.000",0,1.06735615225877E-4,16.66669979691456,-3.40103182769854E-3,31.666721825797527,"<0.000",0,0,252373.26446851194,-1.8610131831064401E-3,8.3333538721002061,-2.83036844981089E-2,66.667423148966847,"<0.000",-0.90833509561913561,"<0.000",-20.00003880262317,-3.0029450315591798E-2,33.333398004371951,3.5890986638624201E-3,31.666732951997727,"<0.005",0,-0.67046852863061601,500.01824225029389,-1.64511426513162E-3,6.6666768987972,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-3.0313127902258802E-4,10.000025282303115,1.3399987218573199E-4,40.000146826106111,"<0.000",0,-0.22222871448558801,900.04185463084957,-2.7302502489528899E-2,23.33340292175507,-4.62264447474721E-4,6.6666768987972,-1.9037423816392399E-3,93.333823084823948,-5.1626841687148596E-3,13.333364923794592,-2.7056662672150001E-2,73.333751857274493,"<0.000",0]
[null,false,"101SMPL.d","2019-10-22T13:54:51","Sample","19286424-015",false,false,false,8.0347275415264499E-2,200.00544041407349,-0.27198765499488498,30.00006154179481,-14.458752971989,33914.812320150748,-2.4540135163921399,30.000105251866998,293.39663620749599,376.67137856277753,20.478960252670898,19010.247719020223,6.9565591521732397,286.66916737950095,0,25044.131299990448,7.2930195109554605E-2,17.33335018157959,9.5814069993207501E-3,420.00599553178751,1.8149310286089401E-2,2720.2127687104626,1.5194970641849799E-2,376.67132770014814,-7.7783337895785995E-2,8502.2434863452891,-9.8421088215416711E-4,91.666951403017563,2.2872941596224401E-2,510.00815585243464,-1.78683177838318E-2,641.68001655519049,48.563367548096998,89527.126024524725,5.0978960059561698E-3,26.666702429453,-7.3233403798387102E-3,15.000027194618774,2.1526839438452201E-3,10.000023692845945,2.2737273872795402E-3,3.3333380520343074,-2.7845785702508898,550.00938867515242,-7.1478579038141296E-2,50.000352784999144,-9.7956951954687108E-3,20.000066459177933,"<0.000",0,9.8614568815502896E-4,23.333461731670372,-6.1431398446613701E-3,8.3333447327214767,"<0.000",0,0,255468.28395364259,-2.85937610531572E-3,3.3333384493986,-2.2790837054536699E-2,211.11386559081939,8.9006542782140504E-4,4.7737532573876713,"<0.000",-27.000063017009751,-2.72414123325866E-2,45.000105028349587,-2.4706491026977402E-4,10.000017732381556,"<0.005",0,-0.79509800934793096,333.34418088221952,-1.6555980136658201E-3,6.6666768987972,-7.8011429138877498E-3,1.666667833924276,2.2830825001867499E-4,10.000024884938822,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.70634694973802E-4,30.000086973109536,-7.2680516548311505E-5,23.333389411369122,"<0.000",0,-0.19151927186604001,1200.0562300297272,-2.6800213884530799E-2,26.666761636732598,-5.2326096359550504E-4,3.3333384493986,-2.1104650175964601E-3,86.667212148499303,-5.3833785467669298E-3,3.3333384493986,-2.6662776282784E-2,96.667177031429944,"<0.000",0]
[null,false,"102SMPL.d","2019-10-22T13:56:21","Sample","19286424-016",false,false,false,"<0.005",0,0.82313952752206698,80.000621080389308,11.559543971016399,33898.115543055792,-4.9679471186996196,16.666707744200412,552.890632814112,416.67192473999768,20.4249737873347,18693.217377461384,3.4077600187094701,250.00274285670011,0,24636.838825590825,-3.1112240702259502E-3,9.333338101704916,1.24235885108651E-2,430.00599538277601,2.28202212937124E-2,2726.8814697829662,5.59802943086204E-2,470.00669414788445,0.193879550381796,9969.683297790958,-1.31953844374611E-3,85.000273709491779,6.6155624777741798E-2,743.35117503974072,-3.9870743050956199E-3,750.01926509085649,53.996588955522803,98808.988618058211,-1.2557845312457501E-2,5.0000058859585836,-2.0723911416143698E-3,23.333373516797423,2.1825913746506301E-4,6.6666856408116368,2.3926261630628398E-3,3.3333380520343074,-3.17835615993378,536.67714908703431,-8.8306545104179499E-2,16.666784261666383,-1.4331577721114701E-2,10.000023692845945,"<0.000",0,-8.3647436335355203E-4,10.000023692845945,-5.2126358271890197E-3,16.484020286659096,1.8875024397430501E-4,1.666667833924276,0,257795.38678354942,-9.2619930259847801E-4,13.333351016044352,-2.6667654082294899E-2,111.11229774027403,7.5352815344251797E-4,3.7275025185818524,"<0.000",-6.0000075399874513,-3.5931586239873602E-2,10.000012566645752,2.4508841312516499E-5,11.66667761901998,"<0.005",0,-0.82109044910278095,300.0123249990516,-8.1616894402550299E-4,13.333350618680063,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-3.8693686904423501E-4,6.6666768987972,2.8290763007107002E-4,53.333518505093672,"<0.000",0,-0.23655285296093101,783.36053056082312,-2.0850423149454601E-2,63.333729753885713,-4.6479957200834299E-4,6.6666768987972,-9.07282150233233E-4,133.3343009153858,-5.1688485659711798E-3,13.3333537975944,-2.65041033053534E-2,106.66739145913634,4.1358020505684701E-5,3.3333384493986]
[null,false,"103SMPL.d","2019-10-22T13:57:50","Sample","19286424-017",false,false,false,-0.26081497225153299,133.33696027604898,0.117441295178854,50.000194758173905,-110.883822174061,33056.504227471778,1.05964358407422,50.000286946689783,273.86190108768898,386.67156279079774,8.5964629156523493,17875.599052708327,-1.85711555427518,213.33508491513501,0,25952.222920573407,-7.6696782285873601E-3,9.3333384195963536,1.4343152420381E-2,463.34081595131534,-1.66892042140238E-3,2630.1970026364374,4.1539259577270601E-2,456.67334829757584,-0.11800280861099401,8568.9366030703713,-7.0566637499629896E-4,100.00034501154704,4.5669644246482096E-3,423.33945597202393,0.15324825596189701,2125.1334347481552,52.512200361949397,100287.14213417044,-1.1412603516288301E-2,6.6666737198828594,-2.6508641739782001E-3,23.333369543154493,-1.8132281650537899E-3,3.3333380520343074,4.4789662364234303E-2,6.6666737198828594,-4.8112674317786297,503.34134941287812,-4.80201278539448E-2,100.00167116526525,-1.15709613520593E-2,16.666690260171539,"<0.000",0,-4.4102684387111103E-4,13.333361744880252,-4.8767224696526596E-3,20.000050167241941,"<0.000",0,0,264817.69136640889,-2.5677953661096199E-3,5.0000106543300946,-2.4722623532134499E-2,166.66834462023883,7.5044371009892698E-4,3.8037508449393229,"<0.000",-19.000032141804219,-3.08203155880632E-2,31.666720236340364,1.0898195638691301E-3,18.333362067738744,"<0.005",0,-0.82677904027365301,300.00838314512231,-1.6857790149047501E-3,6.6666768987972,-3.9820237566726901E-3,8.3333447327214767,7.3415866994350703E-5,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.6118254635912599E-4,16.666712512571923,-1.6007700620528599E-4,16.666701783736023,1.02294391733556E-4,6.6666872302688072,-0.20798298688489,1083.3811366320406,-2.5895292205992598E-2,33.333438138165498,-2.95625423359402E-4,16.666703373193194,-2.53185458030118E-3,73.333627879615207,-4.8273264521709498E-3,30.00012035171012,-2.70049915842831E-2,80.000478823972557,"<0.000",0]
[null,false,"104SMPL.d","2019-10-22T13:59:23","Sample","19286424-018",false,false,false,-0.57141011334633196,66.667574147403656,-0.43618879160836699,23.333394974469218,-91.200517992871298,33801.222640853783,0.35864636306768799,46.666851143039501,148.596131781704,366.6709845264109,-2.1896108456938901,16374.227762639646,0.36434898771517199,236.66862453020676,0,26199.077571842907,5.3225420254285602E-2,16.000013033548992,3.55715497448549E-3,403.33847716443023,2.6778039640963799E-2,2926.9124930066987,4.9169158983041597E-2,480.00777482974871,4.0289267729745303E-2,9602.7239084852608,-1.6718014997922899E-3,83.333607067660381,1.6061867531475201E-2,493.3412732679455,2.2420990202892E-2,1016.7051860450374,63.622516002870697,122321.21041638376,-6.2411348389531203E-3,13.333354592322985,-5.7297430642964499E-3,18.333358491460107,"<0.001",0,"<0.013",0,0.21900016453075499,673.34835320688751,-3.21879752735321E-2,133.33476483825069,-1.4509523628892801E-2,10.000014156102923,"<0.000",0,1.29302356777118E-3,26.666733026503532,-3.7792509787917001E-3,30.000062336523403,"<0.000",0,0,266019.2994214202,-3.19843258065637E-3,1.666667833924276,-2.5979195769019001E-2,133.33441714448173,"<0.000",-1.4584145924681713E-2,"<0.000",-15.000023141502989,-3.24400175857677E-2,25.00003856917165,7.9752856153037302E-4,16.666691452264416,"<0.005",0,-0.64302178308112101,566.68293948011217,-2.0999829317185702E-3,3.333335667848552,-6.8815160564610501E-3,3.3333384493986,1.4616870476011701E-4,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-3.9180475446004098E-4,6.6666768987972,2.2466499688318501E-4,50.000284165139732,"<0.000",0,-0.22379417228617199,933.36848119973217,-2.7502222031401201E-2,23.333391398190585,"<0.000",0,-4.2235778554032002E-4,156.66791682440166,-4.5521227847469302E-3,43.333518654105283,-2.7292132451121699E-2,63.333719422414113,"<0.000",0]
[null,false,"105SMPL.d","2019-10-22T14:00:52","Sample","19286424-019",false,false,false,-0.259542448103806,133.33696027604898,-0.29310816927717098,30.00011478861002,-96.598079695381998,33297.003084565156,1.07883294122578,50.000181247787957,171.09970496628301,366.67158216230695,10.471402438115501,18125.918984247426,3.46353469409026,263.33593611912028,0,25898.668233871442,2.37540271685922E-2,12.666675249735514,9.4631690437900796E-3,433.33968470484496,1.74600623921944E-2,2803.5562733279226,2.88893585681438E-2,423.33883290481322,0.22665226871357499,10603.337079857871,-8.6849446685382603E-4,96.667002191141208,5.5013953104655199E-2,710.01629089291464,0.34099004733845101,3708.7296361192239,55.736548584640602,105850.53133729791,-2.1454080684648498E-3,18.33335133890284,-2.5664202547361001E-3,23.333369940518789,-1.8012149300871999E-3,3.3333380520343074,"<0.013",0,-4.1927704847854299,520.00863889840286,-3.1281896976926203E-2,133.33500127001358,-1.0043213485379499E-2,20.000066459177933,"<0.000",0,9.0130860023281004E-4,23.333385437726196,-4.5007984423171797E-3,23.150705927470732,1.85443405319443E-4,1.666667833924276,0,262026.36328300289,-2.55774178759405E-3,5.0000062833228762,-2.79833162863647E-2,77.778378211803087,"<0.000",0.19791495147474869,"<0.000",-20.000041981537514,-2.8325185591920401E-2,41.666747108100651,2.2531432192711199E-3,25.000049695371843,"<0.005",0,-0.75423087193166405,400.02071956839774,-1.88536114030873E-3,5.0000035017728282,-6.8522824724979699E-3,3.333335667848552,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-7.9640858511798696E-5,20.000040630498916,1.16030335574762E-4,40.000112652776949,"<0.000",0,-0.22579674332342201,900.03068074652583,-2.7445934276740299E-2,23.333402127026488,"<0.000",0,1.52922867019341E-2,763.35333640346926,-3.0563862560460901E-3,113.33385934431563,-2.7447044288670799E-2,53.333470821378569,1.22070760883489E-4,10.000026871760285]
[null,false,"106SMPL.d","2019-10-22T14:02:22","Sample","19286424-020",false,false,false,-9.9373967863818996E-2,166.67074734975083,0.196926736469797,53.333606719966632,-95.651381558545594,33059.84569877616,6.1546208845900701,76.666997720792921,132.453280445258,356.67082572970548,12.710624130235299,18326.123965122093,-1.63532184488276,213.33511670427842,0,25698.415958209593,-3.1994126943177E-2,6.6666704813639326,-2.0591685997233601E-3,363.337435076573,2.2888968381674198E-2,2840.2319929962555,4.8313883882211403E-2,470.00754609692768,0.13807078350568699,10036.439484178909,1.01527136627371E-3,130.0005607306874,5.6410906442159299E-2,716.68305918549902,-2.0154735320495099E-2,641.68144706669693,48.596385576886199,92341.414417507083,-1.13863860899058E-2,6.6666737198828594,-1.58316460801418E-3,25.000041748085987,8.2124514457458705E-5,6.6666856408116368,"<0.013",0,-4.1551177621226199,523.34201032903775,-7.2315807314655298E-2,50.00059517722665,-1.0100557469619201E-2,20.000056922434911,1.6143962850930101E-4,3.3333380520343074,3.0692612879042501E-3,40.000132918355867,-3.3917442650977501E-3,33.333398799100536,"<0.000",0,0,265281.23953658482,"<0.000",0,-2.8019028296204101E-2,77.778375562707708,"<0.000",-1.8297730616265402,"<0.000",-97.337535917696314,9.6685531943848497E-4,165.00700625270099,8.0527461805448295E-4,16.66669502854305,"<0.005",0,-0.66507916236173903,533.35206906042652,-1.27559958664691E-3,10.000014950831508,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-1.6185239993008E-4,16.666702578464609,-5.7396300782179502E-6,30.000128298995964,"<0.000",0,-0.220116514417839,966.71503956227343,-2.69628151305405E-2,26.666740973789377,-4.6812910340045401E-4,6.6666880249973923,4.8558947121834202E-3,363.33836331956041,-5.0375370771726397E-3,20.000042219956086,-2.6612837342144799E-2,103.33388492464196,"<0.000",0]
[null,false,"107SMPL.d","2019-10-22T14:03:53","Sample","Acid rinses ",false,false,false,"<0.125",0,78.873227671876407,143.33409095802767,32239.135066536201,27136.816815682003,309.85112203255801,63.333539019025274,49701.733010531098,356.67128985119922,3554.4514209782401,20795.546424289339,497.38282189667501,183.3344718317021,7.1054273576010003E-15,960.02783298450959,5.3427659246559402,14.666680892308554,2.5674224440792202,296.66996831690318,19.3440628610795,2876.9014056985366,16.398148652150901,536.67543247329047,171.23243661880599,11537.224073795021,0.60418423270311095,106.66713873544626,11.436861817713901,573.34480702859946,6.2745803399059703,425.00630375011946,1111.4960967223401,15039.915293850594,0.38479647449217103,3.333335667848552,12.892123161872901,130.00057026743039,0.33096702769039998,3.3333380520343074,29.372961663957899,11.666680003205736,13666.8808789872,2220.1475702700877,18.395133900409402,166.66905259097982,27.626640535814101,263.33587889866209,0.20375663789780399,16.666728407143626,0.114020189650839,3.3333380520343074,0.80064668664092598,25.000050092736132,"<0.035",0,0,810.02193515466752,1.02502522309151,16.666705757378949,7.50144952145852,622.23671078714335,5.0998637181338097E-2,1.2717552106817038,"<0.027",-379.6742852378743,49.819860596779797,638.34603947878816,1.5522872784691799,28.333388542134731,"<1.520",0,423.59109970234698,1866.8323298332368,0.53671076552107999,13.333360552787374,2.5114022531043001,13.333360155423085,2.7842285075750701,386.67326033105564,1.33441108564335,33.333422640958091,0.114020225916958,6.6666768987972,2.90872682562777E-2,6.6666872302688072,"<0.013",0,7.4625614343464305E-2,10.000015348195801,"<0.004",0,"<0.010",0,238.25809978112699,7118.3498702077777,19.8972867321481,383.33854104074027,7.4524548694425499E-2,13.333376049994785,0.88590358794816304,106.66723251341928,0.24565183528776199,36.666838576393744,0.13776172704537401,30.000164856510889,2.6325027074182598E-2,6.6666880249973923]
[null,false,"108SMPL.d","2019-10-22T14:05:29","Sample","Base rinse",false,false,false,13.6125149780603,200.00544836135964,14.170733345065701,46.666815380253176,20740.398861424899,30872.768094647403,351.42729403698098,123.33403721450709,33695.902434784803,423.33988830437437,1747.31160203488,18142.496133055451,350.19930777489202,226.6685642798459,0,1653.4154518433097,1.82719578037422,13.333349545796713,0.20149450795315499,103.33382452526951,3.9219276305573501,2683.5403042504145,2.1868524290340599,400.00508030248074,20.356482660945201,9169.1581232409244,9.4633654163487693E-2,128.33401488760092,0.98376941535905205,436.67314514508126,1.5640191646711299,1041.6993691903708,44.6558469987311,6401.3060411109391,0.10342775804103101,11.666680003205736,7.3046874120186303E-2,11.666691526770222,"<0.007",0,0.77017631181827495,5.0000058859585836,211.40484347491301,633.3459251123777,-3.2648565645444498E-3,16.666784261666383,3.03248149710445E-2,10.000042766331989,1.77343205819328E-3,3.3333380520343074,0.633346105741735,450.00665629893564,6.5064077751040603E-2,63.333483388024312,"<0.001",0,0,28016.234120285997,0.933238746627292,525.00902029361794,1.9662012262031601E-2,144.44773520039919,6.3643518044251696E-2,54.954000075750415,"<0.001",-378.34117818363734,1.4294922828814101,650.01311287264355,4.82384757951764E-2,31.666720236340357,"<0.044",0,4.0065902888914096,766.69475361767593,1.1131024848152499E-2,11.66668040057003,5.4957340314032897E-2,11.666683182120076,4.8576445648628903E-3,23.333397358654977,7.7162477173253197E-3,6.6666768987972,"<0.000",0,"<0.000",0,"<0.000",0,3.2780600297453999E-2,153.33424657581875,4.4061758558541801E-3,43.333484878140418,"<0.000",0,27.711589654231201,28926.537715243398,2.8927108458065599E-2,40.000146428741814,4.8460745773452997E-3,33.33345085382286,1.2579195965588199,5230.8783196775366,2.0944899678836601E-2,133.33414196966874,0.49735640970474099,3283.6865762320176,7.6112348626790403E-4,6.6666768987972]
[null,false,"109SMPL.d","2019-10-22T14:06:57","Sample","2% Nitric acid rinses ",false,false,false,"<0.169",0,384.75245496786198,510.00802869586101,47446.748437695802,28942.899959342187,196.454837283638,30.00006869435208,68161.840325755998,356.67122627291241,4196.75845523514,17992.466629376508,735.95555068461795,196.66837399202026,7.1054273576010003E-15,706.68318649108426,9.7721994827951004,14.666682561238607,1.4511302117895,83.333723098033815,42.314323225096402,2846.8950991795195,37.031261570674801,526.67521645612692,362.58078410746202,10303.209153634512,1.25931969911463,91.667081738505544,16.507193737211502,336.67042548452179,17.2961861984198,466.67683372917645,105.013320798808,566.69716512231651,2.5284808649104198,8.3333391696213805,3.7889626393722402,15.000015671054291,"<0.258",0,10.822386620394401,1.666667833924276,2489.2843854212101,156.66764979559704,14.432068212881701,50.00059517722665,"<0.276",0,0.107405715466559,3.3333380520343074,1.2263352556201701,13.333361744880252,1.13678383754013,13.333359758058791,"<0.092",0,0,303.33662852640009,6.0389585506569103,36.666757514078057,"<0.032",0,0.21169968896274899,1.9854169738479102,"<0.072",-15.000021234154385,5.1758826066236701,25.00003539025731,6.83412263217458,46.66680584351014,"<4.059",0,788.64057173968399,1300.0952375453141,0.89744030967614397,8.3333471169072322,0.83243389347624297,1.666667833924276,0.128186434335709,6.6666768987972,"<0.107",0,"<0.046",0,7.7673734696700095E-2,6.6666768987972,"<0.034",0,6.63662584330654E-2,3.3333384493986,"<0.010",0,9.0533026055303195E-2,3.3333384493986,84.709560906638202,950.04382667540119,1.81979143450378,13.333364923794592,0.30026414290019499,20.000064074992178,1.40720886167472,63.333664586141744,0.42125792564335102,23.333412855862388,"<0.015",0,"<0.011",0]
[null,false,"110SMPL.d","2019-10-22T14:08:29","Sample","Blank control6",false,false,false,-0.71223130746462404,33.333787073701828,-0.32689129981103499,26.666722695031922,-41.8212412720275,32418.747889318842,-0.30922673109827098,40.000128149984356,732.30555850504095,443.33959554623175,19.5033760134787,18342.775373371886,-1.87529016895731,200.00154832996603,0,24349.904422314401,4.5441030106041103E-3,10.000007629394531,4.8360186908304702E-3,383.33840116850934,3.6616422601480798E-2,2823.5613592928444,3.6810740201091902E-2,420.00596374264416,0.584150931303558,12104.266373660161,-1.08212336288338E-3,88.333599045868723,4.3983716691303798E-2,616.67847032330735,-5.3781362800300898E-3,733.35592001727036,0.445390029633631,1300.0878942529137,-5.6520987092785503E-3,13.333354592322985,-6.1804811515542204E-3,16.666691054900124,"<0.001",0,2.5926220786887202E-3,3.3333380520343074,-3.6418454877821702,520.00874062366177,-2.9282634140443499E-2,133.33523968859802,-1.2812173995027E-2,13.333361744880252,"<0.000",0,1.8831894543406E-3,30.000090152023876,-4.0438626685527501E-3,26.666710774103148,"<0.000",0,0,258309.89198421602,3.0677149098487101E-2,176.66789109506371,-2.92067947888613E-2,44.444771276587268,"<0.000",-1.2716697229196174,"<0.000",-24.66672884424435,-2.6955795281847699E-2,46.666778822738252,1.1673246294258901E-3,18.333378757039032,"<0.005",0,-0.75019098027478504,400.01084903900238,-1.45363761520707E-3,8.3333447327214767,-6.8242577112667497E-3,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.3010217901129701E-4,13.333363731701716,4.80201879684069E-4,70.000264396266175,1.0876006461429201E-6,3.3333384493986,0.37826917102246199,6634.7662819248844,-2.9565623544152E-2,10.000015348195801,-4.06153294398568E-4,10.000015348195801,-2.9209124953830201E-3,56.6670107344735,-5.3125898768027896E-3,6.6666768987972,-2.7376025538340801E-2,56.66692808270065,"<0.000",0]
[null,false,"111SMPL.d","2019-10-22T14:09:59","Sample","Reagent blank14",false,false,false,-0.360837614378685,100.00226721172633,-0.145702187890727,33.333398799100536,25.992783987235999,32258.45736436361,-1.31460132658789,33.333406746386387,474.936302295953,380.00478814038456,18.051829951319299,17308.439514477937,16.434274336394999,346.67100389792017,3.5527136788005001E-15,23241.719217473477,7.8518792850823693E-2,16.66668661435445,3.1188574338760999E-2,506.67494972534547,2.57269561234936E-2,2610.1964434455367,5.4670928621036098E-2,443.34034577001609,0.35615087033910803,10370.07901800352,-2.0936701007335501E-3,68.333530997233638,1.03444973460456E-2,416.67240793497757,-1.4697979716712099E-2,633.34946458574427,0.33843574664003201,1066.7371352529383,-9.7477556305058596E-4,18.333358094095814,-1.4774905742316699E-3,23.333368748425908,"<0.001",0,"<0.014",0,-3.4479003099559402,510.00955457474453,-1.8632297940312199E-2,150.00178950532305,1.3306634914302801E-3,43.333461433647159,1.7143893282441899E-4,3.3333380520343074,1.9882637142439701E-3,30.000080615280854,-3.0358106068581201E-3,34.634750925221773,3.8489989363285003E-4,3.333335667848552,0,253281.20065289934,1.39220995888395E-2,88.333638782297996,-2.8312978062817799E-2,66.667208572240867,6.8655761450075905E-4,3.1370834342080092,"<0.000",-19.000034525989967,-3.0475723171194201E-2,31.666724209983283,1.22993780324097E-3,18.333367630838836,"<0.005",0,-0.64752952789874796,533.35108359694425,-2.0793386743616399E-3,3.3333384493986,-1.74760048143044E-3,11.666685168941541,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.72752433401794E-4,43.333503951626462,3.2008364342605098E-6,3.3333384493986,-2.4254869425475199E-2,2750.2695698351322,-3.0087454747653099E-2,6.6666768987972,"<0.000",0,-2.89146881501117E-3,56.666870862242519,-4.7256962908081897E-3,33.333459198473008,-2.7297456655184099E-2,60.000392496579991,4.2095137204739502E-5,3.3333384493986]
[null,false,"112SMPL.d","2019-10-22T14:11:33","Sample","QC1-4",false,false,false,220.83419766923899,43886.255863746752,7.7281380218462701,383.33808327707521,40.304251592421899,33190.084954647944,563.26944297702096,2810.2211088405879,469.24998667406999,386.67149285468219,663.88900653977396,108222.78480829066,504.25098677805897,4587.2559944700797,0,23708.957980042327,1.4097343294491201E-2,10.666676839192709,0.51412493669433001,3123.619236494811,0.59189692556585105,7591.7119230074895,0.56447664855525503,1626.7456829536106,22.762159644104099,134406.75349250631,0.26271126377944798,4440.5578968839154,1.0755266091070901,6054.3669752125888,5.4254122590320897,43403.207350345612,55.640468881839098,97971.207837613605,0.52486893270367696,635.01199411831305,2.8258616336118902E-3,30.00005478660184,2.8784943300837198,4727.3101102216524,9.7950276705313293,688.34820039548174,-5.46589019812399,446.673679053679,26.759193709314602,51241.908900879767,-9.4573379596278994E-3,20.000047385691889,"<0.000",0,1.5682928893836501E-3,26.666733026503525,-3.72456594047671E-3,28.333429073292578,"<0.000",0,0,248108.79838622417,7.2338267023185103E-3,53.333456516264029,0.93543577052664895,24450.977787190346,0.259465966517681,1990.8454771736015,0.26096942342815199,2857.2488133323864,-2.6485849002896002E-2,46.666775643823911,7.0000696736755897E-4,15.000035141904625,"<0.005",0,-0.68895096783876697,466.68047358623215,0.53514902385946805,4072.1647545869437,2.6964792074085602,4383.9106605860989,"<0.000",0,"<0.000",0,7.4450364886196599E-4,13.333363731701716,"<0.000",0,"<0.000",0,-2.1717949149499701E-4,13.333374460537614,2.2507364830352499E-4,46.66682014862468,5.4638342976612696E-6,3.3333384493986,-0.120251570707163,1816.7965341388406,-2.9498594202091302E-2,10.000026474395993,0.25515908519085201,13906.087439267525,2.6110162742254701,95979.347528129598,0.29852553525505099,13595.83435144274,0.224850705060963,14006.339309518449,5.8537421534485802E-2,4540.6735570131705]
[null,false,"113SMPL.d","2019-10-22T14:13:03","Sample","QC2-4",false,false,false,0.11168136862635,200.00453442345267,-0.17755750255328001,33.333451251187157,468.09417768486202,42602.906266208891,-3.59032542270768,23.333422789969699,552.21313266486902,410.00518187872791,20.495367761667801,18409.540715033476,85.340770309392397,973.36251397884132,0,24249.658655579009,4.9364452809932704,503.34051513671875,1.3192441446383599E-2,426.67273839305727,4.57313104375087E-2,2880.241249398769,2.3044792608691699E-2,383.33803241444576,0.25376969451440101,10102.985607899403,-1.01267312854121E-3,88.33365944524121,2.64408496691995E-2,513.34187060581837,2.3785777350174599E-3,783.35248790724165,0.13936119281144199,733.35870156742203,-6.8394536118907799E-3,11.666679605841445,0.61530124436278399,985.02834123827506,"<0.001",0,"<0.014",0,-0.94293627548470804,590.01081223273059,-1.8651289877156699E-2,150.00179149214458,5.1233433591773503,11013.489826120587,"<0.000",0,2.6381149764238501,18790.160852817284,0.66835984768944601,5672.6016407320649,"<0.000",0,0,247943.83094076568,0.544577383656354,2718.5512900568351,7.3365068803705304E-3,966.69917677912269,"<0.000",-33.708877292040498,"<0.000",-2376.9586358324887,0.995164331159898,4050.4867967393971,0.50254387792545696,2815.2359606538816,0.36104432002066,73.33352406819661,-0.76302896678680399,366.68194954628137,-1.1892525295235201E-3,10.000012566645752,-6.7416555531316603E-3,3.3333384493986,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,0.43043268086261399,17549.580630417953,5.9967423464862301E-5,33.333404362200632,0.43408653433532401,13052.011524259875,9.04952796357958,85561.942423526154,0.51347373806032004,3206.9922414373523,-4.6011096517939898E-4,6.6666880249973923,-2.2298025630835199E-4,153.33444048959353,-3.6658882445908701E-3,80.000369151427805,3.0862013742802102E-2,3273.6854574031809,8.6002598672536398E-5,6.6666768987972]
[null,false,"114SMPL.d","2019-10-22T14:14:33","Sample","QC3-4",false,false,false,-0.69662567069227699,33.333787073701828,0.20951111671669301,46.666841606296479,122.998235591191,32659.146648757454,-2.44683643717857,26.666730642317773,515.32485523185403,370.00415250652799,35.684440847987801,18876.617344939601,7.22298976301997,256.66890715555979,0,22233.796869211441,5.0172108503836001E-2,13.333340644836426,2.5066264137036401E-2,453.33970983813634,5.3945730862118499E-2,2726.8790283767526,4.8407754608834501E-2,410.00509286912643,-7.2451439839742199E-2,7668.4744959480331,-7.2232336310132296E-4,86.666976908838095,1.7435230289790399E-2,433.3388200401443,2.5854719784470098E-3,733.35609485756561,0.16739655112184601,733.36621970011799,-4.7023226223733196E-3,13.333350221315769,1.53241154689739E-3,26.666715145110363,5.6961056760779401E-4,6.6666856408116368,3.1390308205810402E-2,5.0000058859585836,303.01377820497498,9429.206605913143,-4.2052186782034297E-2,100.00119432809639,-1.3960186663748901E-2,10.000014156102923,"<0.000",0,3.6669977257258201E-3,40.000104308126801,9.3268210597138197E-3,132.96861416183987,4.0501856657794801E-4,3.3333384493986,0,240627.38573728921,5.14426669630843E-3,41.666756644843666,-2.0932705348014499E-2,244.44811625621313,4.8594269360149102E-4,1.4858315472926158,"<0.000",-22.666728874046672,-2.6992283159671799E-2,43.333445539075456,-1.6786526964371899E-3,1.666667833924276,"<0.005",0,22.033046859313501,30127.756410712522,-6.9528613936351697E-4,13.3333537975944,-3.1627141703292498E-4,13.3333537975944,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,7.7218097833524797E-3,326.67106141640153,2.8570077417165197E-4,50.00014786918738,3.09392063565517E-3,93.333703875536159,0.53109808366759303,7535.1646836014634,0.14439541843804601,1003.3661502098121,"<0.000",0,-2.0626638865900901E-3,83.333746939891384,-5.1483781087286401E-3,13.333364923794592,-2.60610136121187E-2,123.33426132796812,"<0.000",0]
[null,false,"115SMPL.d","2019-10-22T14:16:02","Sample","QC1-5",false,false,false,212.31996034231099,41547.596143504888,9.84818876876618,470.00663056959758,15.937373789489101,32201.682462865876,529.56066733597299,2603.5263690820379,781.89936469318502,433.33932866643869,628.06428087334496,101601.4121830779,3897.31839040351,33507.781076215986,7.1054273576010003E-15,23341.916969886664,0.33445486061326302,41.333394368489586,0.51137094862321597,3053.5987907608528,0.55993851833228003,7174.7868803920319,1.23427314279694,3110.2769515373766,22.5232006656747,130275.21043140911,0.24920547567498799,4123.813008211856,1.3333462360202499,7244.8196348823158,5.6946674508700097,44405.706762401656,71.096719361902899,121852.71268455788,0.51198331185005597,603.34417079869638,4.41917669119253E-5,25.000043337543158,2.7173624502857399,4333.8711919307034,9.6822987374785292,660.01298556705831,0.75402697140602704,616.67863562685295,24.460581049793099,45292.151891264737,2.9446795015191398,6091.0710762281551,1.09555064275132E-2,203.33494519191564,7.5439846329458493E-2,530.00877862162224,1.8763274936661301E-2,208.33491014935211,"<0.000",0,0,237642.18395885554,2.4483175316624499E-3,28.333374237020205,0.98074726572857496,24517.646489852381,0.24208372943072301,1778.9742083664612,0.24982458895944901,2616.2116611369697,-1.44251224103117E-2,90.000321318701083,1.6711959254226E-2,100.0004276633199,"<0.005",0,3.7724873660797198,6201.3894513559962,0.48965976627843999,3570.3824337763231,2.7630014394700502,4302.2246673371692,1.7098950714025701E-2,696.68342823758576,6.8226694046453502E-3,50.000155021744639,1.5545879656211499E-3,26.666717926660411,1.9829244899880499E-4,13.333374460537614,1.0044683364904101E-3,23.333389808733415,2.9573294559228701E-3,136.66748796898889,-8.4297409154220192E-6,26.666748921075229,5.8881864984888701E-4,20.000050961970523,0.24463250795716501,4934.1336399495667,2.7871755315490701E-2,333.33765804761157,0.24035280612849799,12548.316056505102,2.5591734445641001,90108.159464575452,0.28208831840464399,12318.130513060134,0.23896254163606501,14163.139437390279,5.5326373443955697E-2,4110.5531207141312]
[null,false,"116SMPL.d","2019-10-22T14:17:35","Sample","QC2-5",false,false,false,0.99733178808766398,366.67798385049372,0.22769007186413101,50.000143895544454,539.26888277892601,42576.099625983341,-2.74487319950613,26.666758457818251,494.99520476970901,386.67152464382565,15.7727183081162,17141.564815923779,411.19160444916201,3737.0585584085402,0,23441.968237927034,5.3837785080682199,529.34224446614587,3.44107305554966E-4,343.33704118921793,4.1110932056147402E-2,2736.877909249893,0.102829490583999,550.01006260499264,0.55213184607064403,11337.209827787898,4.7760755644806804E-3,178.33435428637731,0.13824467067773599,1070.0346723690429,0.47682799214333199,4383.8852908585659,2.31036195144249,4400.5974799528158,-6.4064647116603502E-3,11.666679605841443,0.61359347707978795,940.02590050261369,"<0.001",0,8.0004149074589206E-2,8.3333439379928915,-3.6546415553378799,483.34058677145958,-3.2973762524653102E-2,116.66797262929816,5.17333116834976,10583.244064302582,2.1747676640112498E-3,40.000104308126801,2.7120346150136898,18349.708792323243,0.66146491712212097,5324.157228606583,"<0.000",0,0,234481.97567261299,0.59878729204688397,2825.2356553290933,1.3248924098671599E-2,1055.5982189883391,"<0.000",-0.30404097659740614,"<0.000",-2203.9200037928417,0.98118408712764604,3778.7558437572538,0.49423449624964799,2618.5373551864818,0.44496920028984099,85.333607991536454,0.19559925787013399,1566.7766285463911,-1.8202357898855601E-4,16.666711320479045,4.3436566998494103E-2,80.000288089112118,2.5703387698859902E-3,103.33378002046874,2.7658468726931701E-3,20.000050167241938,5.9082775035855105E-4,10.000024884938822,1.0048245820911E-4,6.6666872302688072,"<0.000",0,0.42441420730547103,16364.975342080366,1.27468500209649E-4,36.66679645577873,0.43754034370841299,12441.552142080383,9.1839972011011,82077.856245390329,0.70581211073970695,4103.865498743472,-6.3878665806584007E-5,26.666719516117581,1.31846878891431E-2,610.01346950709615,-1.5127689324286901E-3,166.66793892779043,2.4440563038097499E-2,2760.2505081854688,4.0923105233620202E-4,30.000090946752465]
[null,false,"117SMPL.d","2019-10-22T14:19:04","Sample","QC3-5",false,false,false,0.178140591470733,200.00453442345267,-0.363336492476386,23.333375106254589,36.629992700653197,31737.535334365541,-6.1076430924259597,10.00002289811736,262.517567410521,336.67026653880475,27.516098098323202,18192.616678609327,11.356083658498999,296.66978393987142,3.5527136788005001E-15,22721.105618460391,4.7326617211941799E-2,13.333343187967936,3.49488933878063E-3,350.00433082375429,7.0607302673759698E-2,2910.2431817316537,5.7825306157468302E-2,436.67242035261165,0.47731751135601602,10670.037317351342,-1.0969950871041699E-3,81.666946783657679,1.14189769648984E-2,406.67299300422775,-8.0643873415087598E-3,658.3519288032436,0.22513035478612001,833.38450069343173,-6.2272829204975096E-3,11.666689539948758,-1.21099918682966E-2,6.6666737198828594,2.7070295606891502E-3,10.000014156102923,6.58215951624416E-3,3.3333380520343074,279.40181851843801,8698.8314540361625,-5.0769708919601399E-2,83.335359767124658,-2.3726490019743401E-3,33.333428204058187,"<0.000",0,8.7068234428997002E-3,73.333672384415976,7.97374806951422E-3,119.81789726401523,2.0703313758864701E-4,1.666667833924276,0,234911.51993704555,2.1629423957364902E-3,26.666728655496311,-1.7907465624321699E-2,311.11726396627404,8.1999273026621205E-4,3.8800019679467388,"<0.000",-27.000067626435538,-2.4021825420287201E-2,53.333456516264029,-7.2537937781455802E-4,6.6666768987972,"<0.005",0,25.265847728206602,33533.879330269781,-1.11595514202198E-3,10.000012566645752,-1.0984556792576799E-4,13.333353400230109,1.6552485994745299E-4,6.666686435540222,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,6.3676971672915299E-3,266.66948556895824,-4.8083725671612597E-5,23.333389411369126,3.4058536200436001E-3,100.00064382949506,0.85435456512576702,10153.256706842061,0.16838820506556501,1113.375431000559,-6.4826693681489195E-5,26.66676243146118,1.4480290764552399E-3,203.33498333888772,-4.8260373376823604E-3,26.666761636732591,-2.6259664491688001E-2,110.00070090094158,"<0.000",0]
[null,false,"118SMPL.d","2019-10-22T14:20:33","Sample","Reagent blank15",false,false,false,-0.525779980839709,66.668480138024492,-0.28677348600082497,26.666776736575713,7.4934057348120904,31276.751649825499,1.64428377802022,46.666860679782509,682.27031633520096,406.67264332365022,11.5245376105317,16090.667067011205,6.8558293983177396,260.00223408139396,0,22787.847153924198,9.6586133402154106E-2,18.000016848246258,-3.7360115997066901E-3,313.33681275442024,7.3894343886618197E-2,2943.5828215343331,3.1881295569956399E-2,380.00515689444802,4.61543453394559E-2,8402.1839748314396,1.6410291392818599E-4,101.66704185306467,2.05895306220754E-2,453.33984971036739,-1.0324545823637401E-2,641.68116096439564,-0.13385486199791699,233.33755855768428,-3.2438350040899901E-3,15.000022426247261,1.52333961527837E-3,26.666729450224892,"<0.001",0,6.6372300435196697E-3,3.3333380520343074,-2.2864359224461301,516.67526746776809,-7.8298776188195598E-2,33.333568523332765,-1.22451708346097E-2,13.333361744880252,"<0.000",0,3.2723066673198899E-3,36.66679486632156,4.2627390375399402E-4,59.634824157593279,4.1580526506367402E-4,3.3333384493986,0,233857.11891109069,2.9010399916686099E-3,30.000042468308767,-2.2042019461229901E-2,211.11375962700404,6.7161775798470897E-4,2.788334910050049,"<0.000",-12.000021278857867,-3.2974993137340598E-2,20.000035464763112,5.4703967376673504E-4,13.333363334337422,"<0.005",0,4.8222792383808502,7435.0792321456292,5.7660677008923897E-5,18.333360080917277,"<0.001",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,1.79893702852453E-3,90.000381718073569,2.1648455276230099E-4,43.333523819841083,1.2978466348114099E-4,6.6666872302688072,0.32923932759406399,5584.3470232166919,4.3833829666230902E-2,416.67281483601317,-4.5272078658495699E-4,6.6666880249973923,-2.8621981317147401E-3,53.333559830980107,-4.5859879237959101E-3,36.666778177021271,-2.7337072636902101E-2,53.333493073778946,4.55915430007434E-5,3.3333384493986]
[null,false,"119SMPL.d","2019-10-22T14:22:04","Sample","Acid rinses ",false,false,false,19.7749441035361,200.00815838593599,48.225524839986001,106.6671037673885,25001.412495394499,25784.851503796603,407.40901174105301,100.00047216812065,37238.885064858703,326.67040020221867,2779.5099487952102,19811.037124940092,493.78212487036501,220.00171969331723,-1.7763568394002501E-14,1160.041740933432,5.3792523967011396,12.666678746541342,1.7981733674142899,156.66756714382419,29.027312227514901,2970.2546757918394,15.2999860414506,330.00320578610086,251.665682307173,10770.28781804589,0.89911649326493104,98.333719695602056,17.914379339350099,546.67580107834237,7.1772207134643802,291.6700507825376,1526.94059620024,12271.118946969684,2.7138030292069701,13.333345055579963,16.1568950500752,95.000332370395483,3.4634163854967799,20.000093479949829,"<4.380",0,22878.8084226707,2126.7968433082574,25.968082261341898,133.33523968859802,42.062554303675299,226.66838626064282,"<0.022",0,1.03187456733641,16.666690260171539,4.7129893228435504,81.667002414658626,"<0.062",0,0,450.00654185801932,5.3655163121560898,48.33343115945511,18.1277939893018,833.35880446822091,0.12836341082524599,1.7843682055023091,"<0.049",-296.3383225946875,73.304644704261506,521.67505122708701,4.11231792795608,41.666735981900452,"<2.736",0.66666690508524573,1008.99199637615,2466.9102902268187,0.7254465515595,10.000009785095704,3.9605659707422398,11.666683182120076,2.46260839710766,190.0013418495455,1.2009850119635099,16.666710923114753,0.102619490444171,3.3333384493986,"<0.008",0,"<0.023",0,0.99171511032217496,73.333659668758614,"<0.007",0,"<0.018",0,1161.72635671177,19261.745120474963,108.21061075748101,1156.712247778293,6.7014661953781404E-2,6.6666880249973923,1.29768667309709,86.667089760297188,0.28217941905273403,23.333380669354685,0.636134280732221,66.666933496789127,"<0.007",0]
[null,false,"120SMPL.d","2019-10-22T14:23:40","Sample","Base rinse",false,false,false,"<0.106",0,25.964312794621499,56.666891525185726,29723.740470234501,29483.778043730388,405.43768133804099,96.667145242286551,47004.257740765999,396.67174066098914,2809.8231659356302,19443.948798863934,492.94186680306598,213.33529472348152,3.5527136788005001E-15,1126.7178431145671,2.1814685247045902,10.666671752929688,0.38389415416148898,116.66742150981095,6.3369165763770097,2803.559935437243,3.4012079990165298,400.00560164443249,40.3108881359869,11403.863780286842,0.13813931097009999,118.33384337523813,1.8086306946801201,496.67488629607033,3.0602138429223298,1250.0500322641192,59.316013132922699,5300.8454553971906,0.261161727715012,16.666691054900124,8.8291170946291395E-2,8.3333439379928915,3.5420256241473297E-2,3.3333380520343074,0.41142229460402402,1.666667833924276,354.39696149783902,603.34504817905452,1.5317401998098801,166.66952148086253,0.32860636564312701,40.000094771383779,"<0.001",0,1.1527144484719001,450.00656728933404,0.15653217061176899,76.666921426848759,"<0.002",0,0,14235.925080836934,3.2329253974596699,921.69236399195051,-2.3240695479596198E-2,11.111166990466836,9.2655758740818106E-2,40.710643218102163,"<0.002",-351.00674336641907,2.6726158608769599,610.01126969837242,7.6047037324552896E-2,25.000046119093209,"<0.086",0,19.673545333310599,1600.1006166162726,3.2006316771380702E-2,15.000018849968628,0.116674125377104,11.666680797934321,8.1941393723593894E-3,20.000070432820859,7.5927697480548803E-3,3.3333384493986,"<0.001",0,1.6550584998278101E-3,6.6666768987972,"<0.001",0,0.122068838545695,286.67004475985897,1.1175514622247701E-2,53.333547910051323,1.8259566288518999E-3,3.3333384493986,94.640904414653306,49794.329929721775,0.34349062931333402,126.66749765474351,1.1169583618451599E-2,36.666755527256598,2.54476839397784,5367.5863864880839,4.0006095047733801E-2,116.66744376221135,0.35289670357011299,1210.0507683051017,4.4936649063432197E-3,20.000041822591793]
[null,false,"121SMPL.d","2019-10-22T14:25:08","Sample","2% Nitric acid rinses ",false,false,false,6.33916498151219,33.333787073701828,430.45819421440802,446.67247106622949,57144.063334452301,27140.185956256799,49.860171905280701,6.6666761040686149,105080.29677254301,426.67290369660304,5654.67188086021,18859.97509196345,963.08902846567605,200.0014211733924,2.8421709430404001E-14,553.34555755040708,7.23584845945014,7.3333384195963545,1.5404618015226901,56.666858146585156,64.406610731959603,2713.5462547310253,52.674595456059002,463.34103847531918,574.54075365938695,10002.96831236957,2.4844259248688401,110.00045930345169,33.736531688533802,416.67299285521614,28.6689117501693,466.67441775418757,205.19748279412801,666.7089452355616,5.9081660585828901,11.666681990027199,0.68785668509395004,1.666667833924276,"<0.431",0,18.117131810818599,1.666667833924276,7038.3368062089003,263.33584075169006,48.629647508329697,100.00095193586888,"<0.463",0,0.36080209467760799,6.6666761040686149,5.6762611934953204,36.666785329578538,3.1208888043565501,21.666694159308662,"<0.156",0,0,180.00135789312881,21.2876324703814,76.666902353362687,"<0.054",0,0.427289208574683,2.3795848952792871,"<0.122",-21.000042602419221,12.2634735231736,35.000071004032037,13.986759609230001,56.66681523124155,"<6.841",0,2319.2848673006301,2266.8619026784031,2.7273490547214099,15.000023618340142,"<0.851",0,"<0.032",0,"<0.180",0,"<0.077",0,"<0.020",0,"<0.057",0,1.3525446339932401,40.000131328898696,"<0.017",0,"<0.046",0,130.40019895828101,866.70046200430431,14.785341078690699,63.333569218711524,0.33740658032417198,13.333364923794592,2.8752397889571202,76.667039046679349,0.405454309294632,13.3333537975944,0.13776459180479,6.6666880249973923,0.118464936171698,6.6666880249973923]
[null,false,"122SMPL.d","2019-10-22T14:45:38","Sample","Reagent Blank",false,false,false,-0.53232594803660305,66.668480138024492,-0.68199996695471898,10.000014156102923,-48.739274433783102,30782.560576460353,-4.0770522280926897,20.000065664449348,695.74913289663004,416.67236978800543,11.4802266080818,16390.883239899948,6.2736554238495801,260.00297794734973,0,23221.637023762309,5.8177682061424998E-2,14.666679382324219,-4.26809007898839E-3,316.67018418505512,3.5395983426991601E-2,2673.5363879769584,5.5062945942647698E-2,440.00602066507901,0.92806272840971604,13338.668104132048,-8.9752177622360295E-4,86.666964193180732,2.39071087939445E-2,480.00724077213937,6.3676396147995902E-3,783.35318726842286,0.35111521865996198,1066.7079846073486,-2.0727252335593898E-3,16.666682710249979,1.0018617148631901E-2,40.000117023784163,4.8298130211226701E-4,6.6666761040686149,5.52148801155594E-3,3.333335667848552,-3.36213544120774,496.67468284555252,-3.3750762240088002E-2,116.6686958323376,-1.0786343883989901E-2,16.666709333657582,3.5708207821780403E-4,6.6666761040686149,3.62663591510808E-3,40.000113844869823,8.1643241979597998E-5,58.333491409815984,"<0.000",0,0,239924.0003311038,3.7464105899667199E-2,196.66789715486917,-2.9077501468818798E-2,44.44487459130719,4.2342063848925402E-4,1.0170780924398266,"<0.000",-46.000149359303471,-1.5971499308459799E-2,85.000259404377246,1.7430015764675799E-4,11.666685566305832,"<0.005",0,-0.52352914400550599,666.69624547229762,"<0.000",0,-5.6092775882261503E-3,5.0000062833228762,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,-2.0601641479219999E-4,13.333385586737807,8.00647082976802E-4,90.000562916190972,4.6762233225744302E-4,16.666712512571923,0.60717516072896505,8185.4664406905531,-2.5930381568996401E-2,30.00006869435208,-4.56010199209381E-4,6.6666768987972,-1.77400432933134E-3,93.333778580023179,-4.8391939442947399E-3,26.666740973789377,-2.7736840470864201E-2,33.333452045915742,"<0.000",0]
[null,false,"123SMPL.d","2019-10-22T14:47:07","Sample","NMS 55",false,false,false,0.24734554317597399,266.67574048062579,17.008416857333401,953.36014916459499,773.46377608940099,57189.341595729078,-0.880364181739272,43.333569516734734,-3.5648087536715098,366.67094002161019,-3.8310707547938301,17491.974267929218,-1.89175177023081,233.335440655518,0,28429.114973508593,5.2899339066982097E-2,17.333347002665203,1.312039372026E-2,500.00805427618735,-9.3933266583033394E-3,2803.556883679476,3.5160841678584402E-2,483.34080293763481,42.124364213269899,288548.07773861458,1.0814616768156299E-3,145.00073216854435,9.6226661707639494E-2,1043.3656918501008,1.4442002293714999E-2,1033.3723617105466,100.042324691657,209433.40311870133,-5.9179782547140302E-3,15.000015273689996,1.2075282981889801,2243.4800894536561,8.2714068255606201E-3,23.333394974469218,-2.34091312379066E-2,1.666667833924276,-2.8074674678865699,626.67886435967387,-5.2508135209663E-2,100.00214402879105,-1.09227758273568E-2,20.000037848948867,"<0.000",0,6.0220922026211897E-2,523.34377780541115,-1.1734228371651099E-3,58.730376729015916,1.20245807554756,12055.893781941724,0,292507.69095379574,1.0818742701901201,6351.1825632376094,-1.4877404504624299E-2,477.78712386907165,7.8372861117610899E-4,4.5029055925771679,"<0.000",-37.00027760366185,-1.06281452689172E-2,128.33396720388581,8.9005480482239798E-3,71.666910375154359,0.79387765813210198,189.33447265625,-0.510576531839814,833.37825412651966,-1.9505005725344699E-3,5.0000062833228762,2.5409141676499699E-3,21.666706477601732,5.3172963670446497E-4,26.666783889132976,"<0.000",0,"<0.000",0,"<0.000",0,"<0.000",0,2.1887081470795699E-4,36.66691804925226,8.0354036475785201E-4,110.00082169968653,0.86170023179039701,30562.510937812262,15.620036541248499,171733.25172265188,3.7038468655858199,25944.442158970603,5.0915009974236204E-4,70.000300953781093,6.0390201324401897E-4,216.66843091445517,4.4692915296460399E-3,523.34356163923599,-2.6622118797753599E-2,113.33401352166116,1.4580014296477001E-4,13.333364923794592]

{"sheet":"ConcTable","index":2,"n_rows":124,"n_cols":58,"rows_kept":124,"cut":false,"hidden":false,"merged":[]}
["Data File","Name","Li_7","Be_9","B_11","Al_27","24  Mg","27  Al","28  Si","31  P","34  S","39  K","44  Ca","49  Ti","51  V","52  Cr","55  Mn","56  Fe","59  Co","60  Ni","65  Cu","66  Zn","71  Ga","74  Ge","75  As","78  Se","79  Br","85  Rb","88  Sr","89  Y","90  Zr","98  Mo","101  Ru","105  Pd","109  Ag","114  Cd","115  In","118  Sn","121  Sb","125  Te","127  I","133  Cs","137  Ba","140  Ce","146  Nd","156  Gd","159  Tb","166  Er","184  W","193  Ir","195  Pt","197  Au","202  Hg","205  Tl","208  Pb","209  Bi","232  Th","238  U"]
[{"$f":"SQ!C3","v":"001SMPL.d"},{"$f":"SQ!F3","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A2,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A2,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A2,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A2,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.002"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":3.6567695034817498},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":891.36190962184696},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.55555634200571802},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":849.68450938158105},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":77.176073623918896},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":19.9746438949173},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.8287087693035198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.6207544982561303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.117981180044366},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":9.8959606735844696E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.90403326478725599},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":7.9969448907705101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.38841190358854E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.4308793918635002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":7.1126285550350103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.7937208632075999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.53585867885067E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1.7154649740647201},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.8550858553830899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":2.6062101801452802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.6328311990672501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":5.9943075314910599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.4702019460851199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.2483974038567199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":7.8617516061633405E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":7.2982138130533905E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.7634017385859498E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.0268299467311201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.4555863367139302E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.2438809648298001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":4.04092376138687E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":5.6680493558219696E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.1091781303237698E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.0709293902157798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.20449426210491E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A2,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C4","v":"002SMPL.d"},{"$f":"SQ!F4","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A3,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A3,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A3,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A3,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.12126604149300201},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":3.2338762513137298},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":910.25832704624395},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.8518559303547999},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":666.675071505814},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":79.848084656795606},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":17.5058666370604},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.2430645354592801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":9.5259375584959092E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.124072864083328},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":9.5487265356074796E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.89134909668528695},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.4928421296273197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.5750591845245298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.9480903260690401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":5.0804483449422297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.4200294265425999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.3544648728298699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.5358568650658001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1.28659790403408},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":3.92755978263611E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.3031032259711901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.08984935972018E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.6160965057211899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":5.0406866036726804E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":6.6443623787055001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.13103074290694},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.9461914134977401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.75268034771719E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":6.6743923647380801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.6911198880982699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":4.59199608435596E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.7004167662275499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.7201137131357605E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.1609697350177901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.20449172174856E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A3,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C5","v":"003SMPL.d"},{"$f":"SQ!F5","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A4,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A4,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A4,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A4,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.002"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":2.9851132987144702},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":892.48451952705796},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.4814849915327299},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":712.42717853155796},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":79.661195512964596},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":16.608120287801398},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.3458387558131702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":8.0970454371164704E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.141065652203785},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":7.3785411543202603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.93360732125216395},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.9775101671299203E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":3.8125500152966299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":3.7067614478474301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1.4535675001192299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":4.0643575131326E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":7.4738388901196499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.53585867885067E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1.2865971201234701},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.7130627498595597E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":5.2124166320876696E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.4531387157860399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":7.7262392732444201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":3.3446732795692E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":7.9810928393357801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":6.3423429219636503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.13103074290694},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.94619129749591E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.1611022972144E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.1341370033093598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.6911198880982699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":3.1225418027561E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.70041669865924E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":3.0545911043169201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.4406436062514699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":5.1438165577696502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A4,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C6","v":"004SMPL.d"},{"$f":"SQ!F6","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A5,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A5,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A5,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A5,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.002"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":3.3582566363442599},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":926.29565757115802},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.29629958834908},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":725.49881196866204},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":73.2006220676314},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":14.588195978279099},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.2088069800719796E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.19074376758527E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.132088331793105},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.120661611216011},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.94205654301074604},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.76584276958827E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.2359059781515498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.5170671290644999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.1176987763948499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":7.4738310497676503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.70893362080928E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":3.8037871113835699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1.6439874926123701},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.4985994269721498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":2.6062101801452802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.5429839180667901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.6160965057211899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":7.9810952429182104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.1174609586015899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.18344242028136201},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.7029164790445699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.0268274006618699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.4555863367139302E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":4.2246118279338697E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.9676471167808397E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.2218356260647499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.8008094291697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.2044941745064101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A5,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C7","v":"005SMPL.d"},{"$f":"SQ!F7","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A6,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A6,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A6,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A6,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":6.0632196774693302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":3.1094923552434901},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":930.02417614940498},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.37037089467047901},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":921.58192775977602},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":76.866184949536702},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":10.3239332990896},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":3.3800960809498001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.81037406951421E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.136576983437772},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":9.8091431983936397E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.86177207572249304},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.0889553081547396E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.8292311699338203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":3.8791628181296403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1.4535675001192299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":5.0804512519951796E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.9895312082162998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1.1436421572187501},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.0695827945210593E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":7.2656623981345002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":9.6965817314772498E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.3894414684065799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":7.9810872309767893E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":4.8322611195781599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":5.2412452109864303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":7.2982138130533905E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":9.6520668520536904E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":1.46941697121966E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.70041480674659E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.4436732912845501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.8008056509426799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.6741554099372801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A6,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C8","v":"006SMPL.d"},{"$f":"SQ!F8","v":"deionized (DI) Water rinse bottle"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A7,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A7,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A7,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A7,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.12126439354938701},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.248757448807266},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":916.64229619299294},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.37037089467047901},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":764.71631774310197},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":80.283283174379505},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":13.914889587384},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.3458383528745002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.6681535428523E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.11798143931044699},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":9.8959547131200795E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.77726601844274001},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.6850677163350601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.9308985873061199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.4137042293099101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":2.9071758934479799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":6.0965391767518603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.8684600700000598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.3544648728298699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":0.21443264885548899},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":8.64060005923626E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.3604555992640798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":3.0201592844015602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.4596431686171099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.0268274006618699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":9.1839056366564403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.13361176307704E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.2218356260647499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":9.9044282679816901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A7,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C9","v":"007SMPL.d"},{"$f":"SQ!F9","v":"deionized (DI) Water rinse bottle"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A8,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A8,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A8,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A8,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":6.0632196774693302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.24875692096514601},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":941.92623745106903},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.18518544733523901},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":790.86095591369303},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":72.579325972354098},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":12.568285581023201},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.2088076708239796E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":8.5733547723141207E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.123431598783213},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.105904117304191},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.85755620527854404},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.4427350071738298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.5750576332102499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.6378616443893899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":5.0804483449422297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":6.7264537715979403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.5358568650658001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":0.35738802075704401},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":6.2840685906678906E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.6160965057211899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":4.2005719527371602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.13102996817149801},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.2163686981712799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.0268289919551499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":4.2078875317697298E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.2438809648298001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":2.0204603917815499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.83402656982363E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.8327554782521699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6207255624202299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.6741624178168698E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A8,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C10","v":"008SMPL.d"},{"$f":"SQ!F10","v":"Reagent blank1"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A9,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A9,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A9,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A9,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.18189823826769599},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.49751426894863698},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":853.05908905937804},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.5926047397982601},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":810.46765185634001},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":76.679407533940406},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":158.684039716055},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":6.7601934282782605E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.0016461611079203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.14202754403919199},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.148440276026788},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.0645877313251699},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.3312894809754603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.7776489080443503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.21034736479616101},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1.8026566805453299},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":6.0965377232253797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":6.7264528806488498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.70893362080928E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":1.2679284325329001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":11.0791944576057},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.3565086207546201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":0.24196540553264401},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":7.8186268122329503E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":6.5391075912794704E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":4.0726140395498399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":2.4036156892928702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.1482109047059299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.02442303871262E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.14766362433735E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.31447858633639603},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.2163686981712799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":0.10682198160973801},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.66504985089375E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":9.6520530447024504E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.0268289919551499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":7.5742211332552897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.2438774874516601E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.27737266729454801},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.2218356260647499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.37763733012802E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.54314727993116E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":5.7182482448983598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A9,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":3.2257864996790601E-4}]
[{"$f":"SQ!C11","v":"009SMPL.d"},{"$f":"SQ!F11","v":"Reagent blank2"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A10,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A10,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A10,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A10,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.24252878709877301},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.323384099857709},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":894.06752353457796},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.55555682767318704},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":830.07604321996803},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":71.461500133418696},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":19.525746427330301},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.2088065195706494E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.3345735264113098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.14106560328565701},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.12066154498862899},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.98431584704152697},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.2811815881748597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.2526010946172494E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":6.2067102619810702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.130823123136733},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":3.04826813484945E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.0926741978570099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.70893362080928E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":10.721791474380799},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.3565086207546201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":4.5869300791244599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":7.2656623981345002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.42217170644061E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.0060008386649402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.6883732925468199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.28827042344285903},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":9.7309544574473799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.6705405710110101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.75268034771719E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.4555863367139302E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.4702602120175296E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.4877583185361502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.16347979946803001},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.5344657687297601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.8327534390971299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.6821161338892799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.20449426210491E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":5.7182482448983598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A10,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C12","v":"010SMPL.d"},{"$f":"SQ!F12","v":"Reagent blank3"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A11,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A11,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A11,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A11,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":6.0632196774693302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.22388117489825801},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":879.10539880315798},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.1111131696789001},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":692.81944788345902},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":73.697969889211606},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":22.2189933943976},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.101402923297987},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":4.9535300554820402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.14042435999869901},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.104167910189246},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.84487574279893196},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.0081782519990599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.1342411786221002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.2584866535413499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":4.3607025003576799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":6.0965391767518603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.2705516366220301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.7089297456597398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":10.5073481575243},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.3565086207546201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":6.8804033207392297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.4531335189867599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":4.6866891524939402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.94041358661935E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.81209657872714E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.104824129484287},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":4.8654772287236899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.1611022972144E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":4.2277984098679298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":7.1534191875116204E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":5.6096969285166702E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.12674243314283701},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.40083339731848E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.1091781303237698E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.2605661796616899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.9393272011804598E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":5.7182482448983598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A11,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C13","v":"011SMPL.d"},{"$f":"SQ!F13","v":"Reagent blank4"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A12,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A12,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A12,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A12,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.12126604149300201},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.14925400786731599},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":874.45601744644},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":3.7037159226557002},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":934.656091862023},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":69.597199021227894},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":17.9547326415219},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.6916773965010005E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.0969137514523897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.15805845499967999},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":8.7674589969927594E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.79839900551825205},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.5235156839608702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.12607027743706101},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":6.1205026301673397E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":4.3607433935672103E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":4.0643618737120297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.3452904692158801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.3544648728298699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":11.508056006179},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.57100574716975E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":5.73367162116337E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.4531335189867599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":2.4241468777711499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":8.9914587962204703E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.5962203705539701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.20806472184682E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.55033733240483296},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.21636904617679E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.1611030446504401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.1341370033093598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.4555863367139302E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.8910543826743605E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.4877619296596099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.11388350694990899},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":5.6680493558219696E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.1091781303237698E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.3410489648522902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":7.3483130098369296E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A12,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C14","v":"012CALS.d"},{"$f":"SQ!F14","v":"Reagent blank5"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A13,FQ!$AE:$AE,0)-1,0)","v":-2.5000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A13,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A13,FQ!$AE:$AE,0)-1,0)","v":-1.5349999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A13,FQ!$AE:$AE,0)-1,0)","v":5.0999999999999997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.12126604149300201},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.27363290426444797},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":894.88440550585904},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.9629733385861599},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":882.36519489784303},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":70.653812937119497},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":10.3239341552291},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":6.7601934282782605E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.0966693711886601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.13240900235471201},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":8.07300529106285E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.76460427741173198},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.5541880057390601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.13420393455027199},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":8.0170527728944299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":4.3607433935672103E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.1176989217475E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.2705523850192701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":4.0633946184896099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.53585867885067E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":10.5073413409102},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.4985656234444402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.2934633991517198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.1796987194403501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":2.5857569521811402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.2018039764904399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.68023038348315E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":6.0403185688031203E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.157235806858709},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.4327396583783802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.6705397958921601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.7526828549426499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":4.2278011314216999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":6.3118482021982502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.1219387437258301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.106535821383552},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.26722352615409E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.1609698745545902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.4696636005902299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":5.7182482448983598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A13,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C15","v":"013CALS.d"},{"$f":"SQ!F15","v":"Standard NMS-59"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A14,FQ!$AE:$AE,0)-1,0)","v":50},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A14,FQ!$AE:$AE,0)-1,0)","v":1.25},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A14,FQ!$AE:$AE,0)-1,0)","v":125},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A14,FQ!$AE:$AE,0)-1,0)","v":50},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":1069.6701833038901},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":39.6507501332621},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":2782.3138794364199},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2674.17545825974},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":2730.9694589585702},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":134.70555697659199},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":504.55870153476701},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.19038959718726101},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.65844881713203},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.6505854965681799},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.6427221760043298},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.1164559361340001},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.3134977674383099},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.78181228874162E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":26.1126889374892},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.287673900601852},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.5955422526212102},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.14549895243338501},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":12.8990780574675},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":51.439045818593101},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":17.4262574176018},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":127.418116461905},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":25.404990086742401},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":4.79841554613394E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.4656337694234699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.249496126979006},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":6.1042465573356004E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.9709115294545496},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.4590846567736202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":1.2427278823636401},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.97008070741496E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.62784136938172E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.69753253640384905},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":9.4408506817000498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":12.427278823636399},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":3.5611721714303303E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":3.21431645611022E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.9225292599640603E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":6.7643428787390295E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.01443580462859E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.155824268951961},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.0284529986409299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.1781652031980401},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":11.879440704499499},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.33221988234556},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.23420020498147},{"$f":"OFFSET(SQ!$A$1,MATCH($A14,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":0.24854557647272699}]
[{"$f":"SQ!C16","v":"014CALS.d"},{"$f":"SQ!F16","v":"Reagent blank6"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A15,FQ!$AE:$AE,0)-1,0)","v":2.1999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A15,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A15,FQ!$AE:$AE,0)-1,0)","v":3.2440000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A15,FQ!$AE:$AE,0)-1,0)","v":2.8000000000000001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.307694064192999},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.68638239283458902},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":2680.6799142601199},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":6.1929366120365898},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":2029.4634630676201},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":121.672341527315},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":30.176408589916299},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.9500223474487406E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.9868301631451593E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.17322802382862401},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.143536337393778},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.98425162482287298},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.3266910605888704E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.114891176098735},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.101431742746857},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.10398349272063601},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.42378699552532E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.4627566682771499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":9.2478833695864796E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.1672759971715298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":12.813815158795199},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":0.198492509759465},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.5610308545776801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.2812771062921599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":6.2975706328940104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.7785453828287999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.3394333365820902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.46410106240838E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.4120843543441297E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.42639063093948998},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.1969728230044E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":3.7295907161409101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":2.2248524766043599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.9971772489967104E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":7.6055195774686807E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":5.1289364821306401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.55609977346781E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":9.7993489437569106E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.8717150232133402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":6.4636588465412995E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A15,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":3.5062098544097502E-4}]
[{"$f":"SQ!C17","v":"015CALS.d"},{"$f":"SQ!F17","v":"Reagent blank7"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A16,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A16,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A16,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A16,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":1.3828302817893601},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.82259918414986699},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":2685.2467275221502},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":8.0404548726074605},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":2115.7466584167901},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":123.252319353678},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":26.282744765781899},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.94148716650603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.2516614457800407E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.148132154682169},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.16599502736348601},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.2052774769226},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":6.0836877696260998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.15479759838575699},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.7117778899080998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":6.4914283858002004E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.6690250831454099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.8568131827256199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.17670724238745E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":17.2668370920799},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":8.7830269983792605E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.8629725578135702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.2902390816559499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.2734268885151901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":3.0722727531929002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":9.2407561654441004E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.43871025617795E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.4619618528674904E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.11324010669194},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.0133396277001401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.8229745764115602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":2.24536329734475E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":4.9176136752992203E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.01386661742125E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.115134670457986},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":5.1762088233943697E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":5.0478442258686501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.8201519415316998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":8.0462608015956904E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":2.7844426614840299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A16,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C18","v":"016SMPL.d"},{"$f":"SQ!F18","v":"Reagent blank8"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A17,FQ!$AE:$AE,0)-1,0)","v":-1.0999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A17,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A17,FQ!$AE:$AE,0)-1,0)","v":-0.628},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A17,FQ!$AE:$AE,0)-1,0)","v":7.1999999999999995E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.24348146372642901},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.31469628138390998},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":132.34087066977301},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4938918886458901},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":104.877528418785},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":3.8733226377968299},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.05986090498706},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.104164980413329},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.1293714766725001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.07465054443169E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1525885268885799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.10079983526627501},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1515230564916E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4227136728298302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.2884953345754199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":5.4446686635820002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":7.0489959532899996E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":7.5192985825211402E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.8662692204684503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.3433044853746698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4356955996298999},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.9492124366048301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":8.9555424329226998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":4.7976054489501098E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.33898624267938E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.42493953597266E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":4.27266157878688E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.9032228075965899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.118002129320479},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":4.1536632159549797E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-9.01425683407502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":7.3200894693594294E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.4355605204687802E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8283772652294501E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.1314721478836401E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.12092088454458E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.08820477458432E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7571984652205298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0200059836029799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4955499546754599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.24195326400373E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A17,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.0143418021298198E-5}]
[{"$f":"SQ!C19","v":"017SMPL.d"},{"$f":"SQ!F19","v":"QC1-1"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A18,FQ!$AE:$AE,0)-1,0)","v":9.6530000000000005},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A18,FQ!$AE:$AE,0)-1,0)","v":0.24},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A18,FQ!$AE:$AE,0)-1,0)","v":22.556999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A18,FQ!$AE:$AE,0)-1,0)","v":9.8699999999999992},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":193.47529205639501},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":8.7313481476630095},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":71.8594432687687},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":534.93331460431898},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":214.243832094027},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":581.65424413090602},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":435.93143672398901},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.04324168613828E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.50031823964323496},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.45311000716476502},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.45988192477422901},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":20.046022447567498},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.23400542363395199},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.95817402841591404},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.07564343002461},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":49.404099160137299},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.50055946428345899},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.23219088706818E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.5185330980854501},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":10.3067051313795},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-0.68862310781465697},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":24.6585001079258},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.4962672104679799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.40645522963228E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.3680360237351499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.12826465237853E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":0.92861811224980495},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.224391282569079},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":0.246646541256343},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6966187559756601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9319581035743199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.479718774309917},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.3797567378469999},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.6073484686183499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":3.9621506252572603E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.14145038739902E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.1905363586779898E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.30712173756754E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.7021647651161301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.22368362264776701},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.3438021852834701},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.25984102657326602},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.218682335792486},{"$f":"OFFSET(SQ!$A$1,MATCH($A18,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.9000755301144901E-2}]
[{"$f":"SQ!C20","v":"018SMPL.d"},{"$f":"SQ!F20","v":"QC2-1"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A19,FQ!$AE:$AE,0)-1,0)","v":-7.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A19,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A19,FQ!$AE:$AE,0)-1,0)","v":0.123},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A19,FQ!$AE:$AE,0)-1,0)","v":0.26400000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.75666021823349106},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":1.55014358437079E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":582.96398039961798},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9992407880433301},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":311.59439273812501},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3124248217012697},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":87.615925594226994},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.9748845496121596},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.8504090504882501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.6588800937474202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":7.6299917482125197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-1.24538375525829E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.74585720719811E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.7938546619919797E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":3.7100398333140203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.10687509312948799},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.16406176650329E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.56798150017591098},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.9580234219967402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":6.08524703366937E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.68791319816859},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2033149240918599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":4.8783012043885696},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.4624686572192398},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.65732553977758001},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.56169252250319501},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":2.6130745817911E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.0570269650828901},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.48006782623745398},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":0.37310194280966003},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.168252446082372},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":8.4130913455945098E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.0821678828490699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.40751748870843202},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.7573553367172501E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.39720220550103602},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":9.1721517388611797},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.54101904549429503},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8145149100753895E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":6.6280077568774097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.39509303925444E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":8.2790426594031893E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A19,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.62887511982266E-4}]
[{"$f":"SQ!C21","v":"019SMPL.d"},{"$f":"SQ!F21","v":"QC3-1"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A20,FQ!$AE:$AE,0)-1,0)","v":-1.6E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A20,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A20,FQ!$AE:$AE,0)-1,0)","v":-0.754},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A20,FQ!$AE:$AE,0)-1,0)","v":3.9E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.270415038546232},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.32630139929915902},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":55.834901716241298},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2831939565882702},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":90.909365010202706},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9919634294538602},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":11.4119976603869},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.1891614090905999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.8205117380328401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.1111120442590099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.6083795229040201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4412958316584001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3773351029348E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6202088510536898E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":3.7603204277101399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.35772335251415699},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.8584174553579599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.7708023099304801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.1459677990209801E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":323.12010784299298},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5077159457943198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.6481479303063701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":3.2445135100290899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.1131680597503001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.61752725038453E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":3.6659658821285197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8563436428617899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.0788660016112901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.0344704065749298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.48095951229209E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":38.153483598572002},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.09827118125033E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":5.8245804257145897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":5.1244176683480495E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.5707878215819403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-6.6879015712861603E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.0726452566949501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.62105767522797095},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.13902041081187999},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.2518687576999001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-8.3179478410282896E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0775566738809601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3790987713739099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A20,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.0146131139437903E-5}]
[{"$f":"SQ!C22","v":"020SMPL.d"},{"$f":"SQ!F22","v":"acid rinse"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A21,FQ!$AE:$AE,0)-1,0)","v":6.99},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A21,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A21,FQ!$AE:$AE,0)-1,0)","v":422.48599999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A21,FQ!$AE:$AE,0)-1,0)","v":1488.47},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":11.0936330682948},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":71.540760007783803},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":59296.648170891},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":611.38088345815402},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":78733.251613037806},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":3636.0121946528202},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":628.210705130919},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.4199032356024901},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.40770869063308},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":15.7970473489094},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":16.112348281662001},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":178.76573792929301},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.59925845211397699},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":11.5735567530807},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":10.7234782460794},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":2839.2860979117499},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.0824037067515899},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":13.5869659239501},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.18999644449169},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":5.4981460473739796},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":18735.284703505899},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":29.248496904117498},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":36.6075698226336},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.32429119710526},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.3878455623542301},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":3.5116135077276498},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.046"},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.96098617035791101},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":13.639614478065401},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.212405365421205},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.036"},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":104.620730721415},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.1264898274214401},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<2.030"},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":1305.9919950597},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.35977567906087699},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":7.9888555680972004},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":5.1275011435927196},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.0690037049501999},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":1.0656612935109799},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.006"},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.017"},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.37143848726031},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.3325515712291899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.13585499800329001},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":836.45013263067699},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":96.023031749698802},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.124847049463829},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.4097680688166299},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.4971077075436598},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.15509907939005},{"$f":"OFFSET(SQ!$A$1,MATCH($A21,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.7871261362302E-2}]
[{"$f":"SQ!C23","v":"021SMPL.d"},{"$f":"SQ!F23","v":"base rinse"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A22,FQ!$AE:$AE,0)-1,0)","v":0.06},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A22,FQ!$AE:$AE,0)-1,0)","v":7.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A22,FQ!$AE:$AE,0)-1,0)","v":4.8529999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A22,FQ!$AE:$AE,0)-1,0)","v":12.146000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.067"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":29.197767078346899},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":34955.5373048213},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":208.64778498068901},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":28250.4165367852},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":1771.0389914147399},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":265.30601443026097},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.2530523988704898},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.395631007041648},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.54884848487491},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.8715215005998},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":23.395808556816998},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":7.8761723310717402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.77807042734186904},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.6884696014107601},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":56.0645027588345},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.30775249888131401},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.35235965006294401},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.4446168795665399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":1.8235792339638801},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":278.12525733474001},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.0316642867376999},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":0.117248572441203},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.79448229786027402},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":9.1876739431023299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.87285255049099597},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.79148851608848E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":9.6493933101501297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.36669296464992},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.100691523684122},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.062"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":14.8131320132146},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":5.3271510532154299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":0.12637185377709001},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":6.8903083515204798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":5.4725391207253501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.3680320490913404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.6370517809485606E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.7405850055667297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":5.4666264729550697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":83.523980578749303},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":7.7656780888471097E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":4.1155474570471404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.0528060420324499},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":5.3743020254950501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.30922996893994},{"$f":"OFFSET(SQ!$A$1,MATCH($A22,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.3184581648414997E-3}]
[{"$f":"SQ!C24","v":"022SMPL.d"},{"$f":"SQ!F24","v":"nitric"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A23,FQ!$AE:$AE,0)-1,0)","v":4.343},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A23,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A23,FQ!$AE:$AE,0)-1,0)","v":1052.598},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A23,FQ!$AE:$AE,0)-1,0)","v":16075.18},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.359"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":784.73199675303294},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":203326.372798581},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":666.97231164975699},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":161432.284773269},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":10373.5765550786},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":1612.8242851837299},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":15.6829822765829},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.5015415230230298},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":61.633061939734802},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":64.366271462881301},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":827.45012976282203},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.7475282239642},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":45.855457625475402},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":54.766540934598297},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":140.27421209421601},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":14.991798123456499},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":9.5624676440297698},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.8555259600073999},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":98.629465187157905},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":7498.8441899683003},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":43.922275059922498},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":6.2659257629349101},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.244122196508332},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":8.3773158238667396},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":2.9252080784561398},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.210"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":8.7475824515193796},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.074"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.242"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.165"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":23.7106777292727},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":14.4408112360901},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<9.235"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":1335.73507663436},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.6377171821038301},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":7.6524507030648303},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.044"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.243"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.104"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.027"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.077"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.76088981745585305},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.023"},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.23712115887103},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":315.502055945341},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":10.527048305285399},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.341710298293947},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.5587977819589796},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.7948292336820899},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.757095084822345},{"$f":"OFFSET(SQ!$A$1,MATCH($A23,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.024"}]
[{"$f":"SQ!C25","v":"023SMPL.d"},{"$f":"SQ!F25","v":"Reagent blank9"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A24,FQ!$AE:$AE,0)-1,0)","v":-2.1999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A24,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A24,FQ!$AE:$AE,0)-1,0)","v":-1.3180000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A24,FQ!$AE:$AE,0)-1,0)","v":0.158},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2307462606517601},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6220326556934698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-68.992866237710103},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-0.69399984729588604},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":506.440951394781},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-0.28660989168790302},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":7.9560095942767403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.75166388799218E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.2611136763737101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.71314684300808E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.3128716909602701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.72434564216954},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5952071243863499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.4586506130675801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.5728921855506699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.50175602909621098},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3440139736101599E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-8.9493081450239403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.84830101125032E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-3.55644860698286E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1136270764227301},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.58807760665333E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.8954106089723301E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.0206012065947399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.8724770371517801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.9044572289762999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-6.00841933197197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.9639852830275599E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":8.6345116694569598E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.2701738826260196E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.68394796714292305},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.2159751348364499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.8549055151196301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":3.0234454798714201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1134348752471899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.02533780209501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":1.0229588955966},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.6056740178026002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4838020208632197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7278489935235899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-7.0862313579729804E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6181032653687101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A24,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C26","v":"024SMPL.d"},{"$f":"SQ!F26","v":"Blank control1"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A25,FQ!$AE:$AE,0)-1,0)","v":-2.3E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A25,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A25,FQ!$AE:$AE,0)-1,0)","v":-1.2809999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A25,FQ!$AE:$AE,0)-1,0)","v":0.35299999999999998},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.15036268865307301},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.60376099811740302},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":2518.7331151937201},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":4.8422006866449996},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":1966.12378089412},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":120.936205129892},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":31.208131511031802},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.140345387072493},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.8279876019621202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.18544622907031399},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.18172762913964},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.5511288539275701},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.6055492089780601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.9825361812703906E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.107151520405879},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.40350090951414402},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":3.9459117120023597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.2107971129237201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.6820531075426799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":4.4557379315151302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":18.552606316371101},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.9532263687237503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":7.4251240425254397E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":8.9832144026047601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.5300607729219701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":8.4817288355203101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.26604573543061E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.5689378859270301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.73988722856479E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.76956566717627695},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":6.4134136960588401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":9.9916376125462805E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":7.6125734411597895E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.6127892775240001E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":3.9736076352837798E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.6232806665063701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.64679774706176996},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.19848512955208E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.19108777671899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.5593519418243899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.17209290279733E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.0537742244832299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A25,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.3494943377644405E-5}]
[{"$f":"SQ!C27","v":"025SMPL.d"},{"$f":"SQ!F27","v":"19286427-001"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A26,FQ!$AE:$AE,0)-1,0)","v":3.33},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A26,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A26,FQ!$AE:$AE,0)-1,0)","v":-0.65600000000000003},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A26,FQ!$AE:$AE,0)-1,0)","v":21.213000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":63.4912044107095},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":13.8518457450031},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":45.0398522725541},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":4.6865062831858504},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":10018.263011302601},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":319.40276756943803},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":195.89632945595},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.154232974978737},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.65653229020398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":4.8473380143762998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":7.8188186051413894E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.0380030934308202},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.87939985738687E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":3.3661305492135599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.36544691964916598},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":361.19103426047297},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.62262261292667698},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.4991230801282399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.8864234753814099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":832.71821376140394},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.9907812796367301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.63736512611301},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.1111654022263299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.11544028701620999},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7335742135113501},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.92135421148129E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.21146292861823E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.4686955960942905E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.40072946775499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.5788569559229306E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.184888111688673},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.0911638958172E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":98.510793639238301},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.4273514814212191},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.7161811678264696E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.43622575148809E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":7.0152662065572999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.5699656656711402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.7466056210821408E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.70076972161847E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":8.7244616198562796E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":2.4412296507629798},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":8.9114090888553696E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.2337520332230401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.6201498112512902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":7.6172898091019401E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.64200149031535E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A26,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.07605638477649E-3}]
[{"$f":"SQ!C28","v":"026SMPL.d"},{"$f":"SQ!F28","v":"19286427-002"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A27,FQ!$AE:$AE,0)-1,0)","v":3.1720000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A27,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A27,FQ!$AE:$AE,0)-1,0)","v":-0.77900000000000003},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A27,FQ!$AE:$AE,0)-1,0)","v":22.547999999999998},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":61.352194781947802},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":17.574010773278399},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":51.872284134678502},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":7.5023627215957296},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":10961.6361811192},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":282.77490472845398},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":212.27721886086201},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":6.1196400137132997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":4.95877514295553E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.185779596184666},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.14874406432760201},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.40878802228189},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.8151605621368201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.821558528975186},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.52711690611765605},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":309.15431109251602},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.62998222107972701},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5135060528698001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":3.5429017698576102E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.05641178321035E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":891.66042347063399},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":8.7307212905958595E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6018115877475001},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.22205121528662E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.12600334690963499},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.75218970037525},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3391155980295401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8832052049348892E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":2.0227364529245401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5899942458092401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.6922809414604493E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.37021292173272302},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.0292830465577199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":95.949891551133305},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.0528620204388393},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.9611619452437699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.1195839193823101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":3.14346610317986E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.69558900981027E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.6948085929422605E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6614944370728198E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.4722473922604701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":1.5439548041378801},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":8.9074055350800305E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.2660612487443799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.4353034579836798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.4859787184530799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8190821936929998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A27,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.4626945185020804E-3}]
[{"$f":"SQ!C29","v":"027SMPL.d"},{"$f":"SQ!F29","v":"19286427-003"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A28,FQ!$AE:$AE,0)-1,0)","v":3.1930000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A28,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A28,FQ!$AE:$AE,0)-1,0)","v":-0.72599999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A28,FQ!$AE:$AE,0)-1,0)","v":20.721},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":63.9827978598424},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":14.894807372310799},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":126.05708626360899},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":7.1607279564380999},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11542.941906977199},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":290.69830034052097},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":211.254457091951},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.119547924283051},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.6506435722617897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.11269709702557799},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.21849468517179399},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":20.511420342716701},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.2025354263837801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.7193216554276403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.37287919700609201},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":314.76402872650903},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.58996678583025997},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":5.4843540621438303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.45103830204954E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.27527417351602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":749.70062867053502},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.5203116884592299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.53182158684368},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":9.3124179053257795E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.9419568241586306E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.8455785497480699},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9537891006211504E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.63955112920327E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.36216563378081E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.50778725906179E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.1966413715769904E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.350260984176753},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":6.0112616434124797E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":94.919960079291599},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.8319038850515206},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.9195843185712903E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.2769759819512399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.7941718854695999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.9400535734525699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.3015429130394395E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6536905717973801E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.08558612491557E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.40212067730648599},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":7.5531487883070098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.7268875591105899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.8294304641010297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":9.7505511008169404E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":3.0720756399051801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A28,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.8610768704855697E-3}]
[{"$f":"SQ!C30","v":"028SMPL.d"},{"$f":"SQ!F30","v":"19286427-004"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A29,FQ!$AE:$AE,0)-1,0)","v":3.2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A29,FQ!$AE:$AE,0)-1,0)","v":3.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A29,FQ!$AE:$AE,0)-1,0)","v":-0.77700000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A29,FQ!$AE:$AE,0)-1,0)","v":21.195},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":64.144061974241296},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":14.980462611176099},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":213.408753966067},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":5.7814106197622097},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11256.149313713},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":342.32843028786601},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":191.516783773463},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.32143605796203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.7141705622680599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":5.9396726726476001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.1776965498117E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":4.0929856024818303},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.6335232445378401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.5869318894424098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.32136625203824598},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":254.494300480916},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.58371725834287802},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.0573172453844201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":9.6955966227617804E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.31792555391531E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":715.04456372796506},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":8.3783727154277607E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.57785065140402},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":9.7890871015834692E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.107506402413062},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.8680357940846599},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9668461499897502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.25363232612874E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.0505342492072901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9088209875076801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.6711340623068395E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.36133835160071798},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":8.5293754379636797E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":97.488844712174398},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.9759286034981898},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":3.3809359039061499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.4014098888464602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":3.6841763659091101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.73326699081293E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.2842130570216404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6125106494367699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":6.4498463354536995E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.420323017957279},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":8.0674411168897606E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.3190630998677401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.2641864444097302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":4.3346956522400499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5228408767120002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A29,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.1268826069966397E-3}]
[{"$f":"SQ!C31","v":"029SMPL.d"},{"$f":"SQ!F31","v":"19286427-005"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A30,FQ!$AE:$AE,0)-1,0)","v":3.258},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A30,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A30,FQ!$AE:$AE,0)-1,0)","v":-0.754},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A30,FQ!$AE:$AE,0)-1,0)","v":19.699000000000002},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":63.957129155750799},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":12.4735318549826},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-7.5361235531122501},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":6.7252729930738999},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11851.001813765601},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":285.02617760894799},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":169.68423873056901},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.10476775793160401},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.3831941854075797E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":7.5796661041408003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.101622225492196},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":8.6287165415228806},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":7.0360607956609803E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6262620476379501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.28470451374449501},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":218.117592522671},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.51462182134938295},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.2793311187244E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":3.8208456554634E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4994038273350599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":672.05543941269696},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.44315261297523E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5159689465991899},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":7.9967707132374694E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.107037425950034},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.77888224595105},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.30159097233432E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.39475788091429E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-6.0295728820621702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.8826819432781997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.27723039400175198},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.23428227756802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":95.965010989977401},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.9702426846582703},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.2377648646892999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.6264629660815199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":2.6864459386014E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.5619724608350999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.8923330673012805E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7076918219200899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.3643032768986001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.20663556583241299},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":7.12063582455357E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.1450936139706701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":3.9433202304981298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":7.7093756402037399E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5479158834783701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A30,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.59675386036177E-3}]
[{"$f":"SQ!C32","v":"030SMPL.d"},{"$f":"SQ!F32","v":"19286427-006"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A31,FQ!$AE:$AE,0)-1,0)","v":3.2410000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A31,FQ!$AE:$AE,0)-1,0)","v":3.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A31,FQ!$AE:$AE,0)-1,0)","v":-0.79900000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A31,FQ!$AE:$AE,0)-1,0)","v":22.763000000000002},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":54.875119249938301},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":14.7924429382463},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":77.371123563995795},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":11.185961858426399},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":12044.896450369601},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":362.71816338720799},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":185.749422782268},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.9126608283260906E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":8.0092948920798407E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":5.65384947033368E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.88658971331354E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":2.1553622098460501},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.4520371172484599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.6731062227653104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.34319494723718602},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":417.39551057858199},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.64077697895722596},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.35113016536361E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.0434508053648499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":970.86536304905405},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.81704330515655E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.7801860889902299},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.0573837845707999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.111629564911262},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.8017301900087099},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9012126817211503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7806549619083499E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.6793361889954603E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4005810604407101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.8044289025989301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.246425035311678},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.3180481563130199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":97.237542130627105},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.9614494336536907},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":7.5328601686599104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.3422123313430299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":8.2048317787765997E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.98040354358719E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.2251323739871403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2544925730550099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.1289586725666199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-3.52853721518194E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":5.9756788409716301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.9419429618965499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.5051042123841603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.07145010943114E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.03963220988186E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A31,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":7.0580086175119697E-3}]
[{"$f":"SQ!C33","v":"031SMPL.d"},{"$f":"SQ!F33","v":"19286427-007"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A32,FQ!$AE:$AE,0)-1,0)","v":3.165},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A32,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A32,FQ!$AE:$AE,0)-1,0)","v":-0.78400000000000003},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A32,FQ!$AE:$AE,0)-1,0)","v":21.131},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":64.068500692308902},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":15.1021333090914},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":67.114353244794103},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":11.233547058553},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11350.426802199299},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":335.74419777819202},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":217.43296133464699},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.7991227915685602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.2794896760962003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":4.3144485670675699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":8.0459430694323195E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.4601686468015198},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.0540022254501998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-3.18862478866458E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.289058525586674},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":380.54014810031703},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.545177075176982},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.8891253281115E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.6954960953962501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2346155436415601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":783.78577243793495},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":0.13206852657176399},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.7183582434834399},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.20547561391271E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.119835052259715},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.74796089906184},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8525687509638803E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":8.6272614060890296E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.29112894788909E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1312263804383201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.2396612973842695E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.45461256635144398},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":8.6624672639661103E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":100.343381635691},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.2211883831718904},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":5.9693272981205404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.91465486414158E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":6.0446412371036105E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.0177397786928699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.10517983342937E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6750004576004095E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":6.5072828518225297E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.238008778753936},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":6.86768673902933E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.6407675021586701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.5916267592408397E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":5.8514638283932202E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":4.2067836209728397E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A32,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":7.1505663573191202E-3}]
[{"$f":"SQ!C34","v":"032SMPL.d"},{"$f":"SQ!F34","v":"19286427-008"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A33,FQ!$AE:$AE,0)-1,0)","v":3.2170000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A33,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A33,FQ!$AE:$AE,0)-1,0)","v":-0.51800000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A33,FQ!$AE:$AE,0)-1,0)","v":20.803000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":58.450414751316998},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":17.796398760975801},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":159.34964988054901},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":14.022929382925099},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":13255.411457689699},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":347.557717019775},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":198.25841764348201},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.7374447419496404E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.0229962977707501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":8.0314550165689905E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.10910580927215401},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":5.0217383576821399},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.8007360974317101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0081240826155099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.35537300399028299},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":290.51746641676198},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.61789212781553904},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9804828803798598E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":9.6580410133061904E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2049665034831599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":688.00083429181598},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.4792445764409493E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.54159905220056},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":9.3941801273503701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.11202683708515999},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7517996887885301},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0373084414770401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.161482170186E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":9.9614164262248005E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9017828282609403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.56465053201144E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.34199925347501398},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":3.15953447056586E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":95.514555141757796},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.9689410671023104},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.6721817700927899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.3933003164835501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":8.1955948360988801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.7696098796044298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.7085060905308792E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2591574646687501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.1900383505443998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.118207133402691},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":6.2405243328435098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.4103435176756199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.7002199043130098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6629681897535602E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1769731961250402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A33,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":3.62774315845778E-3}]
[{"$f":"SQ!C35","v":"033SMPL.d"},{"$f":"SQ!F35","v":"19286427-009"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A34,FQ!$AE:$AE,0)-1,0)","v":3.1859999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A34,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A34,FQ!$AE:$AE,0)-1,0)","v":-0.66500000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A34,FQ!$AE:$AE,0)-1,0)","v":20.736000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":62.204869761249803},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":17.111770426736999},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":73.287912316501703},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.7924861579253699},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":10414.121571826099},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":318.87257586417098},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":188.54451019598201},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.2497684118345798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.7897070403559001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.15472884628380801},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":8.8069963366828402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":5.1435651787493804},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.9361517453107798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.9786422421463902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.47095658725037998},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":294.05408904202301},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.49989215567300099},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":3.9576982819256104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.51048445595773E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2327092173658299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":769.42392889263704},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":9.0546105713352795E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5788814372809601},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":7.7341494407914E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.113893904508163},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.79439277057172},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8748937410074299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.25651089814039E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.0005366507700301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5541751981371201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.0661706805041195E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.38522424675620898},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.9211459428094501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":97.249711718179498},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.02311248878115},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.6503640638559697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.14935746447973E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":3.2248195553474199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.26734724940892E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.7637828260273007E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.99976186170535E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":8.6027093013494704E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.16359427796051901},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.1722154506029403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.31988938931252E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.4294641989416602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":8.6762082017945196E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.35092838049771E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A34,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.9623449608090101E-3}]
[{"$f":"SQ!C36","v":"034SMPL.d"},{"$f":"SQ!F36","v":"19286427-010"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A35,FQ!$AE:$AE,0)-1,0)","v":3.0880000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A35,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A35,FQ!$AE:$AE,0)-1,0)","v":-0.80100000000000005},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A35,FQ!$AE:$AE,0)-1,0)","v":21.748000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":66.857371943972893},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":16.088222268114698},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":109.419063821162},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":12.7684664175284},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11319.041249136901},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":317.27421279048201},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":196.235862420177},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.3658333237025503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.6150378606981503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":8.8816971919023199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":6.88743103942666E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.1788354439817099},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.6413969002538899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.64475759801178E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.270304115696543},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":281.62266798591003},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.52403589416738905},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.7674693996054701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.49184532798671E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":8.7460197728880307E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":722.70763405308696},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.3417781863249999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5226497449109599},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":7.1215847143281004E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.4545216722237702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7816201821519599},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-4.9819493090849003E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.3287860971198296E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-4.41436427473757E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.7066295775205499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.25356818709506901},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.44365454111381E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":96.074624975813407},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.8939417893521302},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.2383371025254299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.8040453213440501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":6.2869952668266095E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.8231469423609201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.5861460456890196E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.15494774077605E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.27869960927957399},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.26446067358694E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.7976419479327199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.4790649250225002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":4.18322911082519E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1151554015604701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A35,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.5989241887317601E-3}]
[{"$f":"SQ!C37","v":"035SMPL.d"},{"$f":"SQ!F37","v":"19286427-011"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A36,FQ!$AE:$AE,0)-1,0)","v":3.278},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A36,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A36,FQ!$AE:$AE,0)-1,0)","v":-0.83699999999999997},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A36,FQ!$AE:$AE,0)-1,0)","v":23.213999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":70.725195464700406},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":18.401192428189798},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":38.6047802452917},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":10.8006506799583},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":10902.1372842981},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":348.16706834220997},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":210.49701375356901},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.9792930872762703E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.2524253037442995E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.102286607388701},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.5607166177020202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.6089003542958502},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.1724604001193701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.4811106796350801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.35923699589276398},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":424.157863222245},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.60474546244534},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1532935156376E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.31232322964355E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-3.21396827130775E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":969.13071007413805},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.6017083348509001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5770580440351001},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.1122954468890301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.116256765730642},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.76368772742391},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.89328920596112E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.1963777019966299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":8.7655497605373001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9947152878701901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.2598277004106194E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.48334064971665502},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.0549308348388399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":97.370994282005498},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.1734417463870894},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.6175445601394898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.6363610357147502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":2.7413010795100001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.4567543572628499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.7461917511704092E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6394628181942001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.2032473435831101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.27949523536918902},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":5.2315048137355699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.9482707213708901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.4135033061711302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7312829970648301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A36,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.3221734086962491E-3}]
[{"$f":"SQ!C38","v":"036SMPL.d"},{"$f":"SQ!F38","v":"19286427-012"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A37,FQ!$AE:$AE,0)-1,0)","v":3.12},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A37,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A37,FQ!$AE:$AE,0)-1,0)","v":-0.79300000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A37,FQ!$AE:$AE,0)-1,0)","v":20.529},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":54.299929407721997},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":16.314427699094701},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":185.554196797641},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":5.00569377526491},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11711.4644667096},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":312.38376308351002},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":192.34309940337599},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.0202705128986603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.8413741102069E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":6.4031945181542399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":7.4148029673349006E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":5.3034807113766202},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.0769996253284301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.5516260313095493E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.27899731708871101},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":215.75779274737201},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.638715062466928},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1059736828014002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.12501990961438E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":738.12940090994596},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":9.6960849786132203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5431304910118699},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":5.3425205607059998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":8.8772574720244801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7739615844199399},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-7.5310543878721803E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.0220611135031499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.1387176017906701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6625894011899398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.0232684734673099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.37243173799194901},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.26007266275869E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":92.344249850109406},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.6115700637676902},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":2.8779247119166602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":1.9848449694216602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.0196093597831502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":7.78229906430355E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.0174881843363098E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5018609345317598E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":8.3660902245721403E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.32664853263091898},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.7988932205609201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.8478411774795801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.9819220673739999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6543755765266601E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":8.3603154494653906E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A37,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.9452290533102596E-3}]
[{"$f":"SQ!C39","v":"037SMPL.d"},{"$f":"SQ!F39","v":"19286427-013"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A38,FQ!$AE:$AE,0)-1,0)","v":3.169},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A38,FQ!$AE:$AE,0)-1,0)","v":3.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A38,FQ!$AE:$AE,0)-1,0)","v":-0.88100000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A38,FQ!$AE:$AE,0)-1,0)","v":21.709},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":68.303715744688304},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":15.456940141789},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":127.427414042857},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":15.1448666505791},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":12847.491941205601},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":347.038681279971},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":215.11492039474001},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.8459285203896907E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.9246874663452499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":9.8862512620969498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":6.0558565879990603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":4.2667486148368301},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.1918969342189002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.9986796445696798E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.365964756763349},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":351.72421260578801},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.59248316360884101},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.2415855588389299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.7768338510524198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2134950484808399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":839.74705356869799},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":0.13224469050398499},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5920184733034299},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.0519701664969299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.123802281000383},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7640420919038999},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6895220559379198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":8.3582600656752703E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.28040510998874E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.45227526713165E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.0355184606448499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.29512846956957201},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.4674203475463501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":96.490957886081304},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.9679347088788592},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":3.33987480749985E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.4257598417161401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.08884836885742E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.6341926542619099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.0656136587113001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7259989117928299E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0619249190902201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.323449168421726},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":9.6475276569859894E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.99573022116217E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.0489486373854099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.3988823832937499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-5.34045994941432E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A38,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":6.2171854376920303E-3}]
[{"$f":"SQ!C40","v":"038SMPL.d"},{"$f":"SQ!F40","v":"19286427-014"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A39,FQ!$AE:$AE,0)-1,0)","v":3.234},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A39,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A39,FQ!$AE:$AE,0)-1,0)","v":-0.79200000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A39,FQ!$AE:$AE,0)-1,0)","v":21.076000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":56.487214783800198},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":13.9240879972408},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-88.406565954217996},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":5.1951799936415801},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11850.508083508301},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":312.82109968654902},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":187.92511499188001},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.5863872171037497E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":4.0279608787009098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.10111801268904599},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.105326295410994},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":5.7789780936964998},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.676334462902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.6198617224039299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.28083615315953703},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":274.81878578115698},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.55981444033149297},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.4408466203047999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.09633561954833E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2800671514793602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":715.23075313335596},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.5299253168971801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5885195077746399},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":9.9843753947245796E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.103803981879735},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7794838433956699},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9308665494295096E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.1454752353233599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6003790120762002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.102850066585568},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.39346996429560799},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":4.03438595407296E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":94.660604108147396},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.18817682183907},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":3.3094597956356002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.3506461416042901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":6.3017661526502901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":3.3932606544710702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.8465021151036202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1950904979081598E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":7.3657513822569895E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.28136256364519002},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":7.8922416050916905E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.5689934487376799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":3.8431232107017099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":7.0279056841255499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3877789958971896E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A39,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.8948255120185701E-3}]
[{"$f":"SQ!C41","v":"039SMPL.d"},{"$f":"SQ!F41","v":"19286427-015"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A40,FQ!$AE:$AE,0)-1,0)","v":3.2639999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A40,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A40,FQ!$AE:$AE,0)-1,0)","v":-0.86199999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A40,FQ!$AE:$AE,0)-1,0)","v":22.452000000000002},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":63.5778400523578},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":15.463809759882199},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":15.646297550764301},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":6.7308575351604301},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":12252.514132517899},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":327.93060391586698},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":193.82065469075599},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.111537309019748},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":4.57099904898636E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.10495618810681701},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":6.0805176776318498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":4.0197539876833099},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.8529172107351498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.3384645435586901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.29330559697439701},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":385.43817485738998},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.54520784377707798},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":4.3160124126236802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":9.6486328527090309E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.20064829124345E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":995.33055514510897},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":0.159642847147559},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6460330835327499},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.0700945551892899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.9955382088793704E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7585091948324501},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4914251842484701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.4098208371763599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.2247742713692697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5141406535439601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.26192477809019E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.35384382076486498},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.7498730484774301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":97.538174278575298},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.3163951058076897},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":6.8858801191549503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.61093193850696E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":7.9816714295288501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.9870155471765302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.08355993388261E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1366123160499503E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.3136225033944899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.36032480678668299},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.5865823774178298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.6976327241228499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.2140886600366497E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.10801239200574E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3942135465668002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A40,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.0531338360874098E-3}]
[{"$f":"SQ!C42","v":"040SMPL.d"},{"$f":"SQ!F42","v":"19286427-016"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A41,FQ!$AE:$AE,0)-1,0)","v":3.17},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A41,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A41,FQ!$AE:$AE,0)-1,0)","v":-0.60199999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A41,FQ!$AE:$AE,0)-1,0)","v":22.962},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":64.338055460304801},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":16.4040874430389},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":76.868068087858802},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":7.6335130268620004},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11841.883263007199},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":385.41584639639501},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":224.33958250982201},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":6.6777352045271895E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.8221634618199398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.106653330369191},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":6.9643892330284701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.9969849262681598},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.49445672905838E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.9220869625525202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.29727130601920398},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":507.14327913143399},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.61855460539460305},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.2216987226493399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.14923561221535E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":4.5063187438641801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1040.1174486495199},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":9.9096502541444295E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6889125896286601},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.8070425715449999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.11613449529183301},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7146579957332699},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5866152125861601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.14731470018987E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.2099927773813299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.47545603927825E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.8897535798586493E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.53243144205895299},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.4667040783514101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":94.690574526026793},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.7824801896587807},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":7.0948306285494797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.44276517109254E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.31844969544787E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":4.0816099717369296E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.07307653647256E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7434813304462399E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.18467243124394E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.42315995222007902},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.4953370632290698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.3473642426840502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.9271551246526902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":8.5372148481862502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.76494339540468E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A41,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.8909719953279807E-3}]
[{"$f":"SQ!C43","v":"041SMPL.d"},{"$f":"SQ!F43","v":"19286427-017"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A42,FQ!$AE:$AE,0)-1,0)","v":3.1629999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A42,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A42,FQ!$AE:$AE,0)-1,0)","v":-0.82},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A42,FQ!$AE:$AE,0)-1,0)","v":21.594000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":61.645970349747898},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":14.3407395227092},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-138.958506138053},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":7.9992419686737097},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11424.4994922927},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":349.28878290366401},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":191.04958457604101},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.124747112431222},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.0080777269534402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":6.7016751322083595E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.2535187326002403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.9052331585470599},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.4739341326274601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.11850390017698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.39884147029231998},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":367.60516944843903},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.51141523492097396},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":5.3443387071382504E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":9.11614179423249E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":8.6123784053575195E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1000.43041909909},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":0.104269684675025},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.63832982038819},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.5708481324510901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.107188033235614},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.78157578636596},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8700973788234703E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.1642071568717502E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.3357529619286699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.9595116449265602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.10054947560118099},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.484423917930277},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.1046339725258999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":96.037420250377707},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.0296124763380803},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.1818298522242102E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.1260875259550699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.4637883027044596E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.9789412442350802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.5544363586212892E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6398728190715699E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.25477862286116E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.34392265348295897},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.4729826422427897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.2346392628174402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":7.87616563113364E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":7.1279779939564E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8611252329355199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A42,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":6.8464854892321602E-3}]
[{"$f":"SQ!C44","v":"042SMPL.d"},{"$f":"SQ!F44","v":"19286427-018"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A43,FQ!$AE:$AE,0)-1,0)","v":3.2610000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A43,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A43,FQ!$AE:$AE,0)-1,0)","v":-0.89},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A43,FQ!$AE:$AE,0)-1,0)","v":22.584},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":58.307881652263497},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":17.076267817689299},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-154.321951244579},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":8.4507575660661605},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11392.3883940342},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":340.17320235418401},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":213.42422983313699},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.18630623365305601},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.4954133874775499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":6.8509588528921306E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.0153140559928302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.71752612484448},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.5918779315000401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8835732299261801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.57883737183634099},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":449.088718298103},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.51786985303475397},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":7.0220104826644999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.4469937136235801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1065.3321867919101},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.46081451668502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6399898443665},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.1869903029104499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.13190147990241399},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.76958534445178},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1624313672053993E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.0850317004660798E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.64214614872212E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5864554019876998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.9937919854554405E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.55772527777499803},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.24185547174332E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":96.454347173466999},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.8912798207507802},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.9710398522365704E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.4246810269796401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.4081087637279396E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":3.3979394895640399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.5593446140983702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1246256573700501E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":8.4296045693354105E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.424816300568153},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.5680154041588102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.1542327683060202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":6.0346635598647398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.7885775596816603E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.23610752180716E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A43,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.3341770712614897E-3}]
[{"$f":"SQ!C45","v":"043SMPL.d"},{"$f":"SQ!F45","v":"19286427-019"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A44,FQ!$AE:$AE,0)-1,0)","v":3.2309999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A44,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A44,FQ!$AE:$AE,0)-1,0)","v":-0.78800000000000003},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A44,FQ!$AE:$AE,0)-1,0)","v":22.384},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":60.6357205059766},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":14.413630744665999},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-65.500410782789501},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":9.2349848941454695},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":10427.4251213999},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":391.96183283406998},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":231.35403758936499},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.137519159324024},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.5705271361547098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":7.99584578519595E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.14362458698201899},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":8.4607420944681699},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.7245714452641899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-5.4538513284960602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.33303500867557001},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":534.12191887267898},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.61433417851324601},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":5.1012363930814601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.30512975124235E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3021616647548201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1040.90657753255},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6777943696420299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.72560509798412},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.19395344123343E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.14151945464963001},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.68787937607236},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3903096670689203E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.68959704335846E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-4.9345154445322196E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.72258687215232E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.44474879344177398},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":3.8989287137818302E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":93.922775148059301},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.8970496043797809},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":5.7164781182879997E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.73885616142522E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.8790094103824E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":2.3187219080574402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.4933991995828707E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9815581639204701E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":5.19303638827556E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.40918802838791202},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.28732857131182E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.6620216064773699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.5315579178030302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.01901082697429E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0657624650521799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A44,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.6566481241237593E-3}]
[{"$f":"SQ!C46","v":"044SMPL.d"},{"$f":"SQ!F46","v":"19286427-020"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A45,FQ!$AE:$AE,0)-1,0)","v":3.2189999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A45,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A45,FQ!$AE:$AE,0)-1,0)","v":-0.752},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A45,FQ!$AE:$AE,0)-1,0)","v":22.765999999999998},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":55.7427625042709},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":15.8356088020572},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-83.585249433385798},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.2365670569463501},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":11297.521582478899},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":361.90665717157998},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":221.26724725292701},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.6274076438011799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.5313150332931701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":9.1050640022922305E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.116993639367174},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":3.4309026291137998},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.31005451538503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-1.03397518036099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.34876619829265099},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":524.35399851627596},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.63201110299658103},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":9.6881314738865197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.07791931214295E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-1.15190524097274E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":922.44362454497605},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":9.0236201921961801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6232006097860601},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.8203588212530701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.14109660346219599},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.7567825869491001},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2747974372973402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":8.4297242744752203E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.8128141803936399E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6308679069931996E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.4271923935281504E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.69799903048675604},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.2672875747036299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":96.166682185520798},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.1679436083911607},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.3435205732925601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":2.1166136249820701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.23512278721636E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":3.4438286487464098E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":8.5112433133011697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2403395095031199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":7.4755243093332501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.39193519679530697},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.9783989675278297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.9486232013948999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.4339657424279103E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.13733453336687E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":4.9642342770150103E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A45,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.8620057400208996E-3}]
[{"$f":"SQ!C47","v":"045SMPL.d"},{"$f":"SQ!F47","v":"Acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A46,FQ!$AE:$AE,0)-1,0)","v":5.7610000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A46,FQ!$AE:$AE,0)-1,0)","v":0.16400000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A46,FQ!$AE:$AE,0)-1,0)","v":121.916},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A46,FQ!$AE:$AE,0)-1,0)","v":922.875},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":3.8687109054585198},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":31.672225913667202},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":15705.737498603799},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":151.549038421809},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":15396.389928730699},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":1105.18903560786},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":158.234422953744},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.1217516894230402},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.92040993978909003},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":9.7236532988859601},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":10.081650074075901},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":103.14850235662},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.45313026433106901},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.6325677112876997},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.2821716886485897},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1680.5548599188801},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.9855699836870899},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":8.8931647229512993},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.078"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":3.23758989975516},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":17679.0260020112},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.2869017179681999},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":23.971175638781901},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.130358248149998},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":8.4266794268407094E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":9.8205281977543901},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.028"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.3289527794734799},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.31132364617351},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.032"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.022"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":57.5380453494191},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.5988977053801201},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<1.235"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":387.45883748674498},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":5.4137065842728899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":107.777797049825},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":14.8647977603397},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.08448273184408},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":0.13899736265078599},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":5.9052244498857598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.010"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.10142106418312},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.003"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":159.637816340744},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.6877996017904699},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":7.6181022196138107E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":0.92150088246152795},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.12964932326592399},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.103952121851438},{"$f":"OFFSET(SQ!$A$1,MATCH($A46,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":3.2008177470295301E-2}]
[{"$f":"SQ!C48","v":"046SMPL.d"},{"$f":"SQ!F48","v":"Base rinse"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A47,FQ!$AE:$AE,0)-1,0)","v":0.16300000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A47,FQ!$AE:$AE,0)-1,0)","v":7.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A47,FQ!$AE:$AE,0)-1,0)","v":5.7380000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A47,FQ!$AE:$AE,0)-1,0)","v":22.937000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":1.3355903437646399},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":11.3295360967316},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":20208.894008033101},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":168.62469449056201},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":49619.240925621198},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":1165.55428867363},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":141.63795684123801},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.64071992107780196},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.15137575590032401},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.4221246080421599},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.7544608852361501},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":19.032090910164399},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":8.8992199010887604E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.86071412873568798},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.3433592479697598},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":66.980602089816998},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.157284768439235},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.36300088684939202},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.011"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.263"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":800.43775659474898},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.57539444384755},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":0.46204909647995002},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.37895003429902402},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.82382421916640602},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.003"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.2091043828375001},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.17649112877161399},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.002"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":4.7660932331370702},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.117633873328406},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.132"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":22.990256992175102},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":5.2256447280410803E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.1288556990763499},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":0.19422081050876699},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.003"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":7.54948528825229E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.3387433685423904E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.0381864294931301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.18178689309376E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":43.5237032570911},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.11847174905912},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":9.6873475533201005E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.1412026142237197},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":9.1186904597052004E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.13551258921200299},{"$f":"OFFSET(SQ!$A$1,MATCH($A47,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":2.2079397016571702E-3}]
[{"$f":"SQ!C49","v":"047SMPL.d"},{"$f":"SQ!F49","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A48,FQ!$AE:$AE,0)-1,0)","v":3.161},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A48,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A48,FQ!$AE:$AE,0)-1,0)","v":423.911},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A48,FQ!$AE:$AE,0)-1,0)","v":10378.043},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.168"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":363.97680721863298},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":81143.830222186894},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":220.45967723072999},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":79638.635610576704},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":4526.9893569680999},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":645.36209683667005},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.4973107795672904},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.54595411146047},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":37.551870201168803},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":38.244685076026897},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":419.39433197928003},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.1313041891186699},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":23.9018586391722},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":24.4683533564748},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":134.359440149899},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":8.4862455083365607},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":8.1316680625832305},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.332"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":13.9670271783367},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":19352.069932225899},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":25.012052403413101},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.1832128517448},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.042"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.38049750424508},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":14.6626342993228},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.120"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.4197393347314899},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":0.46650159238074901},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.138"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.094"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":13.0895408928568},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.2504282190425595},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<5.277"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":236.12288688711499},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.6372707899152199},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":38.266289339581398},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":3.5827550725515902},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.139"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.059"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.015"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.044"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.173570883652684},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.47132100912875002},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":59.468920457143398},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.2088061922448698},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.020"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.60221338536613},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.71300626019811397},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.019"},{"$f":"OFFSET(SQ!$A$1,MATCH($A48,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"}]
[{"$f":"SQ!C50","v":"048SMPL.d"},{"$f":"SQ!F50","v":"Blank control2"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A49,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A49,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A49,FQ!$AE:$AE,0)-1,0)","v":-1.524},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A49,FQ!$AE:$AE,0)-1,0)","v":0.39},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.69262247415840095},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-358.70099849266001},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.42131315942116598},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":236.21960379947501},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0835241863831397},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-12.320416542132699},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.8689508567091999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.1304574497763301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.3072382505052997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5540944025919397E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.35692561536431},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.7291371634993799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.0939054802140901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9899789419343898E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.218347688796206},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.3485399390022698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6941632106343298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-3.87260188280653E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":4.0500945146277501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":15.2680514187578},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0240800869242801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6346593792729501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":3.1325165965002201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7190520436239903E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":9.3452316447226597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.07729225926167E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3362286689054201E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8353686608873697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6938344109737101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.47405699542373397},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":3.6902404067675398E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.77862403014527E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.4423698390810301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5859503765949599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.01758829446069E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.45754624908168601},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.43817138539697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":4.9770105088035699E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.14044240674632E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8776360046398405E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A49,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5724610687465703E-6}]
[{"$f":"SQ!C51","v":"049SMPL.d"},{"$f":"SQ!F51","v":"Reagent blank10"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A50,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A50,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A50,FQ!$AE:$AE,0)-1,0)","v":-1.5640000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A50,FQ!$AE:$AE,0)-1,0)","v":0.51400000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.15732618104365201},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":8.2618561995537604E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-163.98176864622599},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":4.44716803467783},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":686.41989155757801},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4381629913793299},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-6.6460736332382204},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-9.7815203176868908E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.5498771390883801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.1551852430543502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5166522866135601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":2.5317466964465098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.6677699326031301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":3.0367089418104901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.26557800499591E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.39208836369484001},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.51097313247544E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.52297042860546E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9931137751869801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":5.80450386312486},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.3301002498170102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1735082254698002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1327836439606099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":8.0496554800625196E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5463030397720101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":2.3445622655732198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.8007642027341499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2602675442524798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3260008857573302E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.412498227252797},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4457439833640499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":8.5524015037946304E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.4780219894864001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7611558387883801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.6950878745832099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.9950427540060001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.50046495497280197},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.6837870300724797E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":4.6686509344759403E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1672573252548297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":5.1544689525962098E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-9.9944781811860703E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A50,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":-6.0135236013692001E-6}]
[{"$f":"SQ!C52","v":"050SMPL.d"},{"$f":"SQ!F52","v":"QC1-2"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A51,FQ!$AE:$AE,0)-1,0)","v":8.5139999999999993},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A51,FQ!$AE:$AE,0)-1,0)","v":0.21099999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A51,FQ!$AE:$AE,0)-1,0)","v":19.602},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A51,FQ!$AE:$AE,0)-1,0)","v":8.9},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":180.71063948691099},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":8.0252794516439199},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-137.37555664339999},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":449.659182698216},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":290.11538547639702},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":509.50512352205499},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":378.75559869518997},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1402861051569E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.44575373577775101},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.44711397777342698},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.46308885528959698},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":17.709543709225699},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.20697248682051},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.88745615928769706},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":4.5006272265695602},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":42.7624616138785},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.44235442630143601},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0786755951357899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.2798906623553399},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":7.9895749025331702},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":7.2606901628067},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":21.224836651076298},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":9.8802994028837493E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6586212441304797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":3.0144324505319499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.7972641457068199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7920896373623203E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":0.81312160603667305},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.210932860609923},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":0.211447995002015},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1281577971100503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.1993426376812397E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.384230682616858},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.45216308043174602},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.1326937795210901},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":6.4137704908775897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.0825343919059203E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-9.7743487724956493E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2160953596818799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.0440004505513901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.51487172990338004},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.0178733120950797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.18995279692808301},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.9855488339795799},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.237213785892985},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.20103154901252401},{"$f":"OFFSET(SQ!$A$1,MATCH($A51,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.0606714830781203E-2}]
[{"$f":"SQ!C53","v":"051SMPL.d"},{"$f":"SQ!F53","v":"QC2-2"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A52,FQ!$AE:$AE,0)-1,0)","v":-1.6E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A52,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A52,FQ!$AE:$AE,0)-1,0)","v":-0.64400000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A52,FQ!$AE:$AE,0)-1,0)","v":0.32500000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.704938386543706},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.60461058794330003},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":150.581516610791},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":3.1913002086678199},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":376.19701182890401},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-8.9023254577374509},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":56.251809267874897},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.6648107431519001},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.7052281861173798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.7334788338699199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":-4.9341544908616003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.12016771202649799},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.84603823941218E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8567023310443499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2365606943680401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.105304081824557},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":4.7448088556275303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.58946258788505801},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5520864190259501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":9.9416565037832196},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.8197510515443901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":4.5993333721576004},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.59789010989418},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.69068760481784297},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.56196853229407795},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":2.9202925226477602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.00164076808796},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.47512606879184399},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":0.41932739529072899},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.36329439199221703},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9413495726622698E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.26148491966328E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":4.2400741980595802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":-2.83477971472861E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.433219167351866},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4557417662253999E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.41878369763303802},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":9.1307112455599704},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.53795003954929499},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":3.2804794880078798E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9010837187002804E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.4329933978373699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":4.54339877296005E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A52,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.12414314160483E-4}]
[{"$f":"SQ!C54","v":"052SMPL.d"},{"$f":"SQ!F54","v":"QC3-2"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A53,FQ!$AE:$AE,0)-1,0)","v":-1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A53,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A53,FQ!$AE:$AE,0)-1,0)","v":-1.367},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A53,FQ!$AE:$AE,0)-1,0)","v":0.104},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.62855404319181296},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-296.41904436941002},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.9525249220734899},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":200.17686173347499},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":2.4833936787439201},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-6.3590943021714796},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9698872312229499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":9.7401640411938304E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":4.1312705881027402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1916488555316203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3590169774331299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.4035511134293104E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2727181055459402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.64867834007262E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":4.6464299743600203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6532315536437999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":3.4749487726388798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.0402157222646201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":318.62301679265801},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8551044963291392E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6753134107673599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":5.9632157250182304E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":2.2840376104703301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7404707630923403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":8.7346395854242791E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":8.5488571819327601E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6665202442374597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-9.2764177251645401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":33.841728693253103},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":3.5666764259081001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5955814053952599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.7925894127242199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.0009296133950602E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":5.68720106045061E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.24807267400656E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.4175038210154402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":2.6605452750054499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.123530768207625},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":4.7704537063917199E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.33480607164631E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.2360013374495798E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.07797854691435E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A53,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C55","v":"053SMPL.d"},{"$f":"SQ!F55","v":"NMS 59"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A54,FQ!$AE:$AE,0)-1,0)","v":51.841999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A54,FQ!$AE:$AE,0)-1,0)","v":1.2410000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A54,FQ!$AE:$AE,0)-1,0)","v":122.699},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A54,FQ!$AE:$AE,0)-1,0)","v":49.094999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":1066.0979902619099},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":41.963181520291499},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-160.76501073614901},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2685.1220200623402},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":197.44593418628301},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":8.0472091744805105},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":456.16334717629002},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":3.3811079688186599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.5513202250309801},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.4540452764267902},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.4118927509558499},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.18464568662683201},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.28653678533005},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":8.6202690506137006E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":25.760848364571402},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5999333842614194E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.67738425212299},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":5.94293313973056E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":12.850193590936801},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":51.758871147536098},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":4.28264701215919},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":128.61641917862099},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":25.0298938809695},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":4.7715397644139201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-6.3966246420843103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":8.2774975791640604E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-6.0427004652145802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.9690419437856699},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":2.55704864500621E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":1.26884163717301},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3416223834855101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-9.2685700658181398E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":4.8731796900789899},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":9.0067545671894594E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":12.3625403674039},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":9.910593523627261E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":3.2108011845518899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":2.72211949359577E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.6352157873009202E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.18475898239088701},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.4271838654837398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.18525671048891},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":12.1633655047833},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.3742392270927699},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.27737775909294},{"$f":"OFFSET(SQ!$A$1,MATCH($A54,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":0.246158444793316}]
[{"$f":"SQ!C56","v":"054SMPL.d"},{"$f":"SQ!F56","v":"Reagent blank11"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A55,FQ!$AE:$AE,0)-1,0)","v":2.4E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A55,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A55,FQ!$AE:$AE,0)-1,0)","v":3.2989999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A55,FQ!$AE:$AE,0)-1,0)","v":0.16700000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.60929411412998402},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.27742890027916201},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-276.97987812893001},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":5.5547202769597304},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":512.75713399519498},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-0.34898946579707502},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1795597859799098},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-2.95788282283268E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.3343821018902403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":5.9037977699753899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8347508976790701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.21818680598010001},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-7.4102289880909099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":3.0199386895454601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.3019770883525903E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.107247553124661},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.3007795671271E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-9.1790426191992302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":7.4151228821325698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.1346417853512E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":6.9719080777642102},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.1402487152841601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":5.6585715556822704E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-6.3544481705795604E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":2.9111457829189302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.8391031578468301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0121690543593802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":8.6611418697088202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":2.16457531289913E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.01005521397051E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4595116257818399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":1.55580544403532},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":3.9215449753577801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.5672414399672901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":7.1082786378447895E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.36995609148834E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-9.6298287296932004E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.384681923819434},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.27781631156573E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":8.0219860726490498E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":6.1296864155832098E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":3.0726365810407299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":9.1112981142287094E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A55,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":2.7978682689768397E-4}]
[{"$f":"SQ!C57","v":"055SMPL.d"},{"$f":"SQ!F57","v":"Blank control3"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A56,FQ!$AE:$AE,0)-1,0)","v":-5.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A56,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A56,FQ!$AE:$AE,0)-1,0)","v":-8.2000000000000003E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A56,FQ!$AE:$AE,0)-1,0)","v":0.20300000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.78067445416023296},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":2113.8075826409399},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":5.8697407802854897},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":2092.28685812514},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":122.018874927563},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":25.275822426967299},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.13026997098266299},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.0066080000694003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.22574452766891301},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.12965835835675099},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.3597532227772999},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":5.6493818431130804E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.5085837279854999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":7.5960260606831698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.22244254279942799},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":8.9858031289419306E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.86456929372627E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":3.7015002829832201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":8.7131655300839397E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":26.975524187702799},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.0430893553099902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.1643545156889999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.64435675611399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":5.82437757862373E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":4.1731315588822001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":3.4003987120378297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.4716630472016E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":7.3131058499069299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.7120978106384701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.66263577333549195},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.05183047319981E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":9.8320505104675607E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":5.9927980729213503E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.1647428209728402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.56406076390156E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.9524787614735599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.22491686541486899},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.4073310040638197E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.7580955767890601E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":5.6442292293266199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":4.7735695160255201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":3.2550438073356801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A56,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C58","v":"056SMPL.d"},{"$f":"SQ!F58","v":"19286426-001"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A57,FQ!$AE:$AE,0)-1,0)","v":-1.2E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A57,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A57,FQ!$AE:$AE,0)-1,0)","v":-0.66500000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A57,FQ!$AE:$AE,0)-1,0)","v":0.14699999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.57294383616963296},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":40.931436531421198},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.62733854132069},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-80.844830297805203},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5100114602574504},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2050022898977599},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.6894468618009398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.9931103964862299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.1923000960916102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.38488840290565E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.23776311292553001},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":6.5029074597298801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.7818699055409198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-6.6097789649473401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":4.4715525326841803E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.67770593172858E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3973587764041498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-1.97613077750061E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4601439211053302},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.57898369601091E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-5.69793584796643E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.3878076489210901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4022953328407901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2432482713328395E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.16430346228544E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":3.39957918179954E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7167188677014198E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.1556990611150599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.36007186909311001},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.0623725972995601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3884738980283197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6373139868205801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-7.7465286641488395E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":2.7415549145091402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2874697473503401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.36468803726893E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8516222490281998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-2.69705155421897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.20542352858952E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A57,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C59","v":"057SMPL.d"},{"$f":"SQ!F59","v":"19286426-002"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A58,FQ!$AE:$AE,0)-1,0)","v":-1.4E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A58,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A58,FQ!$AE:$AE,0)-1,0)","v":-0.97199999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A58,FQ!$AE:$AE,0)-1,0)","v":0.34300000000000003},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.156829179004355},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-8.0976355392201799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":83.501196534003199},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-0.18798500070090099},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":49.137132311690003},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0807461533227203},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2449367247137104},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5074037146459403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.6176120880817E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.7740591384102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.5625476506253398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.24111093314125401},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.8651842334056399E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.6911652126535101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4481976750854301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1852101256193101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":3.53146926644907E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-5.4310921048946297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.79084168705515E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.5219657376365202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2935067509470102},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":8.9986436088454105E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.80260514346303E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2112953301473398E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0577118017894298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9421906326296001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-7.3295263004024697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":7.5424643187473799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.3761477310425995E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.169250135344166},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-6.3411865180481501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9274594723134396E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-7.8764535778008899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1319424340593193E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4892819626929897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.31317940817651E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9444643140747902E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1958049798510601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2173876675242398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.23696536858316E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A58,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C60","v":"058SMPL.d"},{"$f":"SQ!F60","v":"19286426-003"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A59,FQ!$AE:$AE,0)-1,0)","v":-1.6E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A59,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A59,FQ!$AE:$AE,0)-1,0)","v":-1.085},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A59,FQ!$AE:$AE,0)-1,0)","v":0.48399999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":1.01775032914504},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.32660887626479201},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-69.898320304579102},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9434270314094801},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-241.64099825780301},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8269236767893302},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":1.5843288584170501},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.4908382569827403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8327992314239708E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3516849611446702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.6086776345952901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.55216726910575098},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.13925896188029E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.0377747991117899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.19171546311938E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-8.6247868384080301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1551215216098503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":6.7286683389814602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5488055443014899},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":7.7643618598775598E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9707286289431399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3318108901493699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2821267330364599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5364656894546799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0682520310518399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.48462207435896E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.1176054501460899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":6.2731138717565197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.16982196515511},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1738262489110399E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":6.7430702525778097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-4.9967187916331504E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-7.8855266375631096E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2490855386862898E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3748918135840998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8519630807241401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.17695107384018E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.42483388297394E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.35259474592843E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5238018807107999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A59,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C61","v":"059SMPL.d"},{"$f":"SQ!F61","v":"19286426-004"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A60,FQ!$AE:$AE,0)-1,0)","v":-1.4999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A60,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A60,FQ!$AE:$AE,0)-1,0)","v":-1.248},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A60,FQ!$AE:$AE,0)-1,0)","v":1.2589999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":2.0424807238025098},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":1.1068874639424999},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-110.617407709334},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1716394990351899},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-437.89653671064002},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5059748262398802},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9840302647682599},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.1235351654538107E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.7899057366056199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":-2.92974912495858E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.7587895346219303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.13700399442049699},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.24299602668996E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.8158198707270597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3068436505789901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1.8062742920763899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.32814042946602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.07580070605872E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3329478170536201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8784705711633398},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.0956392946013602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7818915642756201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.0807018952387204E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5512096023245601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5251726393568401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6633085426632301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.8411683047797099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":4.9881732412727503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.5838558985527801E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.13719418307745601},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":4.3435525034208998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":3.0678765687231898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7247086552804601E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-8.5372363181579101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":8.3439073865856301E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0691497317163099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.13165732733739299},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6606183185292601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7519012417714199E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.10275139350303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-2.40067049658015E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6561867273503799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A60,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C62","v":"060SMPL.d"},{"$f":"SQ!F62","v":"19286426-005"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A61,FQ!$AE:$AE,0)-1,0)","v":-1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A61,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A61,FQ!$AE:$AE,0)-1,0)","v":-1.226},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A61,FQ!$AE:$AE,0)-1,0)","v":0.94399999999999995},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.758024478345102},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-76.525089045549706},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8516995484908598},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":336.91746432616799},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-13.386157681103301},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":5.1623464074139003},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5581558814900302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":5.6751805315864304E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.12643629094331E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":9.51488627293187E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":5.7133004210687698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.16435230757484E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.54002521789946E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5130845751356601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":3.8009951259526699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":7.9664295561296707E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.71851905972076E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5081637156821595E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2915723765259601},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5959757454041001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7774209288806097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.7203489119341403E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1858623664500802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5684105645734801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.80549792462441E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":2.37524238011501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":8.10905485824674E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5059592278733502E-7},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.35506262231453101},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.10107487125978E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8830392025981202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7459592555642298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.5381868129514501E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3856560740613302E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7669786133657998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0601460994735802E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7851332277234899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8627317281723999E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1683565989633501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2749444000563797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7136355397538101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A61,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.2144582851602599E-5}]
[{"$f":"SQ!C63","v":"061SMPL.d"},{"$f":"SQ!F63","v":"19286426-006"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A62,FQ!$AE:$AE,0)-1,0)","v":-1.7999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A62,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A62,FQ!$AE:$AE,0)-1,0)","v":-1.302},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A62,FQ!$AE:$AE,0)-1,0)","v":1.526},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.760236281691402},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.91521662874446097},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-68.3183416208756},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.47485627517565},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":212.538722994247},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-7.3559174801228302},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1220054569305598},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-5.01463943678317E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":9.4549932567033999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.81266439490935E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.58915834169957E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.15043216189403799},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.39321685610857E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":8.4452036297394006E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.5817713927441E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.117050101393971},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0392172138836603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.3482600960083E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5584408337568298},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.5208343102980101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3227340966136601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.1410657256581102E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2856177567077401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5299417179133802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.13871824222495E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":2.5529175549324603E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9369915405294403E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.4357405820243798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3648646377815697E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":6.3416564096460595E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8470543911524802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2423341359052796E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-5.4556899954364899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5766716522354999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.106143926992859},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4320399125142701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.1774017139406599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.59950435967317E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9884266637212002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8510512285248001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A62,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C64","v":"062SMPL.d"},{"$f":"SQ!F64","v":"19286426-007"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A63,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A63,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A63,FQ!$AE:$AE,0)-1,0)","v":-1.3129999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A63,FQ!$AE:$AE,0)-1,0)","v":2.7080000000000002},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.74899526561985297},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":1.8258125133629299},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-103.20958654142601},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.5723047731001101},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-306.88266991675101},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-7.4634995119967602},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":7.1817849800448998},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3138679060120399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2715552717453799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.10976497124102801},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.129420405484486},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.28737284007},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":9.9505544323867895E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.16699504542369401},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.1720099920434401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.20922609833151101},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.3569124606410601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":9.6832743520774297E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.8689053982021901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5319091741921104E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1306696230444198},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.0084554613012503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1440550747193396E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":4.3381278296259797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0253156790188201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.83567181576315E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.85814368163647E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5765592634744101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.3683816885506901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":6.6400862750498196E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.0878698604444401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.47687390527718498},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":5.9952868413603796E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.62165267060361E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4746786561950099E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":-7.8126683078602196E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4105891474400999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.2528660872432501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.105162127388673},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7174251015931301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1830577676428201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.79112276931747E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.86473584403366E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9441427131799501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A63,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C65","v":"063SMPL.d"},{"$f":"SQ!F65","v":"19286426-008"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A64,FQ!$AE:$AE,0)-1,0)","v":-1.7000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A64,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A64,FQ!$AE:$AE,0)-1,0)","v":-1.3819999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A64,FQ!$AE:$AE,0)-1,0)","v":1.139},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.46082358635628301},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.52144339087492697},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-14.789031327780799},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0147260855427902},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":698.37702405646098},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1540635054894999},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":0.64995578515212304},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1804741050557201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.16085853002375E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.6288091986637299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.6898611475156997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.151741492396407},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8262637327984899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.4400901612495698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.0559756196997901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-8.0997837394790301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-7.66214461864702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":4.5799545059044197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.1579627024928402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-4.92646086507607},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.7220836707189401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2052664729299601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.19815477903161E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1718489612472698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.59854242777433E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8111974740462502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.0551110479311301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-8.5387902830085698E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.37793879258756102},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4093129215789602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7398027037874798E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":1.7838173522191801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.40725843031943E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.9725217750341801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0620346232599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.12179158781199401},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4855675516584502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.17949161696387E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9902052038234097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2021499017629504E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9142940309322199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A64,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C66","v":"064SMPL.d"},{"$f":"SQ!F66","v":"19286426-009"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A65,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A65,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A65,FQ!$AE:$AE,0)-1,0)","v":-1.36},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A65,FQ!$AE:$AE,0)-1,0)","v":1.0680000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.46561168732029201},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":1.1581586814774001},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-65.699115471445594},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.25211443064756},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":242.435200399211},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6714055283828202},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.91285326688747},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4559112967322801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-8.3420505805847697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.24808468170487E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.4925699463072399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.17542021075971501},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2130009859051701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.8352091007377394E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-5.25463044588108E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":2.8355914886362198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3653715115524596E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":5.9246915214927996E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.5316953547642401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6194926854637002},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.3193307897595699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1358314078349E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.8901574316485001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1326917579419401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":3.81698025089019E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1905413780560698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1582620097078301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":2.33883908944823E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.80668637162259E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-8.3866134913015001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.22802483474938001},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.93274985613688E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-7.8257008605326101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-3.69985326268206E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":4.6313121440964803E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4842695346867199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.7097764478513998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0808663040592001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.116404601861646},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.35885744185385E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.3364371585016105E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9608767021449302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4101000998318998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9670300293798899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A65,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.3830221146173099E-5}]
[{"$f":"SQ!C67","v":"065SMPL.d"},{"$f":"SQ!F67","v":"19286426-010"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A66,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A66,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A66,FQ!$AE:$AE,0)-1,0)","v":-1.371},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A66,FQ!$AE:$AE,0)-1,0)","v":16.254000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":10.2138216125671},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-140.24662886139001},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.74340012927380394},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":171.66428053306501},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8016203597834597},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8903679597884293E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.04198032815289E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7153493220310197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":7.1431657746650898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.9447536971464702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.27038052447655803},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3633262693556001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":3.1797252235001502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":4.1719434085151802E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":9.0138647022067406E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":6.6320177471237696E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7239164357813302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8287588486283399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3090787823456697E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2896663011826703},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":3.2364998052594702E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.38360005612561E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.1147030601966601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5048439061731197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.24597146022117E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2663037988181201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.7379552802921001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.1386657491701899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.2850824995313602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.49689248921856799},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4134429848171596E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":4.9244806163176403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2432710203754096E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.1695149004875802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2510702324114E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":7.9320341402283904E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.11327092825150301},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0826403366021399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.0047212819483201E-7},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.6966438386480801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7022820859217397E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.01871570323912E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A66,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.23312473399421E-4}]
[{"$f":"SQ!C68","v":"066SMPL.d"},{"$f":"SQ!F68","v":"19286426-011"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A67,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A67,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A67,FQ!$AE:$AE,0)-1,0)","v":-1.375},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A67,FQ!$AE:$AE,0)-1,0)","v":2.0619999999999998},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.60282224154328401},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":1.16920563161325},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-119.914427325017},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":4.4433468856183298},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":209.65166012568801},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":0.68488175681207497},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-2.58991605464356},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4170938953389599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1755985788824599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.85637197616196E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":7.4047949742315697E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-8.7941538459466208E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.36635885334298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.5632803826223399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-9.0579966267413797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.161416852211101},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.77752740731156E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6815317711616298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.6476526811472198E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2964287688720099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-7.0775172754493099},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":3.8482471663735902E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.5616487586259901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.6724956583619001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3380951191996702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.8671439245234001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2115294110476702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.68780736819757E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":9.0544610847182599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.9484577898933398E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":3.4165884962744199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.47360378818073001},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.0849629006687901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-2.95981062606192E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2448091278955403E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4859738243281999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":7.8346614465408498E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":7.40446754546393E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.11698127106152},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7052773463793499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8776991812948103E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.91590295206914E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1332983406930602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0134483342374899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A67,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.1019762150678202E-5}]
[{"$f":"SQ!C69","v":"067SMPL.d"},{"$f":"SQ!F69","v":"19286426-012"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A68,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A68,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A68,FQ!$AE:$AE,0)-1,0)","v":-1.3939999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A68,FQ!$AE:$AE,0)-1,0)","v":1.1850000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.30650675582356002},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.92868399816966296},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-137.994185808997},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.55136221961427},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":372.72708314184803},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-7.3655836208841698},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":1.6387204889060301},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7082617181093999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.0091096720182401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.51817196016197E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.60884524371753E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9650714707744697E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1246509982186396E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.5511707858797895E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1036977825444597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.238011531201237},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0276717413783901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":4.6371624734710204E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4735340818618106E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-8.2410636601880896},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-7.3361568266781E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7226367767162899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.1900915077375299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7263704282824799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2538626892640301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5520384489605499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":3.7114857279018298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3867715003163701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.4595377479449201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.30396242801794998},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.0738059391357899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8646725375306602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1689610818715801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":6.0319574796144804E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.142118678402917},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5329406558501699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.16594259886363E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8365181767424901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2696324923113997E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0862355259846801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A68,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C70","v":"068SMPL.d"},{"$f":"SQ!F70","v":"19286426-013"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A69,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A69,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A69,FQ!$AE:$AE,0)-1,0)","v":-1.444},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A69,FQ!$AE:$AE,0)-1,0)","v":0.749},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.43469831968887801},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.124377366545111},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-293.66738732201901},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7867934039626301},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":305.40119640381999},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-13.872003444314},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1164152459869596},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4363384793711099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.2986134611166398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.01833432745887E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.7036675935696498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0181144145649399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-7.5028166897307897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.1962189110827399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.3630372598170201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.171461784153807},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3369141989615699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-3.05349037226412E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6013694010272898E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5692928610012604},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.5761346297465399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4195631586534702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":8.5527346599234197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2649764935650597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.53605362439056E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.77555238602476E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.9594963688157399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9633207534335602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.192993489986514},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2559022808995903E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9051402782163603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7627979684141602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.33125799534978E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.22632300702026E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0511949351100099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.105539196124048},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.76331461374042E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9505333970677597E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0932527085190702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0665853382798097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0892666194180601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A69,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.1528958343013002E-5}]
[{"$f":"SQ!C71","v":"069SMPL.d"},{"$f":"SQ!F71","v":"19286426-014"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A70,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A70,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A70,FQ!$AE:$AE,0)-1,0)","v":-1.4790000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A70,FQ!$AE:$AE,0)-1,0)","v":0.59},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.59091761551120903},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.40567842540740301},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-197.24843853780001},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":7.6889489317300899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":27.403493771698798},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-9.5696283910300703},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-0.67969646774262704},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.5711078451230797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-7.9393383566158694E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.6167176141092501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.32647515543349E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5655002084477E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.1219012410213802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.5221185135233199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.61875165560186E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":3.5378648125250399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":3.8968579773675802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8863521380381998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.5683980392107297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":7.3399429065673595E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-10.6308025157174},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6216628963355701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-8.7405197784197908E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.1821726454720699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2306397667641001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.8527873146247301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5819853207174298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2351531975500501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":9.4186509979829494E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.08113155306303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-1.48286242197377E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.24037063448790899},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.17654678946264E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9079856633922597E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.02138221994895E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.7085362540903998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.131585624671298},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8711733310004699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.17714075146749E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2009585021870999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2791621734742302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0494511517217101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A70,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C72","v":"070SMPL.d"},{"$f":"SQ!F72","v":"19286426-015"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A71,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A71,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A71,FQ!$AE:$AE,0)-1,0)","v":-1.476},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A71,FQ!$AE:$AE,0)-1,0)","v":0.36699999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.30080649938221699},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.15875044381682801},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-204.93133536206099},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6317325486530201},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":100.59176846186401},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6208351338937499},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":4.2251566134132101},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.7936771384070608E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8048729080406306E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.65081460275637E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.4565700201501599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-2.62044895718423E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-7.4062404098164901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.8441641903820802E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.30750133674061E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.20524360227247401},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6991196980278499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":9.7236601682690002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.85673434102852E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5479132611223997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-10.6730429892161},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6515108930045201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.35058928865306E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.21336983983548E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9003628553673998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5535889047850302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9142811851389401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.3602698777993901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9941475554245899E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.50278588958024995},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4282720601333998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0881331097740396E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.33766578567205E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.00585197709172E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.10546181462802},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.88542361832979E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8007687558300699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4985596809511003E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1383025500431101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A71,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C73","v":"071SMPL.d"},{"$f":"SQ!F73","v":"19286426-016"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A72,FQ!$AE:$AE,0)-1,0)","v":-1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A72,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A72,FQ!$AE:$AE,0)-1,0)","v":-1.508},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A72,FQ!$AE:$AE,0)-1,0)","v":0.91700000000000004},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.635029569377772},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-281.47544405334799},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.80552343609646},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":210.951678610931},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-14.877706407946301},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5247564934269304},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2541666244788903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-1.10616304029982E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.4064323829496701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.7855277804864601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.233899305234994},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2198022850435399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.9960150254132601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":7.60773507220698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1925525194049299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4746745869969799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8729080631059298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-6.5975713857224505E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-7.9120929694571203},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2609038406376599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.20748706830699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.2812152043393402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0188229122295697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5162721550133198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.5402357792331903E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0900436350563499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1941935431882402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.39078558534653601},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8908141951543809E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.41433109146785E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.0966242312382801E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.13311886267510201},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6306951503706499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-7.5044590203584304E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3141937560472396E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3643356385051303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0895967843570599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A72,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C74","v":"072SMPL.d"},{"$f":"SQ!F74","v":"19286426-017"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A73,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A73,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A73,FQ!$AE:$AE,0)-1,0)","v":-1.4890000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A73,FQ!$AE:$AE,0)-1,0)","v":4.343},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.15100878618063801},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":3.8007039587851601},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-216.149891022719},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":3.8562853550188301},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":39.553394273198997},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":2.3176845352248798},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-0.82201875963400906},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-1.39906348965352E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.38810936746625E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":9.0615640025609508E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":6.7551678483896593E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.57605679374248497},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4299502882570699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.4147871379164899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.9872115481993999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-1.35098821531237E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6820959344048902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-7.7031870752383696E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8289706628246299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4558656718680698},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.7018372520199201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2665642285394396E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1708037328257798E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6456653660070699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.94328134863951E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.82692996761087E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.1328669675700299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.6521633631717602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.24096525023062201},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.06018213960042E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0189859111316802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3318170073350899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9841324650426304E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.0973468914928101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.114458656104776},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.03255737462029E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.17797121357317E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0654718219375E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7861658646661102E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0953670825216099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A73,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.0666566986465102E-5}]
[{"$f":"SQ!C75","v":"073SMPL.d"},{"$f":"SQ!F75","v":"19286426-018"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A74,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A74,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A74,FQ!$AE:$AE,0)-1,0)","v":-1.526},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A74,FQ!$AE:$AE,0)-1,0)","v":4.1139999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":2.9386822975193398},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":2.75914586945107},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-322.54507994164101},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.63618521778449699},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":628.01729213575197},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-15.263643987612101},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":2.8735437542697202},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7027802606029398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-9.9127491224868195E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.7332276609676203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.2118738175292401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.345349552024107},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.7661583879939498E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.0817397414501199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5238660115122499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.102878566737324},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.5351845452156502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0258479395680899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7680636049779201E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4570358989267298},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6382874220407399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":8.5364821169174009E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.5993468005425101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-9.0176074144511997E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8584952436674599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9098141699986902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.90477505429944E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":3.0388719823007702E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.3073252113581599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.5311342833586901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0723303480283299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2394598320507001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-2.09334904399994E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-2.30773996752691E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2562974508641201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":7.1396093048285106E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.123712070589236},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8752161059000299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.27792455824306E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.2758544103096501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4230844225651798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1133480396335001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A74,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C76","v":"074SMPL.d"},{"$f":"SQ!F76","v":"19286426-019"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A75,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A75,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A75,FQ!$AE:$AE,0)-1,0)","v":-1.532},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A75,FQ!$AE:$AE,0)-1,0)","v":3.036},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":3.2698977133990499},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-312.154857909246},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.83042628447925904},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":551.64806084084603},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-0.85159460733400705},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":0.26279384743744},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8337158296575399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.17295631586209E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.2591362076641602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.8777390310254997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.310808538607662},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":3.0771056752912899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.5828673146180905E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6093259283981899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.18772675058887001},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.07373166636126E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-7.7485136326430103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3365991379207998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1774967349878702},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.6024298238263401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0195360742298499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.62033521352003E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4167980652586201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.54465583170854E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4963302450600099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":1.2237867821963101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.33597688357486E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1544742563929503E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.407822158641582},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2188874848016301E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8695871295327595E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-9.5163554995427802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7029020166671501E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0370019377001999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.10032722846916101},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6131815203176201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1844269829123599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7741865112580798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3551062143987497E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0971440750570602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A75,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C77","v":"075SMPL.d"},{"$f":"SQ!F77","v":"19286426-020"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A76,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A76,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A76,FQ!$AE:$AE,0)-1,0)","v":-1.5249999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A76,FQ!$AE:$AE,0)-1,0)","v":2.077},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.59830338771091396},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":1.75517665699724},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-337.378546130887},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-0.45084237246483},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":227.00693756269399},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-10.4874675869941},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":1.6747822344444501},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.7460097336352501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.98312818676614E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.93549446665294E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.1783136551051799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.57598143015000403},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.26486757429352E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":3.7542097904223699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.7265135865975301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-3.14218014998872E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.4241210033365799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7220779212731402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1115159003561101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-9.9755234699290405},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":5.7266408923935301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3246741370461398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.54755475143148E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-1.51783359453523E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.8659414536072801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.98137948572527E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.9021004375504501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.25775722058702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.4209661679568299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.426431472874088},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4185209528912005E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.2554012260785E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6730837842082296E-7},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4860020240075201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":7.8175843087171802E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.13268316324768101},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.08342364472192E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":5.8171436180499003E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.6294584125523201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2046449701532299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1745413799778903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A76,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.1004948608940702E-5}]
[{"$f":"SQ!C78","v":"076SMPL.d"},{"$f":"SQ!F78","v":"Acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A77,FQ!$AE:$AE,0)-1,0)","v":3.5950000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A77,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A77,FQ!$AE:$AE,0)-1,0)","v":105.819},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A77,FQ!$AE:$AE,0)-1,0)","v":1094.3019999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":17.233853730411798},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":116.475007902665},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":43427.172986949401},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":445.05743097929201},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":50255.8412883727},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":3382.83367384756},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":819.81668921268601},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.6706234257749397},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.1482252092128702},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":16.752729919690701},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":11.3801546491756},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":145.912494867698},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.55229758820255204},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.0741150657480496},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":4.95617569323257},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1483.23879086582},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.165396870045177},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":9.4408598909078005},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.087"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":21.6393257885012},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":12855.992526432099},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":11.052195864358},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":21.548564595905699},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.10447574547437399},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.39414249714639799},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.68221455115740604},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.029"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.69494189646901905},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":8.5372818944151003},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.034"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.023"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":55.295838018337498},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.9403251996505499},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<1.292"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":366.55762857018902},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.285290612265656},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":0.52548842602181201},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.1610767759464702},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.47350211574464},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":0.19375424403893399},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":2.4713899932923899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.011"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.006"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.18845282482947E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.009"},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":200.581469967278},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":13.217170488912},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":7.9592670071809798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":0.70352540505800798},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.48011963389302198},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.31247045854595601},{"$f":"OFFSET(SQ!$A$1,MATCH($A77,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.5917458101580601E-2}]
[{"$f":"SQ!C79","v":"077SMPL.d"},{"$f":"SQ!F79","v":"Base rinse"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A78,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A78,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A78,FQ!$AE:$AE,0)-1,0)","v":0.111},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A78,FQ!$AE:$AE,0)-1,0)","v":16.100000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":2.29447018916848},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":26.858786368691401},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":24683.478791399801},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":520.59801501121603},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":38264.809670021699},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":1648.8494909719},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":372.44801599731198},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.59468357231837199},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.314767804301745},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.59782011121627},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.97456863735299},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":19.686809344137199},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":7.7170671720803594E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.91481899121501498},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.5804134664964},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":35.613245490559599},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":9.4906221309018796E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.18946219938594},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.006"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":0.13872723953464},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":156.10810958857499},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":0.55478271189651895},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":0.180856398791855},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.565610954885857},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.15075259539018901},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.43815949295874002},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8792805188038499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":3.1665087808000598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.07454049045922},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.5999195783714101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.029"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":4.4784484797936202},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":5.4767484901668299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":1.45784729619875E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":7.9559071689342602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.4086914705116499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":2.7074944551889499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":4.0206156332712002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.31505204821905E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":12.974470486867901},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.6422574354642898E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":2.3703898005241601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":0.916690058554814},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.64530674048761E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.35652585451211899},{"$f":"OFFSET(SQ!$A$1,MATCH($A78,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.0199321418556001E-3}]
[{"$f":"SQ!C80","v":"078SMPL.d"},{"$f":"SQ!F80","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A79,FQ!$AE:$AE,0)-1,0)","v":4.2329999999999997},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A79,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A79,FQ!$AE:$AE,0)-1,0)","v":353.90199999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A79,FQ!$AE:$AE,0)-1,0)","v":10003.688},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.243"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":558.882567875419},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":92704.581841829204},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":157.01851703379199},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":103730.340090731},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":6364.2336699233101},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":916.09578688639704},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.8744370368817798},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.7789387607233},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":44.420953152054103},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":42.385206895276298},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":398.464774861298},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.19015975768527},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":25.551014835127798},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":17.664731575243199},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":67.364676157157007},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":11.0802245094239},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.53610757328592},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.312"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":78.495726945744707},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":4794.8138462704901},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":23.247225615078801},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.19770859778165},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.039"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.5770770592599899},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":3.0794777681040499},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.111"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.9556270646040499},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.039"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.127"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.087"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":10.3983773735358},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.2602969088845501},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<4.861"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":459.96490495897802},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.21446285549329799},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":3.0119737715272099},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.023"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.128"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.055"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.040"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.024"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.012"},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.325583475939347},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":85.509174840706606},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.7363230712527602},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.11989906852672801},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":0.79498623992488804},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.36017955475052799},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.14450283034650299},{"$f":"OFFSET(SQ!$A$1,MATCH($A79,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"}]
[{"$f":"SQ!C81","v":"079SMPL.d"},{"$f":"SQ!F81","v":"Blank control4"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A80,FQ!$AE:$AE,0)-1,0)","v":-2.1999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A80,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A80,FQ!$AE:$AE,0)-1,0)","v":-1.5569999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A80,FQ!$AE:$AE,0)-1,0)","v":0.85799999999999998},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.15882922732186999},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.35311748788039199},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-174.89496021349501},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":8.8354840455750896},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":94.823267182356602},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":5.6156938629900699},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1782012929093399},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.3736835251514099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":8.9651817322323902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":7.8579740419963601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.5860825352282501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.70163607809190598},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2910611994337001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":8.6146350249957704E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.5142158234920399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.50775711061140105},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6682738333033899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":5.6773969821747499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6322849844029399},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.0855988207369E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9080569540014E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.7213432370428001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":7.3345983860963901E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.4068185286491899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-3.27864587514336E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":8.9960281973894103E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":4.4703944655938299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.1897444219477299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.182196509753662},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.6842413458975199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.0299171879105804E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2685914755960404E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.41285244142556E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.31336442158325E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.04829506032006E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.38365706977439901},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0413218427988602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3010562961238903E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6715004363840498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2216276840783903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A80,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C82","v":"080SMPL.d"},{"$f":"SQ!F82","v":"Reagent blank12"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A81,FQ!$AE:$AE,0)-1,0)","v":-2.3E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A81,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A81,FQ!$AE:$AE,0)-1,0)","v":-1.575},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A81,FQ!$AE:$AE,0)-1,0)","v":1.0169999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.31105490555389198},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-8.6795675984169499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-238.15488465138699},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3655136948542199},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-130.63977677222101},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":2.9596935662088302},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":0.61880390465107404},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.1563696261187594E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5894143955219199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.2145320678656102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.33128896609425E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.38784644281178798},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9790813882741904E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.1112522309077705E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":6.6394881796540501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.41321354396378202},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":5.5325121118085302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.55428402087031E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2817095654948699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-12.034275836854899},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":8.3984738394850007E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3451549310671404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.49713108103763E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3377964126229598E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":5.8822470145388998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-3.11274990253993E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.8225890004373801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.19842782172648E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.38466160565230201},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4590738297188599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-4.05740819868606E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3263944011164199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.41095812534857E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.0772250730527E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0369762958314399E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2060535787936897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":5.36527597648513E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8792304851299302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6340855055337403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1366219175852597E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A81,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C83","v":"081SMPL.d"},{"$f":"SQ!F83","v":"QC1-3"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A82,FQ!$AE:$AE,0)-1,0)","v":9.7309999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A82,FQ!$AE:$AE,0)-1,0)","v":0.23300000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A82,FQ!$AE:$AE,0)-1,0)","v":22.873000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A82,FQ!$AE:$AE,0)-1,0)","v":10.724},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":197.81832298116299},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":6.3909094061313496},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-93.138838817223601},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":480.43485257375602},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":462.919005812107},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":576.77539152750001},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":495.46042329935898},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6667022342811501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.52727764224069595},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.52375590459027599},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.46748201089008901},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":21.2276529828763},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.245412345030427},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.1093954834330699},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.1467899294425701},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":51.364981848890203},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.55212452854668503},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.9678021940767901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.7210905669524599},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":9.30825986968709},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3503011528544597},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":24.964970910000801},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":4.1688376086970097E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.31243665713158E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.1065925910642501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9122716261944201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":0.91214823426579605},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.234491333152404},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":0.255479950718906},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.2599717969611802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.2909141247443501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1808831343047603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.52629006440069603},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.4451946963097102},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2573532966646095E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":3.4937499364128401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2569510132956899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2042718878817003E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":7.2802878712691001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6353808023005298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3011342913611803E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.23708614773220299},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.3893493702582398},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.27422665197048202},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.20509947686583099},{"$f":"OFFSET(SQ!$A$1,MATCH($A82,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.7516913150104402E-2}]
[{"$f":"SQ!C84","v":"082SMPL.d"},{"$f":"SQ!F84","v":"QC2-3"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A83,FQ!$AE:$AE,0)-1,0)","v":-1.2999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A83,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A83,FQ!$AE:$AE,0)-1,0)","v":-0.53100000000000003},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A83,FQ!$AE:$AE,0)-1,0)","v":0.82499999999999996},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.628812112409744},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.341494073482354},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":339.31771029999697},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.3566869356857798},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":436.21890683952103},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":11.415445687568001},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":64.719726488485193},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.7694081948412803},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8459491928068E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":4.7845010901305003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.8397001297722801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":2.2989123845781299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":9.1088650283184498E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.12051383204180401},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":3.36700815424344E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":9.6521990490367401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.6025377095565099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.59688824504159399},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-9.7031292257478601},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":2.4546343304715602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":4.9944486697752399},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.7001291034620798},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.66721199691865796},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.58295027693464896},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0711051121000101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.0994816319977301},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.50079650603003201},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":0.42677206738677897},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.24915296839933801},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-4.39216016839656E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-4.53847974537492E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.40638711024189},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.18627872046952E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.430540626568853},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":9.26817753715512},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.50778669339083604},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.6551144174156999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.01972863267527E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3847104987963601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":2.18659355341938E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A83,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":3.98774555988047E-5}]
[{"$f":"SQ!C85","v":"083SMPL.d"},{"$f":"SQ!F85","v":"QC3-3"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A84,FQ!$AE:$AE,0)-1,0)","v":-1.7999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A84,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A84,FQ!$AE:$AE,0)-1,0)","v":-1.3280000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A84,FQ!$AE:$AE,0)-1,0)","v":0.187},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.62340573953096201},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.36347979230731903},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-251.597045216048},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.28595752219104},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":180.028757840168},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-6.6895039659280302},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":3.51644514250034},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1980701647880397E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":8.6074580040670901E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.3534530144258098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.9848683556103598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.13543235247594301},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.6897499830455801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.9916176200969601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.9994787516733302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.16745977183076799},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3606318026776199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.3658815064521899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.804094104184E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2618341914710797E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":326.15001875227199},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.6824203017732998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3070494203753902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":6.4969009662943597E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.0936337657803798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.1186342610217E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.7849817391527401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.61454858020146E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.1917526009620903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0638356180765999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":29.5926558903009},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5507139584898201E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0197831498678498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5166854937430398E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.9902623941116198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.83595006320479E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.5487052654667099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.51043310887365201},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.116733141337615},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":5.5147615813231902E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.90651144510853E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0716033705805298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9542216127766498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A84,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.21425400475665E-4}]
[{"$f":"SQ!C86","v":"084SMPL.d"},{"$f":"SQ!F86","v":"NMS 59"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A85,FQ!$AE:$AE,0)-1,0)","v":49.865000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A85,FQ!$AE:$AE,0)-1,0)","v":1.2090000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A85,FQ!$AE:$AE,0)-1,0)","v":115.532},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A85,FQ!$AE:$AE,0)-1,0)","v":47.195},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":1004.5734810858301},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":41.632360625809198},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-219.24541445748699},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2595.2052213636498},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-465.42507930091199},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7927195578988799},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":399.64115920755899},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.3507534096621798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.4037355889184702},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.3350088963965998},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.3693860213934701},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.20513476884495899},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.2414898028394701},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.14598583436295401},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":25.073917137869799},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.40760645897833497},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.4811198997189301},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":4.2748775330670702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":12.430366753469199},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":49.603752087061402},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7512622656319099},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":126.26736918844},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":25.371474096690299},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.6344469648963799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":7.56085770695892E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":8.2501528838539603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":5.7132412580748705E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":4.8613857469534301},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":3.4797219931425299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":1.2789665874485501},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.2945541139963201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-8.6861684250294697E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":5.5412016199912397},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":8.9737674675670204E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":12.3769733082837},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":-8.2608097325087204E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":1.92850580039275E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":9.7875151819251096E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.11031997909634E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":6.2627869607840604E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.21318762539885899},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":8.5592187246599297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.18186362627246},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":12.1697687497121},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":1.36899841948917},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":1.26643465034848},{"$f":"OFFSET(SQ!$A$1,MATCH($A85,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":0.25813996292958602}]
[{"$f":"SQ!C87","v":"085SMPL.d"},{"$f":"SQ!F87","v":"Reagent blank13"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A86,FQ!$AE:$AE,0)-1,0)","v":1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A86,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A86,FQ!$AE:$AE,0)-1,0)","v":2.9489999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A86,FQ!$AE:$AE,0)-1,0)","v":0.21299999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.94572442767334097},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.147789626637396},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-270.946060821836},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.47498270286028799},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-49.561347833383302},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":5.0855182651429098},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-2.26852157376269},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.17290500288829E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.5923479569075701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.8401416074890199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.8330357175563297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.14348748784103901},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.51266882307546E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.1351093961480699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":4.26751905571503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":8.0968013333974909E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0129508905742796E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5295148438006302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":1.9739932245686699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.3692789567180701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-9.0985053910223694},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.0733804509421601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.03510354972638E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.9864467846848602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":6.2928250150319205E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3195296235691799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":6.8040398423379903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.0218734915267203E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":2.1396539502241802E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.6220178766382899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-5.4135995607583399E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.43736875747727399},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4440163574471996E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1088181232958299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":1.7230461019003101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7664650751074298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0772601740355997E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.02585313200064E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":4.1856149873341603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1028595681384701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.1869374592051396E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.8669022685369897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.1169228287494599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":4.78501302276036E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A86,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.3759912154841299E-4}]
[{"$f":"SQ!C88","v":"086SMPL.d"},{"$f":"SQ!F88","v":"Blank control5"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A87,FQ!$AE:$AE,0)-1,0)","v":-4.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A87,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A87,FQ!$AE:$AE,0)-1,0)","v":-8.2000000000000003E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A87,FQ!$AE:$AE,0)-1,0)","v":0.19700000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.87619866229130405},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.91212575218985803},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":1637.4178708966799},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":8.2297101347621098},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":1791.1895872206701},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":107.85756440185099},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":24.313789336799001},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.4850906347703101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.4191173547548996E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.27780916360322999},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.139857891637583},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":1.5409540970059701},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":6.2746054985686897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.9090364242712199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":9.5804681126961597E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.27342486087787998},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":1.6661352361936999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.66515168885807E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":3.6934386431786499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":4.3458313119352603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":20.100763185252799},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":9.6713139170295101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.8852552854408601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.19446218935923E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":7.1059286784700398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":3.5116188924587399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":3.0893933932825999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":3.8385747039204401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.9864251613338798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":1.0355091333684501},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.5104610780678199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":8.8000029288708894E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":5.4439859377283697E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":3.53759266382376E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0535021038907101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.31904884281034901},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.11958508161315E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":5.82797215220839E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.4040968648245997E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":5.4557597140578798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":2.8360009986096901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A87,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C89","v":"087SMPL.d"},{"$f":"SQ!F89","v":"19286424-001"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A88,FQ!$AE:$AE,0)-1,0)","v":-1.0999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A88,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A88,FQ!$AE:$AE,0)-1,0)","v":-0.48799999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A88,FQ!$AE:$AE,0)-1,0)","v":0.316},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.11499585145772},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.232909265368933},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":112.964426194016},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3270832943264801},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":340.39964676246802},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":13.8350140730922},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-0.35525727598773599},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.7145065258997303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.1568189757301599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.5022185815487001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.2860995256321798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.15297831166977099},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-4.23786899548978E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.5209851519857299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-9.4952334245034103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":52.689281508778798},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.26937237736702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":6.57298137187271E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":7.7179344272655503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":6.8324193900649394E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-0.46771231647584},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0569403158917E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":6.4148297545301904E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9857981997355702E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6237945974462302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.0142846147415401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8380753999235299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-7.9413556540595693E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.19195628735153E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.48436263555099301},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2323668903842799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8176489012563897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":7.5853391110017494E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.2171980588944798E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.6829441643680099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.4657997355874E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.91899075670585E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8363136411364002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3123206798936599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0477332341054101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.560771708229E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7242949835092401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.19094769355762E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A88,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.3196097545209995E-5}]
[{"$f":"SQ!C90","v":"088SMPL.d"},{"$f":"SQ!F90","v":"19286424-002"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A89,FQ!$AE:$AE,0)-1,0)","v":-1.2999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A89,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A89,FQ!$AE:$AE,0)-1,0)","v":-0.61099999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A89,FQ!$AE:$AE,0)-1,0)","v":0.193},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.101378451448013},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.63558191860037405},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":143.63425941518099},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6156514883370399},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":252.99043218767699},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":15.354846649750399},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":4.3143575534070902},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.9618809968932601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-3.69698319240292E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.6736670336472501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":7.6732308623999601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":7.65512189554753E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.56065459277985E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.70259932179154E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0426793833779298E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":57.256171544946298},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.39808047438395E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":8.9199798906808708E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0831157470965998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":0.48949526818019601},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":4.5365160599544603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":6.8250897079884507E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.43164357802159E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4095302089607099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8741769510122301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2478222716295499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.4239741256570401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1339529947328498E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.1333269501246601E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.431948090248576},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4374024477927E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7908590152390797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2527421358099501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.92469645453475E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.11449777587047499},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.57466155197401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.47237026465035E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0629326967945704E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2025287954616198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9132887439822999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A89,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C91","v":"089SMPL.d"},{"$f":"SQ!F91","v":"19286424-003"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A90,FQ!$AE:$AE,0)-1,0)","v":-1.6E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A90,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A90,FQ!$AE:$AE,0)-1,0)","v":-0.746},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A90,FQ!$AE:$AE,0)-1,0)","v":0.13500000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.12977723952023101},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":141.532248890402},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4694514026023899},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":244.561876790387},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":2.9528492404728399},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5880458954619801},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":6.0678459678491104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.4016353619010899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1426611467152404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.4948446479090199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":6.3023229494429497E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":6.1842335630566899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.1973285855846499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.12033503441906E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":59.723492183071798},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.26425594596237E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4769452287546294E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.07501641696592E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5788707604231798},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.34327965899304E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.28696102895618E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7434703082958301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6171279247009499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1865693017547E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.51660139503286E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.4900512280168699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9872721292815299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.48086612462051E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.55625816857838295},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2323775624149299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8132670354615404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.27649184036425E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.13626057400623E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.3099073291900101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.12661445175602101},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.31233726550345E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0477531215263101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.67837241481194E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3736083503913298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0776643615659798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A90,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.1597630354344902E-5}]
[{"$f":"SQ!C92","v":"090SMPL.d"},{"$f":"SQ!F92","v":"19286424-004"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A91,FQ!$AE:$AE,0)-1,0)","v":-1.7999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A91,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A91,FQ!$AE:$AE,0)-1,0)","v":-0.85799999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A91,FQ!$AE:$AE,0)-1,0)","v":0.11899999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.41607272702869802},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.843699009876686},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":106.29691087435999},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2906232317461201},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":444.99434730808599},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":16.148937880267699},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":3.6708252872743099},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.4606417523409403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.06205697703569E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":5.32857425041017E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.1198339391764002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.262987133703456},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0333412841291401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.7917081321837399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7981017833861801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":56.761338036477298},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-8.5698198295035104E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.9346658963756399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":4.0820861896633696E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.2946406202412201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":1.0650504340212701},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.18192112259068E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-9.6855596247497795E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4427413040941801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3459254074252E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1781758106908996E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6362469691052699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.3583517005669605E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-8.12668144680527E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.5033624836059199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.63834468199051697},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2447935034213801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5806331276084001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.93883919546846E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.4375848834157501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.114214714876695},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6445509106595299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2132515089746098E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0358924234250599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3347360591524404E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2921438346455499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A91,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.3091757218143797E-5}]
[{"$f":"SQ!C93","v":"091SMPL.d"},{"$f":"SQ!F93","v":"19286424-005"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A92,FQ!$AE:$AE,0)-1,0)","v":-1.6E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A92,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A92,FQ!$AE:$AE,0)-1,0)","v":-0.89400000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A92,FQ!$AE:$AE,0)-1,0)","v":0.38200000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.398188018513979},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.55668355742216402},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":141.21508749780801},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.45710615819773},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":71.023608233560495},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":20.746761604381501},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":6.5764943349017697},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":6.0072000509229298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.20112208204914E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.87402217306861E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.3891593622340002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.175307495437032},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0569152283892199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.6602432452371898E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.0212068560227703E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":51.851422039889698},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6960836267530403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2421242179551001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":4.8954622396369399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-0.25573375557355099},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2633337853126803E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-9.6716220358399503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.12276152712819E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7767722612142901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1935570527720804E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.82851144000909E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7024319902994101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.5600541651643401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.76590685553673099},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8568252094687901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7267347225851703E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.3275518885726001E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3942043373982299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.4667649176246E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.3678602464260196E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.14208371237432299},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0552858069553401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1676406524622904E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2013080638443299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.44220825581518E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A92,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C94","v":"092SMPL.d"},{"$f":"SQ!F94","v":"19286424-006"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A93,FQ!$AE:$AE,0)-1,0)","v":-1.7999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A93,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A93,FQ!$AE:$AE,0)-1,0)","v":-1.0069999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A93,FQ!$AE:$AE,0)-1,0)","v":0.125},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":6.9500092405132802E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.213004749564632},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":89.6049469134596},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":3.1907994904518602},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":470.40682058822898},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":12.899783748150201},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":0.13142791518696501},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.02403723754362E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.42408669001236E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.6855012797463499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.6421586631579198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.41701996706279798},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9853064187544499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.32994505853800099},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.169192505652733},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":56.884917027458201},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1185369272718E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.1622382384070202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.8226688169287598E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-4.46099007527389},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8253836671989E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-8.0842609813342806E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.2886740427191801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-5.8509167542042395E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4932972679640899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9896335250851702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1999569465367098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.60179524190696E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.68834090895709599},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.84913289832703E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5758822111163196E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.9314184217407199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":6.0421543596519397E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.169234295893369},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1561099322557401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7574541570272201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.21932784986812E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6345028818968103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3953160921326801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A93,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C95","v":"093SMPL.d"},{"$f":"SQ!F95","v":"19286424-007"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A94,FQ!$AE:$AE,0)-1,0)","v":-1.7999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A94,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A94,FQ!$AE:$AE,0)-1,0)","v":-1.1220000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A94,FQ!$AE:$AE,0)-1,0)","v":9.2999999999999999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.55503872355013495},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":9.0863390604191102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":109.183683192537},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.9977369538224004},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-25.958976005226798},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":18.912750230140201},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":4.6195573221071697},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.2278548292052398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.3275567674456303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.00951137369176E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.78437835909212E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":2.24643394801218E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2026799659497699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.32819925733242E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-5.28608097294585E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":50.211688930238203},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1953841227122396E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-6.0835738251059698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":3.2727730714858902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":0.72639477148988596},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6348773869319798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.42050575637515E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2598123122222801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3425638245980701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1627440077983001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0310152656825099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":5.6196975854247097E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6862187583302201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.5099471730034599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.76205838729831299},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8474965920112899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7343014745678701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1579191023369699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.99257718484664E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.16153652254162701},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4378031791754198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.21144711590118E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7821614041324897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.49710569139064E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A94,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C96","v":"094SMPL.d"},{"$f":"SQ!F96","v":"19286424-008"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A95,FQ!$AE:$AE,0)-1,0)","v":-1.7999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A95,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A95,FQ!$AE:$AE,0)-1,0)","v":-1.0720000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A95,FQ!$AE:$AE,0)-1,0)","v":0.218},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.38962642121885499},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.24565103128416199},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":116.142624602943},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":2.2164200210383602},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":592.33189736793304},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":14.5706016190589},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":1.21071396462454},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.5716469920202797E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.5620870397969099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":-3.88972175921326E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.8813243297412598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.110842927754086},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6698266983526198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.7523662083417003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6981772893254601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":47.549788172596998},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8054094716993599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1920747202801303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7283537394026901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.04170466968248E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.03057816837743},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5968908790201997E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7198566672288899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.2693222800870204E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0195436958682098E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5336698154769998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9810157703569999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6833558268181099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":6.2150810442962803E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.55488506170901997},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0832111356764399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1250122558604299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.2821144608408899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8556989695787602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.12936742160919E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":2.1737841292329501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.17740812935547601},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4605195728048999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.23286207288658E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7587234810265901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0216566002337397E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4615965736673798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A95,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C97","v":"095SMPL.d"},{"$f":"SQ!F97","v":"19286424-009"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A96,FQ!$AE:$AE,0)-1,0)","v":-1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A96,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A96,FQ!$AE:$AE,0)-1,0)","v":-1.0069999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A96,FQ!$AE:$AE,0)-1,0)","v":0.14000000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.34896931149115601},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":66.747631660152805},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4188765958207297},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":271.94564953147602},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":15.0327143877544},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":4.8354488772082096},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-1.17121756409924E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":8.8376274236588704E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.3491264962729699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.3536996617385503E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.15748513197682401},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.4289701223735201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.077018044692E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7075486019014304E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":58.330366461816702},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1137490365825303E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.03851437682934E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0435180693809899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3254886262471799},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2703450030859295E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1583676521545904E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":4.7213489916372196E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7346968460542797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.84645119871276E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.17480532325003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.2085949551079105E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2484734268198499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1020078115338599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.44628789341641001},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4192346485162701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7599402363653603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-3.82134778906494E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.6292471270417997E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.15987452376101},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9412723085362502E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6120080999644903E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5021246582514498E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.01226790616987E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5252185938097398E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A96,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.2619310782601998E-5}]
[{"$f":"SQ!C98","v":"096SMPL.d"},{"$f":"SQ!F98","v":"19286424-010"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A97,FQ!$AE:$AE,0)-1,0)","v":-1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A97,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A97,FQ!$AE:$AE,0)-1,0)","v":-1.0980000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A97,FQ!$AE:$AE,0)-1,0)","v":0.124},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.182379813364907},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":120.023029508836},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.64552927024242},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":536.83169475176896},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":17.805310490846299},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":3.2923611313080201},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5363517245603199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":7.9627781403368292E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.55280407484226E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.3631905040949699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.164475764021427},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7219768163612499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.7551960513912804E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1369383247547601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":60.6177812647473},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-8.2498648970415998E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6862526895532299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-1.89160247599678},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.71492189469556E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-7.9245087568573895E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.0980688329104099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3539923364308798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":3.9465630328992601E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-8.0459325412828801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.89462308035716E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.6229106769906897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.76790990986228E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.02145983873297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.78606326453042796},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.28871331871296E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7272016648107398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.5792598488680699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":4.3900753443078998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1466247038513101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2116733271350201E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.17937442343872601},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3784484158616202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2102338516694599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5734591803161502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8549457812457897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5870285224132698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A97,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C99","v":"097SMPL.d"},{"$f":"SQ!F99","v":"19286424-011"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A98,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A98,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A98,FQ!$AE:$AE,0)-1,0)","v":-1.133},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A98,FQ!$AE:$AE,0)-1,0)","v":0.108},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.54443931149138702},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-2.40475254697445E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":101.30666937318099},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.1184995165073299},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":339.415271408867},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":6.4562709685143496},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-0.10054277274019099},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.9256056715541099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.19084796945957E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.5857898431718903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.6942268081687699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.15179187201196001},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8133388313523998E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.65513356581418E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-9.6761120671011395E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":51.319098943439997},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3787900891108902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1279047323278999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6738978893274701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":3.9203780544130899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9493934185892998},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8360693714192699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5729027064768902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.0337589421650501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7183334707881202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1749755366610398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1648895588655199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.8407494037985201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9438260270990401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.60453729145942E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.78734265864626096},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0692389190488502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.70673709253807E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.34341298617075E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-2.21619299663652E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.16838388589894E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.207516667245323},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3255286645130699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2134003805127797E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4934710250776602E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6338752396586202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5762222081979699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A98,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C100","v":"098SMPL.d"},{"$f":"SQ!F100","v":"19286424-012"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A99,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A99,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A99,FQ!$AE:$AE,0)-1,0)","v":-1.077},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A99,FQ!$AE:$AE,0)-1,0)","v":0.20100000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.23569023592000299},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":70.425805054992196},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0515919578806399},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":292.72563461401398},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":18.174079045692999},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-0.109139369938706},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.5823129150755198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5639639334433699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.2746160769816299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.02328932043352E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.23799537597873899},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.3506349199791299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.3863566464275701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":4.4376861718892898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":68.195409545047696},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-7.0363225267994003E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2178270557066403E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.5949238272744202E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8505330830877802},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1162224719527004E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.57950354888515E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.8698954468399499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5922484634989399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8499391421690501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5702151149098001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0429979418488601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.54450569815945E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.74283970483998896},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4264217147102901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.34644167040297E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.54407535323705E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.22008244450563E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":9.4326943964631103E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.1300784046316101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.200455369764658},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3392194528117499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4120620873807997E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2718390365180901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7214746044793596E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6163240925194699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A99,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C101","v":"099SMPL.d"},{"$f":"SQ!F101","v":"19286424-013"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A100,FQ!$AE:$AE,0)-1,0)","v":-1.9E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A100,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A100,FQ!$AE:$AE,0)-1,0)","v":-1.0980000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A100,FQ!$AE:$AE,0)-1,0)","v":0.252},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.37212943367554702},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":6.23954844299629E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":103.39914517572301},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8769921385627004},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":464.16743505983601},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":20.872363626989902},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":7.4936719118016297},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":8.1687245021134003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.5763151828133801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.3704874277666698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.10961894394613E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":5.4356954680691003E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":6.5105728042999403E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":7.1797467257777595E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0538490871706101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":54.988360971931002},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5646820793255101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9984574790696004E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9874796774363102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-5.5128168554714101},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4919961410820899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.26667032581077E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.0261283076962699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5988765823299498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8511910320128901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0961808096018698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.02762406262417E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0445021816893798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":9.4889031865033602E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.62167976998439101},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-7.7927840985483399E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7772115816638898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8350838398908197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.21635347189292E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.18991735718693301},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5075994453163501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.22513785697986E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.0446575492652E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5762839621025899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6937773555841302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A100,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C102","v":"100SMPL.d"},{"$f":"SQ!F102","v":"19286424-014"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A101,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A101,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A101,FQ!$AE:$AE,0)-1,0)","v":-1.2250000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A101,FQ!$AE:$AE,0)-1,0)","v":0.14799999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.71431680789830898},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.18988686713866701},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":47.578054212232701},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.3198457299978603},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":475.42125750757799},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":11.7073112093908},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-0.31457743384321502},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":3.6185461267816903E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":6.5392186482708097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.9901453496161102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.0289660723890601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.14185459962062399},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":7.5233250274761305E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.50592893234277E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3479179752680199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":49.851859640658397},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6283354445825296E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8739771701807496E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.5789461132830801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.9003839222464002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1878032200884898},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.00008112973563E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.120572359478E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.06735615225877E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.40103182769854E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8610131831064401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.83036844981089E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0029450315591798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":3.5890986638624201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.67046852863061601},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.64511426513162E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0313127902258802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.3399987218573199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.22222871448558801},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7302502489528899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.62264447474721E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9037423816392399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1626841687148596E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7056662672150001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A101,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C103","v":"101SMPL.d"},{"$f":"SQ!F103","v":"19286424-015"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A102,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A102,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A102,FQ!$AE:$AE,0)-1,0)","v":-1.1439999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A102,FQ!$AE:$AE,0)-1,0)","v":0.129},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":8.0347275415264499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.27198765499488498},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-14.458752971989},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4540135163921399},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":293.39663620749599},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":20.478960252670898},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":6.9565591521732397},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.2930195109554605E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":9.5814069993207501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.8149310286089401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.5194970641849799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-7.7783337895785995E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-9.8421088215416711E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.2872941596224401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.78683177838318E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":48.563367548096998},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":5.0978960059561698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-7.3233403798387102E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.1526839438452201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.2737273872795402E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7845785702508898},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-7.1478579038141296E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-9.7956951954687108E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.8614568815502896E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1431398446613701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.85937610531572E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2790837054536699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":8.9006542782140504E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.72414123325866E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4706491026977402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.79509800934793096},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6555980136658201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-7.8011429138877498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.2830825001867499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.70634694973802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2680516548311505E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.19151927186604001},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6800213884530799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2326096359550504E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1104650175964601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3833785467669298E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6662776282784E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A102,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C104","v":"102SMPL.d"},{"$f":"SQ!F104","v":"19286424-016"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A103,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A103,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A103,FQ!$AE:$AE,0)-1,0)","v":-1.17},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A103,FQ!$AE:$AE,0)-1,0)","v":1.7809999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.82313952752206698},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":11.559543971016399},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.9679471186996196},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":552.890632814112},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":20.4249737873347},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":3.4077600187094701},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1112240702259502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.24235885108651E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.28202212937124E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.59802943086204E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.193879550381796},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.31953844374611E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":6.6155624777741798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9870743050956199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":53.996588955522803},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2557845312457501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0723911416143698E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.1825913746506301E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.3926261630628398E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.17835615993378},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-8.8306545104179499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4331577721114701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-8.3647436335355203E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2126358271890197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.8875024397430501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-9.2619930259847801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6667654082294899E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.5352815344251797E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5931586239873602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.4508841312516499E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.82109044910278095},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-8.1616894402550299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8693686904423501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.8290763007107002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.23655285296093101},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0850423149454601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6479957200834299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-9.07282150233233E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1688485659711798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.65041033053534E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A103,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.1358020505684701E-5}]
[{"$f":"SQ!C105","v":"103SMPL.d"},{"$f":"SQ!F105","v":"19286424-017"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A104,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A104,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A104,FQ!$AE:$AE,0)-1,0)","v":-1.2509999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A104,FQ!$AE:$AE,0)-1,0)","v":0.13},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.26081497225153299},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.117441295178854},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-110.883822174061},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.05964358407422},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":273.86190108768898},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":8.5964629156523493},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.85711555427518},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-7.6696782285873601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.4343152420381E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":-1.66892042140238E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.1539259577270601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-0.11800280861099401},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-7.0566637499629896E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.5669644246482096E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.15324825596189701},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":52.512200361949397},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1412603516288301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6508641739782001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8132281650537899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":4.4789662364234303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8112674317786297},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.80201278539448E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.15709613520593E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":-4.4102684387111103E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8767224696526596E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5677953661096199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4722623532134499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.5044371009892698E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.08203155880632E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.0898195638691301E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.82677904027365301},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6857790149047501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9820237566726901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":7.3415866994350703E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6118254635912599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6007700620528599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.02294391733556E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.20798298688489},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5895292205992598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.95625423359402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.53185458030118E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8273264521709498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.70049915842831E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A104,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C106","v":"104SMPL.d"},{"$f":"SQ!F106","v":"19286424-018"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A105,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A105,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A105,FQ!$AE:$AE,0)-1,0)","v":-1.234},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A105,FQ!$AE:$AE,0)-1,0)","v":0.17599999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.57141011334633196},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.43618879160836699},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-91.200517992871298},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":0.35864636306768799},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":148.596131781704},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1896108456938901},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":0.36434898771517199},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.3225420254285602E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.55715497448549E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.6778039640963799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.9169158983041597E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":4.0289267729745303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6718014997922899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.6061867531475201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.2420990202892E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":63.622516002870697},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2411348389531203E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7297430642964499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":0.21900016453075499},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.21879752735321E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4509523628892801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.29302356777118E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7792509787917001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-3.19843258065637E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5979195769019001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.24400175857677E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.9752856153037302E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.64302178308112101},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0999829317185702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8815160564610501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.4616870476011701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-3.9180475446004098E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.2466499688318501E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.22379417228617199},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7502222031401201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2235778554032002E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5521227847469302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7292132451121699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A105,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C107","v":"105SMPL.d"},{"$f":"SQ!F107","v":"19286424-019"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A106,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A106,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A106,FQ!$AE:$AE,0)-1,0)","v":-1.1779999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A106,FQ!$AE:$AE,0)-1,0)","v":0.25600000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.259542448103806},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.29310816927717098},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-96.598079695381998},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.07883294122578},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":171.09970496628301},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":10.471402438115501},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":3.46353469409026},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.37540271685922E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":9.4631690437900796E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":1.74600623921944E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.88893585681438E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.22665226871357499},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-8.6849446685382603E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.5013953104655199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.34099004733845101},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":55.736548584640602},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1454080684648498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5664202547361001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8012149300871999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1927704847854299},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1281896976926203E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0043213485379499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":9.0130860023281004E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5007984423171797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.85443405319443E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":-2.55774178759405E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.79833162863647E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8325185591920401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":2.2531432192711199E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.75423087193166405},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.88536114030873E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8522824724979699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-7.9640858511798696E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.16030335574762E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.22579674332342201},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7445934276740299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.52922867019341E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0563862560460901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7447044288670799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A106,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.22070760883489E-4}]
[{"$f":"SQ!C108","v":"106SMPL.d"},{"$f":"SQ!F108","v":"19286424-020"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A107,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A107,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A107,FQ!$AE:$AE,0)-1,0)","v":-1.149},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A107,FQ!$AE:$AE,0)-1,0)","v":0.222},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-9.9373967863818996E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.196926736469797},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-95.651381558545594},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":6.1546208845900701},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":132.453280445258},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":12.710624130235299},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.63532184488276},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1994126943177E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0591685997233601E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.2888968381674198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.8313883882211403E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.13807078350568699},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.01527136627371E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":5.6410906442159299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0154735320495099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":48.596385576886199},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-1.13863860899058E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.58316460801418E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":8.2124514457458705E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-4.1551177621226199},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2315807314655298E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0100557469619201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.6143962850930101E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.0692612879042501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3917442650977501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8019028296204101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":9.6685531943848497E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.0527461805448295E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.66507916236173903},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.27559958664691E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6185239993008E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-5.7396300782179502E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.220116514417839},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.69628151305405E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6812910340045401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":4.8558947121834202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0375370771726397E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6612837342144799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A107,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C109","v":"107SMPL.d"},{"$f":"SQ!F109","v":"Acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A108,FQ!$AE:$AE,0)-1,0)","v":4.1340000000000003},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A108,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A108,FQ!$AE:$AE,0)-1,0)","v":129.107},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A108,FQ!$AE:$AE,0)-1,0)","v":797.74199999999996},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.125"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":78.873227671876407},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":32239.135066536201},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":309.85112203255801},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":49701.733010531098},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":3554.4514209782401},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":497.38282189667501},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.3427659246559402},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.5674224440792202},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":19.3440628610795},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":16.398148652150901},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":171.23243661880599},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.60418423270311095},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":11.436861817713901},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":6.2745803399059703},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1111.4960967223401},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.38479647449217103},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":12.892123161872901},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":0.33096702769039998},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":29.372961663957899},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":13666.8808789872},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":18.395133900409402},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":27.626640535814101},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.20375663789780399},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.114020189650839},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.80064668664092598},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.035"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.02502522309151},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":7.50144952145852},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":5.0998637181338097E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.027"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":49.819860596779797},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.5522872784691799},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<1.520"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":423.59109970234698},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.53671076552107999},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.5114022531043001},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.7842285075750701},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.33441108564335},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":0.114020225916958},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":2.90872682562777E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.4625614343464305E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.004"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.010"},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":238.25809978112699},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":19.8972867321481},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":7.4524548694425499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":0.88590358794816304},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.24565183528776199},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.13776172704537401},{"$f":"OFFSET(SQ!$A$1,MATCH($A108,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":2.6325027074182598E-2}]
[{"$f":"SQ!C110","v":"108SMPL.d"},{"$f":"SQ!F110","v":"Base rinse"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A109,FQ!$AE:$AE,0)-1,0)","v":2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A109,FQ!$AE:$AE,0)-1,0)","v":2E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A109,FQ!$AE:$AE,0)-1,0)","v":0.67100000000000004},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A109,FQ!$AE:$AE,0)-1,0)","v":6.9790000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":13.6125149780603},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":14.170733345065701},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":20740.398861424899},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":351.42729403698098},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":33695.902434784803},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":1747.31160203488},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":350.19930777489202},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.82719578037422},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.20149450795315499},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.9219276305573501},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.1868524290340599},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":20.356482660945201},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":9.4633654163487693E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.98376941535905205},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.5640191646711299},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":44.6558469987311},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.10342775804103101},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":7.3046874120186303E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.007"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":0.77017631181827495},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":211.40484347491301},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2648565645444498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":3.03248149710445E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.77343205819328E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":0.633346105741735},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":6.5064077751040603E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.933238746627292},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.9662012262031601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.3643518044251696E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":1.4294922828814101},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":4.82384757951764E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.044"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":4.0065902888914096},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":1.1131024848152499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":5.4957340314032897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":4.8576445648628903E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":7.7162477173253197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":3.2780600297453999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":4.4061758558541801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":27.711589654231201},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.8927108458065599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":4.8460745773452997E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.2579195965588199},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":2.0944899678836601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.49735640970474099},{"$f":"OFFSET(SQ!$A$1,MATCH($A109,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":7.6112348626790403E-4}]
[{"$f":"SQ!C111","v":"109SMPL.d"},{"$f":"SQ!F111","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A110,FQ!$AE:$AE,0)-1,0)","v":3.7719999999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A110,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A110,FQ!$AE:$AE,0)-1,0)","v":504.24200000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A110,FQ!$AE:$AE,0)-1,0)","v":11536.584000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.169"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":384.75245496786198},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":47446.748437695802},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":196.454837283638},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":68161.840325755998},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":4196.75845523514},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":735.95555068461795},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.7721994827951004},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.4511302117895},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":42.314323225096402},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":37.031261570674801},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":362.58078410746202},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.25931969911463},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":16.507193737211502},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":17.2961861984198},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":105.013320798808},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.5284808649104198},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":3.7889626393722402},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.258"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":10.822386620394401},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":2489.2843854212101},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":14.432068212881701},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.276"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.107405715466559},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.2263352556201701},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.13678383754013},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.092"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":6.0389585506569103},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.032"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.21169968896274899},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.072"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":5.1758826066236701},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":6.83412263217458},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<4.059"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":788.64057173968399},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.89744030967614397},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":0.83243389347624297},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":0.128186434335709},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.107"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.046"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":7.7673734696700095E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.034"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":6.63662584330654E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.010"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":9.0533026055303195E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":84.709560906638202},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":1.81979143450378},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.30026414290019499},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.40720886167472},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.42125792564335102},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.015"},{"$f":"OFFSET(SQ!$A$1,MATCH($A110,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.011"}]
[{"$f":"SQ!C112","v":"110SMPL.d"},{"$f":"SQ!F112","v":"Blank control6"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A111,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A111,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A111,FQ!$AE:$AE,0)-1,0)","v":-1.5980000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A111,FQ!$AE:$AE,0)-1,0)","v":0.27100000000000002},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.71223130746462404},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.32689129981103499},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-41.8212412720275},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-0.30922673109827098},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":732.30555850504095},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":19.5033760134787},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.87529016895731},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.5441030106041103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":4.8360186908304702E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.6616422601480798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.6810740201091902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.584150931303558},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.08212336288338E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":4.3983716691303798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3781362800300898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.445390029633631},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6520987092785503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1804811515542204E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":2.5926220786887202E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6418454877821702},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9282634140443499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.2812173995027E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.8831894543406E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0438626685527501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":3.0677149098487101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.92067947888613E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6955795281847699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.1673246294258901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.75019098027478504},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.45363761520707E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8242577112667497E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3010217901129701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":4.80201879684069E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.0876006461429201E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.37826917102246199},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9565623544152E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.06153294398568E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9209124953830201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.3125898768027896E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7376025538340801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A111,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C113","v":"111SMPL.d"},{"$f":"SQ!F113","v":"Reagent blank14"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A112,FQ!$AE:$AE,0)-1,0)","v":-0.02},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A112,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A112,FQ!$AE:$AE,0)-1,0)","v":-1.5620000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A112,FQ!$AE:$AE,0)-1,0)","v":0.53700000000000003},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.360837614378685},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.145702187890727},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":25.992783987235999},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-1.31460132658789},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":474.936302295953},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":18.051829951319299},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":16.434274336394999},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.8518792850823693E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.1188574338760999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":2.57269561234936E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.4670928621036098E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.35615087033910803},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0936701007335501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.03444973460456E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4697979716712099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.33843574664003201},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-9.7477556305058596E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4774905742316699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.4479003099559402},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8632297940312199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":1.3306634914302801E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.7143893282441899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.9882637142439701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0358106068581201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":3.8489989363285003E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.39220995888395E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8312978062817799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.8655761450075905E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0475723171194201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.22993780324097E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.64752952789874796},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0793386743616399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.74760048143044E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.72752433401794E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.2008364342605098E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4254869425475199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.0087454747653099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.89146881501117E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7256962908081897E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7297456655184099E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A112,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.2095137204739502E-5}]
[{"$f":"SQ!C114","v":"112SMPL.d"},{"$f":"SQ!F114","v":"QC1-4"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A113,FQ!$AE:$AE,0)-1,0)","v":10.589},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A113,FQ!$AE:$AE,0)-1,0)","v":0.25700000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A113,FQ!$AE:$AE,0)-1,0)","v":23.12},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A113,FQ!$AE:$AE,0)-1,0)","v":11.116},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":220.83419766923899},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":7.7281380218462701},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":40.304251592421899},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":563.26944297702096},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":469.24998667406999},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":663.88900653977396},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":504.25098677805897},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":1.4097343294491201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.51412493669433001},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.59189692556585105},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.56447664855525503},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":22.762159644104099},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.26271126377944798},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.0755266091070901},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.4254122590320897},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":55.640468881839098},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.52486893270367696},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":2.8258616336118902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.8784943300837198},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":9.7950276705313293},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-5.46589019812399},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":26.759193709314602},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-9.4573379596278994E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.5682928893836501E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-3.72456594047671E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":7.2338267023185103E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":0.93543577052664895},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.259465966517681},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":0.26096942342815199},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6485849002896002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.0000696736755897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.68895096783876697},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.53514902385946805},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.6964792074085602},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":7.4450364886196599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.1717949149499701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.2507364830352499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":5.4638342976612696E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":-0.120251570707163},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9498594202091302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.25515908519085201},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.6110162742254701},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.29852553525505099},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.224850705060963},{"$f":"OFFSET(SQ!$A$1,MATCH($A113,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.8537421534485802E-2}]
[{"$f":"SQ!C115","v":"113SMPL.d"},{"$f":"SQ!F115","v":"QC2-4"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A114,FQ!$AE:$AE,0)-1,0)","v":-1.2999999999999999E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A114,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A114,FQ!$AE:$AE,0)-1,0)","v":-0.36699999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A114,FQ!$AE:$AE,0)-1,0)","v":0.33600000000000002},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.11168136862635},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.17755750255328001},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":468.09417768486202},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-3.59032542270768},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":552.21313266486902},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":20.495367761667801},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":85.340770309392397},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.9364452809932704},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.3192441446383599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":4.57313104375087E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":2.3044792608691699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.25376969451440101},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.01267312854121E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.64408496691995E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.3785777350174599E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.13936119281144199},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-6.8394536118907799E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.61530124436278399},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-0.94293627548470804},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8651289877156699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":5.1233433591773503},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.6381149764238501},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.66835984768944601},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.544577383656354},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":7.3365068803705304E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":0.995164331159898},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.50254387792545696},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":0.36104432002066},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.76302896678680399},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1892525295235201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-6.7416555531316603E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.43043268086261399},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":5.9967423464862301E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.43408653433532401},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":9.04952796357958},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.51347373806032004},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.6011096517939898E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2298025630835199E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6658882445908701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":3.0862013742802102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A114,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":8.6002598672536398E-5}]
[{"$f":"SQ!C116","v":"114SMPL.d"},{"$f":"SQ!F116","v":"QC3-4"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A115,FQ!$AE:$AE,0)-1,0)","v":-1.4E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A115,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A115,FQ!$AE:$AE,0)-1,0)","v":-1.25},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A115,FQ!$AE:$AE,0)-1,0)","v":0.17100000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.69662567069227699},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.20951111671669301},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":122.998235591191},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.44683643717857},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":515.32485523185403},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":35.684440847987801},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":7.22298976301997},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.0172108503836001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":2.5066264137036401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":5.3945730862118499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":4.8407754608834501E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2451439839742199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2232336310132296E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.7435230289790399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":2.5854719784470098E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.16739655112184601},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-4.7023226223733196E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.53241154689739E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":5.6961056760779401E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":3.1390308205810402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":303.01377820497498},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.2052186782034297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.3960186663748901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.6669977257258201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":9.3268210597138197E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":4.0501856657794801E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":5.14426669630843E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0932705348014499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.8594269360149102E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6992283159671799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-1.6786526964371899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":22.033046859313501},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-6.9528613936351697E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-3.1627141703292498E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":7.7218097833524797E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.8570077417165197E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.09392063565517E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.53109808366759303},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.14439541843804601},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0626638865900901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-5.1483781087286401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.60610136121187E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A115,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C117","v":"115SMPL.d"},{"$f":"SQ!F117","v":"QC1-5"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A116,FQ!$AE:$AE,0)-1,0)","v":9.94},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A116,FQ!$AE:$AE,0)-1,0)","v":0.23300000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A116,FQ!$AE:$AE,0)-1,0)","v":21.565000000000001},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A116,FQ!$AE:$AE,0)-1,0)","v":16.158000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":212.31996034231099},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":9.84818876876618},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":15.937373789489101},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":529.56066733597299},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":781.89936469318502},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":628.06428087334496},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":3897.31839040351},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":0.33445486061326302},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.51137094862321597},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":0.55993851833228003},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":1.23427314279694},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":22.5232006656747},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.24920547567498799},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.3333462360202499},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":5.6946674508700097},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":71.096719361902899},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.51198331185005597},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":4.41917669119253E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.7173624502857399},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":9.6822987374785292},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":0.75402697140602704},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":24.460581049793099},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":2.9446795015191398},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":1.09555064275132E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":7.5439846329458493E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":1.8763274936661301E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.4483175316624499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":0.98074726572857496},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.24208372943072301},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":0.24982458895944901},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.44251224103117E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.6711959254226E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":3.7724873660797198},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.48965976627843999},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.7630014394700502},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.7098950714025701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":6.8226694046453502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":1.5545879656211499E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.9829244899880499E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":1.0044683364904101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":2.9573294559228701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-8.4297409154220192E-6},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":5.8881864984888701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.24463250795716501},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":2.7871755315490701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.24035280612849799},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.5591734445641001},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.28208831840464399},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.23896254163606501},{"$f":"OFFSET(SQ!$A$1,MATCH($A116,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":5.5326373443955697E-2}]
[{"$f":"SQ!C118","v":"116SMPL.d"},{"$f":"SQ!F118","v":"QC2-5"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A117,FQ!$AE:$AE,0)-1,0)","v":-3.0000000000000001E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A117,FQ!$AE:$AE,0)-1,0)","v":1E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A117,FQ!$AE:$AE,0)-1,0)","v":-0.23699999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A117,FQ!$AE:$AE,0)-1,0)","v":1.4850000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.99733178808766398},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":0.22769007186413101},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":539.26888277892601},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-2.74487319950613},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":494.99520476970901},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":15.7727183081162},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":411.19160444916201},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.3837785080682199},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.44107305554966E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":4.1110932056147402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":0.102829490583999},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.55213184607064403},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":4.7760755644806804E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":0.13824467067773599},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":0.47682799214333199},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":2.31036195144249},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4064647116603502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.61359347707978795},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":8.0004149074589206E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.6546415553378799},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2973762524653102E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":5.17333116834976},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":2.1747676640112498E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":2.7120346150136898},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.66146491712212097},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":0.59878729204688397},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":1.3248924098671599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":0.98118408712764604},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":0.49423449624964799},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":0.44496920028984099},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":0.19559925787013399},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.8202357898855601E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":4.3436566998494103E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.5703387698859902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":2.7658468726931701E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":5.9082775035855105E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.0048245820911E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.42441420730547103},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.27468500209649E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.43754034370841299},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":9.1839972011011},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.70581211073970695},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-6.3878665806584007E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.31846878891431E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5127689324286901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":2.4440563038097499E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A117,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.0923105233620202E-4}]
[{"$f":"SQ!C119","v":"117SMPL.d"},{"$f":"SQ!F119","v":"QC3-5"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A118,FQ!$AE:$AE,0)-1,0)","v":-1.4E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A118,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A118,FQ!$AE:$AE,0)-1,0)","v":-1.1679999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A118,FQ!$AE:$AE,0)-1,0)","v":0.14499999999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.178140591470733},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.363336492476386},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":36.629992700653197},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-6.1076430924259597},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":262.517567410521},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":27.516098098323202},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":11.356083658498999},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":4.7326617211941799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":3.49488933878063E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":7.0607302673759698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.7825306157468302E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.47731751135601602},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0969950871041699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.14189769648984E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-8.0643873415087598E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.22513035478612001},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-6.2272829204975096E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":-1.21099918682966E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":2.7070295606891502E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":6.58215951624416E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":279.40181851843801},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.0769708919601399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3726490019743401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":8.7068234428997002E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":7.97374806951422E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":2.0703313758864701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.1629423957364902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.7907465624321699E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":8.1999273026621205E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-2.4021825420287201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":-7.2537937781455802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":25.265847728206602},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.11595514202198E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0984556792576799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":1.6552485994745299E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":6.3676971672915299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8083725671612597E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":3.4058536200436001E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.85435456512576702},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.16838820506556501},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-6.4826693681489195E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.4480290764552399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8260373376823604E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6259664491688001E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A118,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C120","v":"118SMPL.d"},{"$f":"SQ!F120","v":"Reagent blank15"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A119,FQ!$AE:$AE,0)-1,0)","v":-1.7000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A119,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A119,FQ!$AE:$AE,0)-1,0)","v":-1.2929999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A119,FQ!$AE:$AE,0)-1,0)","v":2.7E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.525779980839709},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.28677348600082497},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":7.4934057348120904},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":1.64428377802022},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":682.27031633520096},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":11.5245376105317},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":6.8558293983177396},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":9.6586133402154106E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-3.7360115997066901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":7.3894343886618197E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.1881295569956399E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":4.61543453394559E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.6410291392818599E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.05895306220754E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0324545823637401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":-0.13385486199791699},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2438350040899901E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.52333961527837E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":6.6372300435196697E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2864359224461301},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-7.8298776188195598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.22451708346097E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.2723066673198899E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":4.2627390375399402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":4.1580526506367402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":2.9010399916686099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.2042019461229901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":6.7161775798470897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-3.2974993137340598E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":5.4703967376673504E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":4.8222792383808502},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":5.7660677008923897E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.79893702852453E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":2.1648455276230099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.2978466348114099E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.32923932759406399},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":4.3833829666230902E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5272078658495699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8621981317147401E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.5859879237959101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7337072636902101E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A119,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.55915430007434E-5}]
[{"$f":"SQ!C121","v":"119SMPL.d"},{"$f":"SQ!F121","v":"Acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A120,FQ!$AE:$AE,0)-1,0)","v":6.6070000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A120,FQ!$AE:$AE,0)-1,0)","v":0.27200000000000002},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A120,FQ!$AE:$AE,0)-1,0)","v":214.15799999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A120,FQ!$AE:$AE,0)-1,0)","v":902.46600000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":19.7749441035361},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":48.225524839986001},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":25001.412495394499},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":407.40901174105301},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":37238.885064858703},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":2779.5099487952102},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":493.78212487036501},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.3792523967011396},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.7981733674142899},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":29.027312227514901},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":15.2999860414506},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":251.665682307173},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.89911649326493104},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":17.914379339350099},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":7.1772207134643802},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":1526.94059620024},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":2.7138030292069701},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":16.1568950500752},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":3.4634163854967799},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":"<4.380"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":22878.8084226707},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":25.968082261341898},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":42.062554303675299},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.022"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.03187456733641},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":4.7129893228435504},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.062"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":5.3655163121560898},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":18.1277939893018},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.12836341082524599},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.049"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":73.304644704261506},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":4.11231792795608},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<2.736"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":1008.99199637615},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":0.7254465515595},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":3.9605659707422398},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":2.46260839710766},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":1.2009850119635099},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":0.102619490444171},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.008"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.023"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.99171511032217496},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.007"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.018"},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":1161.72635671177},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":108.21061075748101},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":6.7014661953781404E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":1.29768667309709},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.28217941905273403},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.636134280732221},{"$f":"OFFSET(SQ!$A$1,MATCH($A120,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.007"}]
[{"$f":"SQ!C122","v":"120SMPL.d"},{"$f":"SQ!F122","v":"Base rinse"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A121,FQ!$AE:$AE,0)-1,0)","v":0.09},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A121,FQ!$AE:$AE,0)-1,0)","v":8.9999999999999993E-3},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A121,FQ!$AE:$AE,0)-1,0)","v":2.77},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A121,FQ!$AE:$AE,0)-1,0)","v":12.518000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.106"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":25.964312794621499},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":29723.740470234501},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":405.43768133804099},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":47004.257740765999},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":2809.8231659356302},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":492.94186680306598},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":2.1814685247045902},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":0.38389415416148898},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":6.3369165763770097},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.4012079990165298},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":40.3108881359869},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":0.13813931097009999},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":1.8086306946801201},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":3.0602138429223298},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":59.316013132922699},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":0.261161727715012},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":8.8291170946291395E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":3.5420256241473297E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":0.41142229460402402},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":354.39696149783902},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":1.5317401998098801},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":0.32860636564312701},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":1.1527144484719001},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":0.15653217061176899},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.002"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":3.2329253974596699},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.3240695479596198E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":9.2655758740818106E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.002"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":2.6726158608769599},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":7.6047037324552896E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.086"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":19.673545333310599},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":3.2006316771380702E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":0.116674125377104},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":8.1941393723593894E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":7.5927697480548803E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":1.6550584998278101E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":0.122068838545695},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":1.1175514622247701E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":1.8259566288518999E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":94.640904414653306},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":0.34349062931333402},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":1.1169583618451599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.54476839397784},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":4.0006095047733801E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.35289670357011299},{"$f":"OFFSET(SQ!$A$1,MATCH($A121,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":4.4936649063432197E-3}]
[{"$f":"SQ!C123","v":"121SMPL.d"},{"$f":"SQ!F123","v":"2% Nitric acid rinses "},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A122,FQ!$AE:$AE,0)-1,0)","v":6.9749999999999996},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A122,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A122,FQ!$AE:$AE,0)-1,0)","v":764.67399999999998},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A122,FQ!$AE:$AE,0)-1,0)","v":15568.478999999999},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":6.33916498151219},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":430.45819421440802},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":57144.063334452301},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":49.860171905280701},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":105080.29677254301},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":5654.67188086021},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":963.08902846567605},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":7.23584845945014},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.5404618015226901},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":64.406610731959603},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":52.674595456059002},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":574.54075365938695},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":2.4844259248688401},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":33.736531688533802},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":28.6689117501693},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":205.19748279412801},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":5.9081660585828901},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":0.68785668509395004},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.431"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":18.117131810818599},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":7038.3368062089003},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":48.629647508329697},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.463"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":0.36080209467760799},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":5.6762611934953204},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":3.1208888043565501},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.156"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":21.2876324703814},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.054"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":0.427289208574683},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.122"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":12.2634735231736},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":13.986759609230001},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<6.841"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":2319.2848673006301},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":2.7273490547214099},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.851"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.032"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.180"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.077"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.020"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.057"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":1.3525446339932401},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.017"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.046"},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":130.40019895828101},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":14.785341078690699},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":0.33740658032417198},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":2.8752397889571202},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":0.405454309294632},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":0.13776459180479},{"$f":"OFFSET(SQ!$A$1,MATCH($A122,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":0.118464936171698}]
[{"$f":"SQ!C124","v":"122SMPL.d"},{"$f":"SQ!F124","v":"Reagent Blank"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A123,FQ!$AE:$AE,0)-1,0)","v":-2.1000000000000001E-2},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A123,FQ!$AE:$AE,0)-1,0)","v":0},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A123,FQ!$AE:$AE,0)-1,0)","v":-1.4179999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A123,FQ!$AE:$AE,0)-1,0)","v":0.26300000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":-0.53232594803660305},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":-0.68199996695471898},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":-48.739274433783102},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-4.0770522280926897},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":695.74913289663004},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":11.4802266080818},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":6.2736554238495801},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.8177682061424998E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":-4.26809007898839E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":3.5395983426991601E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":5.5062945942647698E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":0.92806272840971604},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":-8.9752177622360295E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":2.39071087939445E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":6.3676396147995902E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":0.35111521865996198},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0727252335593898E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.0018617148631901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":4.8298130211226701E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":5.52148801155594E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-3.36213544120774},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-3.3750762240088002E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.0786343883989901E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":3.5708207821780403E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":3.62663591510808E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":8.1643241979597998E-5},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":3.7464105899667199E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-2.9077501468818798E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":4.2342063848925402E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.5971499308459799E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":1.7430015764675799E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.52352914400550599},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":-5.6092775882261503E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":-2.0601641479219999E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":8.00647082976802E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":4.6762233225744302E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":0.60717516072896505},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.5930381568996401E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":-4.56010199209381E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.77400432933134E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":-4.8391939442947399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.7736840470864201E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A123,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"}]
[{"$f":"SQ!C125","v":"123SMPL.d"},{"$f":"SQ!F125","v":"NMS 55"},{"$f":"OFFSET(FQ!$Q$1,MATCH(C$1&\"_\"&$A124,FQ!$AE:$AE,0)-1,0)","v":19.997},{"$f":"OFFSET(FQ!$Q$1,MATCH(D$1&\"_\"&$A124,FQ!$AE:$AE,0)-1,0)","v":0.48899999999999999},{"$f":"OFFSET(FQ!$Q$1,MATCH(E$1&\"_\"&$A124,FQ!$AE:$AE,0)-1,0)","v":48.77},{"$f":"OFFSET(FQ!$Q$1,MATCH(F$1&\"_\"&$A124,FQ!$AE:$AE,0)-1,0)","v":21.725000000000001},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(G$1&\"*\",SQ!$1:$1,0)-1)","v":0.24734554317597399},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(H$1&\"*\",SQ!$1:$1,0)-1)","v":17.008416857333401},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(I$1&\"*\",SQ!$1:$1,0)-1)","v":773.46377608940099},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(J$1&\"*\",SQ!$1:$1,0)-1)","v":-0.880364181739272},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(K$1&\"*\",SQ!$1:$1,0)-1)","v":-3.5648087536715098},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(L$1&\"*\",SQ!$1:$1,0)-1)","v":-3.8310707547938301},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(M$1&\"*\",SQ!$1:$1,0)-1)","v":-1.89175177023081},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(N$1&\"*\",SQ!$1:$1,0)-1)","v":5.2899339066982097E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(O$1&\"*\",SQ!$1:$1,0)-1)","v":1.312039372026E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(P$1&\"*\",SQ!$1:$1,0)-1)","v":-9.3933266583033394E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(Q$1&\"*\",SQ!$1:$1,0)-1)","v":3.5160841678584402E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(R$1&\"*\",SQ!$1:$1,0)-1)","v":42.124364213269899},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(S$1&\"*\",SQ!$1:$1,0)-1)","v":1.0814616768156299E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(T$1&\"*\",SQ!$1:$1,0)-1)","v":9.6226661707639494E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(U$1&\"*\",SQ!$1:$1,0)-1)","v":1.4442002293714999E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(V$1&\"*\",SQ!$1:$1,0)-1)","v":100.042324691657},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(W$1&\"*\",SQ!$1:$1,0)-1)","v":-5.9179782547140302E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(X$1&\"*\",SQ!$1:$1,0)-1)","v":1.2075282981889801},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(Y$1&\"*\",SQ!$1:$1,0)-1)","v":8.2714068255606201E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(Z$1&\"*\",SQ!$1:$1,0)-1)","v":-2.34091312379066E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AA$1&\"*\",SQ!$1:$1,0)-1)","v":-2.8074674678865699},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AB$1&\"*\",SQ!$1:$1,0)-1)","v":-5.2508135209663E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AC$1&\"*\",SQ!$1:$1,0)-1)","v":-1.09227758273568E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AD$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AE$1&\"*\",SQ!$1:$1,0)-1)","v":6.0220922026211897E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AF$1&\"*\",SQ!$1:$1,0)-1)","v":-1.1734228371651099E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AG$1&\"*\",SQ!$1:$1,0)-1)","v":1.20245807554756},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AH$1&\"*\",SQ!$1:$1,0)-1)","v":1.0818742701901201},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AI$1&\"*\",SQ!$1:$1,0)-1)","v":-1.4877404504624299E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AJ$1&\"*\",SQ!$1:$1,0)-1)","v":7.8372861117610899E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AK$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AL$1&\"*\",SQ!$1:$1,0)-1)","v":-1.06281452689172E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AM$1&\"*\",SQ!$1:$1,0)-1)","v":8.9005480482239798E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AN$1&\"*\",SQ!$1:$1,0)-1)","v":0.79387765813210198},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AO$1&\"*\",SQ!$1:$1,0)-1)","v":-0.510576531839814},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AP$1&\"*\",SQ!$1:$1,0)-1)","v":-1.9505005725344699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AQ$1&\"*\",SQ!$1:$1,0)-1)","v":2.5409141676499699E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AR$1&\"*\",SQ!$1:$1,0)-1)","v":5.3172963670446497E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AS$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AT$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AU$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AV$1&\"*\",SQ!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AW$1&\"*\",SQ!$1:$1,0)-1)","v":2.1887081470795699E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AX$1&\"*\",SQ!$1:$1,0)-1)","v":8.0354036475785201E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AY$1&\"*\",SQ!$1:$1,0)-1)","v":0.86170023179039701},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(AZ$1&\"*\",SQ!$1:$1,0)-1)","v":15.620036541248499},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(BA$1&\"*\",SQ!$1:$1,0)-1)","v":3.7038468655858199},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(BB$1&\"*\",SQ!$1:$1,0)-1)","v":5.0915009974236204E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(BC$1&\"*\",SQ!$1:$1,0)-1)","v":6.0390201324401897E-4},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(BD$1&\"*\",SQ!$1:$1,0)-1)","v":4.4692915296460399E-3},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(BE$1&\"*\",SQ!$1:$1,0)-1)","v":-2.6622118797753599E-2},{"$f":"OFFSET(SQ!$A$1,MATCH($A124,SQ!$C:$C,0)-1,MATCH(BF$1&\"*\",SQ!$1:$1,0)-1)","v":1.4580014296477001E-4}]

{"sheet":"QCReview","index":3,"n_rows":40,"n_cols":49,"rows_kept":40,"cut":false,"hidden":false,"merged":[]}
["Data File","Name","Acq. Date-Time","Li_7","Be_9","B_11","Al_27","24  Mg","28  Si","31  P","39  K","44  Ca","49  Ti","51  V","52  Cr","55  Mn","56  Fe","59  Co","60  Ni","65  Cu","66  Zn","71  Ga","74  Ge","75  As","78  Se","79  Br","85  Rb","88  Sr","90  Zr","98  Mo","105  Pd","109  Ag","114  Cd","115  In","118  Sn","121  Sb","125  Te","127  I","133  Cs","137  Ba","184  W","195  Pt","197  Au","202  Hg","205  Tl","208  Pb","209  Bi","232  Th","238  U"]
[]
[null,null,"QC target",10,0.25,25,10,200,500,500,500,500,5,0.5,0.5,0.5,20,0.25,1,5,50,0.5,0.5,2.5,10,250,25,5,2,0.5,0.5,1,0.25,0.25,1,0.5,0.5,10,0.5,2.5,0.5,0.5,5,0.5,0.25,2.5,0.25,0.25,0.05]
[null,null,"Level",1,1,1,1,1,2,1,1,1,2,1,1,1,1,1,1,1,1,1,2,1,1,3,1,2,2,2,2,1,1,1,2,2,2,3,1,1,2,2,2,2,1,1,1,1,1]
[null,null,"HPA"]
[null,null,"HPB"]
[null,null,"HPC"]
[null,null,"QCs"]
["017SMPL.d","QC1","2019-10-08 16:25:02",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":9.6530000000000005},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0.24},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":22.556999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":9.8699999999999992},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":193.47529205639501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":71.8594432687687},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":534.93331460431898},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":581.65424413090602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":435.93143672398901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":1.04324168613828E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":0.50031823964323496},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":0.45311000716476502},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":0.45988192477422901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":20.046022447567498},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":0.23400542363395199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":0.95817402841591404},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":5.07564343002461},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":49.404099160137299},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":0.50055946428345899},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":1.23219088706818E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":2.5185330980854501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":10.3067051313795},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-0.68862310781465697},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":24.6585001079258},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":1.4962672104679799E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":9.40645522963228E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":1.3680360237351499E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":1.12826465237853E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":0.92861811224980495},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":0.224391282569079},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":0.246646541256343},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-2.6966187559756601E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-1.9319581035743199E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":0.479718774309917},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":2.3797567378469999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":3.9621506252572603E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":3.1905363586779898E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":-1.30712173756754E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":2.7021647651161301E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":0.22368362264776701},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":2.3438021852834701},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":0.25984102657326602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":0.218682335792486},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A9,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":4.9000755301144901E-2}]
["018SMPL.d","QC2","2019-10-08 16:26:32",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-7.0000000000000001E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":0.123},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.26400000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":-0.75666021823349106},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":582.96398039961798},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":-2.9992407880433301},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":-4.3124248217012697},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":87.615925594226994},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":4.9748845496121596},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":1.8504090504882501E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":2.6588800937474202E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":7.6299917482125197E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":-1.24538375525829E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":-1.74585720719811E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":9.7938546619919797E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":3.7100398333140203E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":0.10687509312948799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-1.16406176650329E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":0.56798150017591098},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":1.9580234219967402E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":6.08524703366937E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-2.68791319816859},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-2.2033149240918599E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":4.8783012043885696},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":2.4624686572192398},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":0.65732553977758001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":0.56169252250319501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":2.6130745817911E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":1.0570269650828901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":0.48006782623745398},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":0.37310194280966003},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":0.168252446082372},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":8.4130913455945098E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":1.0821678828490699E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":0.40751748870843202},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":0.39720220550103602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":9.1721517388611797},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.54101904549429503},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":-9.8145149100753895E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":6.6280077568774097E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":3.39509303925444E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":8.2790426594031893E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A10,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":1.62887511982266E-4}]
["019SMPL.d","QC3","2019-10-08 16:28:05",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.6E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-0.754},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":3.9E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":-0.270415038546232},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":55.834901716241298},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":-2.2831939565882702},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":-5.9919634294538602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":11.4119976603869},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":4.1891614090905999E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":3.8205117380328401E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":3.1111120442590099E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":2.6083795229040201E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":-2.4412958316584001E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":-2.3773351029348E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":-3.6202088510536898E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":3.7603204277101399E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":0.35772335251415699},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":2.8584174553579599E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":1.7708023099304801E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":5.1459677990209801E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":323.12010784299298},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-5.5077159457943198E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":2.6481479303063701E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":3.1131680597503001E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":1.61752725038453E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":-1.8563436428617899E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":4.0788660016112901E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":3.0344704065749298E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":8.48095951229209E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":38.153483598572002},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":2.09827118125033E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":5.8245804257145897E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":7.5707878215819403E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":2.0726452566949501E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":0.62105767522797095},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.13902041081187999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":1.2518687576999001E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-8.3179478410282896E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-4.0775566738809601E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":-2.3790987713739099E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A11,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":4.0146131139437903E-5}]
["050SMPL.d","QC1","2019-10-08 17:14:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":8.5139999999999993},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0.21099999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":19.602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":8.9},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":180.71063948691099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":-137.37555664339999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":449.659182698216},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":509.50512352205499},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":378.75559869518997},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":-5.1402861051569E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":0.44575373577775101},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":0.44711397777342698},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":0.46308885528959698},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":17.709543709225699},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":0.20697248682051},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":0.88745615928769706},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":4.5006272265695602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":42.7624616138785},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":0.44235442630143601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":-2.0786755951357899E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":2.2798906623553399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":7.9895749025331702},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":7.2606901628067},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":21.224836651076298},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":9.8802994028837493E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":-4.6586212441304797E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":3.0144324505319499E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":-4.7920896373623203E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":0.81312160603667305},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":0.210932860609923},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":0.211447995002015},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-7.1281577971100503E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":7.1993426376812397E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-0.384230682616858},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":0.45216308043174602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":2.1326937795210901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":-9.7743487724956493E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":3.0440004505513901E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":-0.51487172990338004},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":4.0178733120950797E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":0.18995279692808301},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":1.9855488339795799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":0.237213785892985},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":0.20103154901252401},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A12,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":4.0606714830781203E-2}]
["051SMPL.d","QC2","2019-10-08 17:16:13",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.6E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-0.64400000000000002},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.32500000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":0.704938386543706},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":150.581516610791},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":3.1913002086678199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":-8.9023254577374509},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":56.251809267874897},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":4.6648107431519001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":2.7052281861173798E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":2.7334788338699199E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":-4.9341544908616003E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":-0.12016771202649799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":1.84603823941218E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":-1.8567023310443499E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":-1.2365606943680401E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":-0.105304081824557},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":4.7448088556275303E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":0.58946258788505801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":-2.5520864190259501E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":9.9416565037832196},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":2.8197510515443901E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":4.5993333721576004},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":2.59789010989418},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":0.69068760481784297},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":0.56196853229407795},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":2.9202925226477602E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":1.00164076808796},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":0.47512606879184399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":0.41932739529072899},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-0.36329439199221703},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-3.9413495726622698E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":1.26148491966328E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":0.433219167351866},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":0.41878369763303802},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":9.1307112455599704},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.53795003954929499},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":3.2804794880078798E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-5.9010837187002804E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":2.4329933978373699E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":4.54339877296005E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":1.12414314160483E-4}]
["052SMPL.d","QC3","2019-10-08 17:17:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.9E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-1.367},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.104},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":-296.41904436941002},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":1.9525249220734899},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":2.4833936787439201},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":-6.3590943021714796},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":-5.9698872312229499E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":9.7401640411938304E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":4.1312705881027402E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":-3.1916488555316203E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":-3.3590169774331299E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":5.4035511134293104E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":-3.2727181055459402E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":-1.64867834007262E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":4.6464299743600203E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-2.6532315536437999E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":3.4749487726388798E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":2.0402157222646201E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":318.62301679265801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-9.8551044963291392E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":1.6753134107673599E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":5.9632157250182304E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":2.2840376104703301E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":-5.7404707630923403E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":8.7346395854242791E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":8.5488571819327601E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-5.6665202442374597E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":-9.2764177251645401E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":33.841728693253103},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":3.5666764259081001E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":-1.5955814053952599E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":5.68720106045061E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":3.4175038210154402E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":2.6605452750054499E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.123530768207625},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":4.7704537063917199E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-3.33480607164631E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":3.2360013374495798E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":1.07797854691435E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":"<0.000"}]
["081SMPL.d","QC1","2019-10-08 17:14:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":9.7309999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0.23300000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":22.873000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":10.724},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":197.81832298116299},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":-93.138838817223601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":480.43485257375602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":576.77539152750001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":495.46042329935898},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":-5.6667022342811501E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":0.52727764224069595},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":0.52375590459027599},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":0.46748201089008901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":21.2276529828763},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":0.245412345030427},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":1.1093954834330699},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":5.1467899294425701},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":51.364981848890203},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":0.55212452854668503},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":1.9678021940767901E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":2.7210905669524599},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":9.30825986968709},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-9.3503011528544597},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":24.964970910000801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":4.1688376086970097E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":-1.31243665713158E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":1.1065925910642501E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":-3.9122716261944201E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":0.91214823426579605},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":0.234491333152404},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":0.255479950718906},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":2.2599717969611802E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":5.2909141247443501E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-8.1808831343047603E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":0.52629006440069603},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":2.4451946963097102},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":-1.2569510132956899E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":7.2802878712691001E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":-2.6353808023005298E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":-3.3011342913611803E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":0.23708614773220299},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":2.3893493702582398},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":0.27422665197048202},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":0.20509947686583099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":4.7516913150104402E-2}]
["082SMPL.d","QC2","2019-10-08 17:16:13",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.2999999999999999E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-0.53100000000000003},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.82499999999999996},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":0.628812112409744},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":339.31771029999697},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":2.3566869356857798},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":11.415445687568001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":64.719726488485193},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":4.7694081948412803},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":-3.8459491928068E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":4.7845010901305003E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":1.8397001297722801E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":2.2989123845781299E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":9.1088650283184498E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":0.12051383204180401},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":3.36700815424344E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":9.6521990490367401E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":1.6025377095565099E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":0.59688824504159399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":"<0.013"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-9.7031292257478601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":2.4546343304715602E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":4.9944486697752399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":2.7001291034620798},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":0.66721199691865796},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":0.58295027693464896},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":-5.0711051121000101E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":1.0994816319977301},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":0.50079650603003201},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":0.42677206738677897},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-0.24915296839933801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-4.39216016839656E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":0.40638711024189},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":0.430540626568853},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":9.26817753715512},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.50778669339083604},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":1.6551144174156999E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-2.01972863267527E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-1.3847104987963601E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":2.18659355341938E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":3.98774555988047E-5}]
["083SMPL.d","QC3","2019-10-08 17:17:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.7999999999999999E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-1.3280000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.187},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":0.62340573953096201},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":-251.597045216048},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":2.28595752219104},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":-6.6895039659280302},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":3.51644514250034},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":-4.1980701647880397E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":8.6074580040670901E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":3.3534530144258098E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":4.9848683556103598E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":0.13543235247594301},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":1.6897499830455801E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":2.9916176200969601E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":1.9994787516733302E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":0.16745977183076799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-2.3606318026776199E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":1.3658815064521899E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":-1.804094104184E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":-4.2618341914710797E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":326.15001875227199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":1.6824203017732998E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":-4.3070494203753902E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":3.0936337657803798E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":1.1186342610217E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":2.7849817391527401E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":-2.61454858020146E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":3.1917526009620903E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":-3.0638356180765999E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":29.5926558903009},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-1.5507139584898201E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":-4.0197831498678498E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":7.9902623941116198E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":3.5487052654667099E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":0.51043310887365201},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.116733141337615},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":5.5147615813231902E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-2.90651144510853E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-4.0716033705805298E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":-2.9542216127766498E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":1.21425400475665E-4}]
["112SMPL.d","QC1","2019-10-08 17:14:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":10.589},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0.25700000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":23.12},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":11.116},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":220.83419766923899},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":40.304251592421899},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":563.26944297702096},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":663.88900653977396},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":504.25098677805897},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":1.4097343294491201E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":0.51412493669433001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":0.59189692556585105},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":0.56447664855525503},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":22.762159644104099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":0.26271126377944798},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":1.0755266091070901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":5.4254122590320897},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":55.640468881839098},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":0.52486893270367696},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":2.8258616336118902E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":2.8784943300837198},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":9.7950276705313293},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-5.46589019812399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":26.759193709314602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":-9.4573379596278994E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":1.5682928893836501E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":-3.72456594047671E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":7.2338267023185103E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":0.93543577052664895},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":0.259465966517681},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":0.26096942342815199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-2.6485849002896002E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":7.0000696736755897E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-0.68895096783876697},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":0.53514902385946805},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":2.6964792074085602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":-2.1717949149499701E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":5.4638342976612696E-6},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":-0.120251570707163},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":-2.9498594202091302E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":0.25515908519085201},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":2.6110162742254701},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":0.29852553525505099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":0.224850705060963},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":5.8537421534485802E-2}]
["113SMPL.d","QC2","2019-10-08 17:16:13",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.2999999999999999E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":1E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-0.36699999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.33600000000000002},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":0.11168136862635},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":468.09417768486202},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":-3.59032542270768},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":20.495367761667801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":85.340770309392397},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":4.9364452809932704},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":1.3192441446383599E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":4.57313104375087E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":2.3044792608691699E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":0.25376969451440101},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":-1.01267312854121E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":2.64408496691995E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":2.3785777350174599E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":0.13936119281144199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-6.8394536118907799E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":0.61530124436278399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":"<0.014"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-0.94293627548470804},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-1.8651289877156699E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":5.1233433591773503},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":2.6381149764238501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":0.66835984768944601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":0.544577383656354},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":7.3365068803705304E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":0.995164331159898},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":0.50254387792545696},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":0.36104432002066},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":-0.76302896678680399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-1.1892525295235201E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":-6.7416555531316603E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":0.43043268086261399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":0.43408653433532401},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":9.04952796357958},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.51347373806032004},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":-4.6011096517939898E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-2.2298025630835199E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-3.6658882445908701E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":3.0862013742802102E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":8.6002598672536398E-5}]
["114SMPL.d","QC3","2019-10-08 17:17:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.4E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":1E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-1.25},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.17100000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":-0.69662567069227699},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":122.998235591191},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":-2.44683643717857},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":35.684440847987801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":7.22298976301997},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":5.0172108503836001E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":2.5066264137036401E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":5.3945730862118499E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":4.8407754608834501E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":-7.2451439839742199E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":-7.2232336310132296E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":1.7435230289790399E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":2.5854719784470098E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":0.16739655112184601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-4.7023226223733196E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":1.53241154689739E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":5.6961056760779401E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":3.1390308205810402E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":303.01377820497498},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-4.2052186782034297E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":-1.3960186663748901E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":3.6669977257258201E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":9.3268210597138197E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":5.14426669630843E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":-2.0932705348014499E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":4.8594269360149102E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-2.6992283159671799E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":-1.6786526964371899E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":22.033046859313501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-6.9528613936351697E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":-3.1627141703292498E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":7.7218097833524797E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":3.09392063565517E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":0.53109808366759303},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.14439541843804601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":-2.0626638865900901E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-5.1483781087286401E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":-2.60610136121187E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":"<0.000"}]
["115SMPL.d","QC1","2019-10-08 17:14:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":9.94},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0.23300000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":21.565000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":16.158000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":212.31996034231099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":15.937373789489101},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":529.56066733597299},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":628.06428087334496},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":3897.31839040351},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":0.33445486061326302},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":0.51137094862321597},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":0.55993851833228003},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":1.23427314279694},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":22.5232006656747},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":0.24920547567498799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":1.3333462360202499},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":5.6946674508700097},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":71.096719361902899},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":0.51198331185005597},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":4.41917669119253E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":2.7173624502857399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":9.6822987374785292},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":0.75402697140602704},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":24.460581049793099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":2.9446795015191398},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":7.5439846329458493E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":1.8763274936661301E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":2.4483175316624499E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":0.98074726572857496},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":0.24208372943072301},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":0.24982458895944901},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-1.44251224103117E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":1.6711959254226E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":3.7724873660797198},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":0.48965976627843999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":2.7630014394700502},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":2.9573294559228701E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":5.8881864984888701E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":0.24463250795716501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":2.7871755315490701E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":0.24035280612849799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":2.5591734445641001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":0.28208831840464399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":0.23896254163606501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":5.5326373443955697E-2}]
["116SMPL.d","QC2","2019-10-08 17:16:13",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-3.0000000000000001E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":1E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-0.23699999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":1.4850000000000001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":0.99733178808766398},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":539.26888277892601},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":-2.74487319950613},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":15.7727183081162},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":411.19160444916201},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":5.3837785080682199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":3.44107305554966E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":4.1110932056147402E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":0.102829490583999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":0.55213184607064403},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":4.7760755644806804E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":0.13824467067773599},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":0.47682799214333199},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":2.31036195144249},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-6.4064647116603502E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":0.61359347707978795},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":"<0.001"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":8.0004149074589206E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":-3.6546415553378799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-3.2973762524653102E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":5.17333116834976},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":2.7120346150136898},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":0.66146491712212097},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":0.59878729204688397},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":1.3248924098671599E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":0.98118408712764604},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":0.49423449624964799},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":0.44496920028984099},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":0.19559925787013399},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-1.8202357898855601E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":4.3436566998494103E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":0.42441420730547103},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":0.43754034370841299},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":9.1839972011011},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.70581211073970695},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":-6.3878665806584007E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":1.31846878891431E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-1.5127689324286901E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":2.4440563038097499E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":4.0923105233620202E-4}]
["117SMPL.d","QC3","2019-10-08 17:17:40",{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1)","v":-1.4E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1)","v":0},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1)","v":-1.1679999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1)","v":0.14499999999999999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1)","v":0.178140591470733},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1)","v":36.629992700653197},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1)","v":-6.1076430924259597},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1)","v":27.516098098323202},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1)","v":11.356083658498999},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1)","v":4.7326617211941799E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1)","v":3.49488933878063E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1)","v":7.0607302673759698E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1)","v":5.7825306157468302E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1)","v":0.47731751135601602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1)","v":-1.0969950871041699E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1)","v":1.14189769648984E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1)","v":-8.0643873415087598E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1)","v":0.22513035478612001},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1)","v":-6.2272829204975096E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1)","v":-1.21099918682966E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1)","v":2.7070295606891502E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1)","v":6.58215951624416E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1)","v":279.40181851843801},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1)","v":-5.0769708919601399E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1)","v":-2.3726490019743401E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1)","v":8.7068234428997002E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1)","v":7.97374806951422E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1)","v":2.1629423957364902E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1)","v":-1.7907465624321699E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1)","v":8.1999273026621205E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1)","v":"<0.000"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1)","v":-2.4021825420287201E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1)","v":-7.2537937781455802E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1)","v":"<0.005"},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1)","v":25.265847728206602},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1)","v":-1.11595514202198E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1)","v":-1.0984556792576799E-4},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1)","v":6.3676971672915299E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1)","v":3.4058536200436001E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1)","v":0.85435456512576702},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1)","v":0.16838820506556501},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1)","v":-6.4826693681489195E-5},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1)","v":1.4480290764552399E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1)","v":-4.8260373376823604E-3},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1)","v":-2.6259664491688001E-2},{"$f":"OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1)","v":"<0.000"}]
[]
[null,"Level","Accuracy"]
[null,1,"HPA",{"$f":"IFERROR(IF($B26=D$4,D9/D$3,\"\"),\"\")","v":0.96530000000000005},{"$f":"IFERROR(IF($B26=E$4,E9/E$3,\"\"),\"\")","v":0.96},0.90227999999999997,0.98699999999999988,0.96737646028197499,"",1.0698666292086381,1.1633084882618121,0.87186287344797797,"",1.0006364792864699,0.90622001432953003,0.91976384954845802,1.0023011223783749,0.93602169453580797,0.95817402841591404,1.015128686004922,0.98808198320274598,1.001118928566918,"",1.0074132392341801,1.0306705131379501,"",0.98634000431703195,"","","","",0.92861811224980495,0.89756513027631601,0.986586165025372,"","","","",0.95943754861983399,0.95190269513879999,"","","","",0.89473449059106802,0.93752087411338803,1.0393641062930641,0.874729343169944,0.98001510602289799]
[null,2,"HPB",{"$f":"IFERROR(IF($B27=D$4,D10/D$3,\"\"),\"\")","v":""},"","","","",1.1659279607992359,"","","",0.99497690992243193,"","","","","","","","","",1.135963000351822,"","","","",0.97566024087771397,1.2312343286096199,1.31465107955516,1.12338504500639,"","","",1.0570269650828901,0.96013565247490795,0.74620388561932005,"","","",0.81503497741686404,0.79440441100207204,1.8344303477722359,1.0820380909885901,"","","","",""]
[null,3,"HPC","","","","","","","","","","","","","","","","","","","","","","",1.292480431371972,"","","","","","","","","","","",3.8153483598572002,"","","","","","","","","","",""]
[null,1,"HPA",0.85139999999999993,0.84399999999999997,0.78408,0.89,0.90355319743455498,"",0.89931836539643195,1.01901024704411,0.75751119739037998,"",0.89150747155550203,0.89422795554685397,0.92617771057919396,0.88547718546128495,0.82788994728204002,0.88745615928769706,0.90012544531391203,0.85524923227756999,0.88470885260287202,"",0.911956264942136,0.79895749025331697,"",0.84899346604305193,"","","","",0.81312160603667305,0.843731442439692,0.84579198000806,"","","","",0.90432616086349205,0.85307751180843605,"","","","",0.75981118771233203,0.79421953359183195,0.94885514357194001,0.80412619605009605,0.81213429661562397]
[null,2,"HPB","","","","","",0.301163033221582,"","","",0.93296214863038007,"","","","","","","","","",1.178925175770116,"","","","",0.91986667443152004,1.29894505494709,1.3813752096356859,1.1239370645881559,"","","",1.00164076808796,0.95025213758368798,0.83865479058145798,"","","",0.86643833470373199,0.83756739526607604,1.8261422491119941,1.07590007909859,"","","","",""]
[null,3,"HPC","","","","","","","","","","","","","","","","","","","","","","",1.2744920671706321,"","","","","","","","","","","",3.3841728693253104,"","","","","","","","","","",""]
[null,1,"HPA",0.97309999999999997,{"$f":"IFERROR(IF($B32=E$4,E15/E$3,\"\"),\"\")","v":0.93200000000000005},0.91492000000000007,1.0724,0.98909161490581499,"",0.96086970514751202,1.153550783055,0.99092084659871793,"",1.0545552844813919,1.047511809180552,0.93496402178017801,1.061382649143815,0.98164938012170799,1.1093954834330699,1.029357985888514,1.027299636977804,1.1042490570933701,"",1.088436226780984,0.930825986968709,"",0.998598836400032,"","","","",0.91214823426579605,0.93796533260961601,1.021919802875624,"","","","",1.0525801288013921,0.97807787852388406,"","","","",0.94834459092881196,0.95573974810329587,1.0969066078819281,0.82039790746332397,0.95033826300208801]
[null,2,"HPB","","","","","",0.67863542059999393,"","","",0.95388163896825606,"","","","","","","","","",1.193776490083188,"","","","",0.99888973395504799,1.3500645517310399,1.3344239938373159,1.1659005538692979,"","","",1.0994816319977301,1.001593012060064,0.85354413477355795,"","","",0.81277422048378001,0.861081253137706,1.853635507431024,1.0155733867816721,"","","","",""]
[null,3,"HPC","","","","","","","","","","","","","","","","","","","","","","",1.304600075009088,"","","","","","","","","","","",2.9592655890300898,"","","","","","","","","","",""]
[null,1,"HPA",1.0589,1.028,0.92480000000000007,1.1115999999999999,1.104170988346195,"",1.1265388859540419,1.327778013079548,1.008501973556118,"",1.02824987338866,1.1837938511317021,1.1289532971105101,1.1381079822052049,1.0508450551177919,1.0755266091070901,1.0850824518064179,1.112809377636782,1.0497378654073539,"",1.1513977320334878,0.97950276705313288,"",1.070367748372584,"","","","",0.93543577052664895,1.037863866070724,1.043877693712608,"","","","",1.0702980477189361,1.0785916829634241,"","","","",1.020636340763408,1.044406509690188,1.194102141020204,0.89940282024385199,1.170748430689716]
[null,2,"HPB","","","","","",0.936188355369724,"","","",0.98728905619865404,"","","","","","","","","",1.230602488725568,"","","","",1.0246686718354701,1.319057488211925,1.336719695378892,1.089154767312708,"","","",0.995164331159898,1.0050877558509139,0.72208864004132001,"","","",0.86086536172522798,0.86817306867064803,1.809905592715916,1.0269474761206401,"","","","",""]
[null,3,"HPC","","","","","","","","","","","","","","","","","","","","","","",1.2120551128199,"","","","","","","","",{"$f":"IFERROR(IF($B37=AI$4,AI20/AI$3,\"\"),\"\")","v":""},"","",2.2033046859313501,"","","","","","","","","","",""]
[null,1,"HPA",{"$f":"IFERROR(IF($B38=D$4,D21/D$3,\"\"),\"\")","v":0.99399999999999999},0.93200000000000005,0.86260000000000003,1.6158000000000001,1.0615998017115549,"",1.0591213346719459,1.25612856174669,7.7946367808070196,"",1.0227418972464319,1.1198770366645601,2.46854628559388,1.1261600332837349,0.99682190269995197,1.3333462360202499,1.138933490174002,1.4219343872380579,1.0239666237001119,"",1.0869449801142959,0.96822987374785296,"",0.97842324199172392,"","","","",0.98074726572857496,0.96833491772289204,0.99929835583779603,"","","","",0.97931953255687998,1.1052005757880201,"","","","",0.96141122451399197,1.02366937782564,1.128353273618576,0.95585016654426003,1.1065274688791138]
[null,2,"HPB",{"$f":"IFERROR(IF($B39=D$4,D22/D$3,\"\"),\"\")","v":""},"","","","",1.0785377655578521,"","","",1.076755701613644,"","","","","","","","","",1.2271869541595759,"","","","",1.0346662336699519,1.3560173075068449,1.3229298342442419,1.1975745840937679,"","","",0.98118408712764604,0.98846899249929598,0.88993840057968199,"","","",0.84882841461094205,0.87508068741682599,1.8367994402202199,1.4116242214794139,"","","","",""]
[null,3,"HPC",{"$f":"IFERROR(IF($B40=D$4,D23/D$3,\"\"),\"\")","v":""},"","","","","","","","","","","","","","","","","","","","","",1.1176072740737519,"","","","","","","","","","","",2.5265847728206601,"","","","","","","","","","",""]

{"sheet":"Container1","index":4,"n_rows":37,"n_cols":49,"rows_kept":37,"cut":false,"hidden":false,"merged":[]}
[null,"Name","Acq. Date-Time","Li_7","Be_9","B_11","Al_27","24  Mg","28  Si","31  P","39  K","44  Ca","49  Ti","51  V","52  Cr","55  Mn","56  Fe","59  Co","60  Ni","65  Cu","66  Zn","71  Ga","74  Ge","75  As","78  Se","79  Br","85  Rb","88  Sr","90  Zr","98  Mo","105  Pd","109  Ag","114  Cd","115  In","118  Sn","121  Sb","125  Te","127  I","133  Cs","137  Ba","184  W","195  Pt","197  Au","202  Hg","205  Tl","208  Pb","209  Bi","232  Th","238  U"]
[null,null,"RL",10,0.25,25,10,200,500,500,500,500,5,0.5,0.5,0.5,20,0.25,1,5,50,0.5,0.5,2.5,10,250,25,5,2,0.5,0.5,1,0.25,0.25,1,0.5,0.5,10,0.5,2.5,0.5,0.5,5,0.5,0.25,2.5,0.25,0.25,0.05]
[null,null,"Low QC",{"$f":"MIN(QCReview!D26:D100)*D2","v":8.5139999999999993},{"$f":"MIN(QCReview!E26:E100)*E2","v":0.21099999999999999},{"$f":"MIN(QCReview!F26:F100)*F2","v":19.602},{"$f":"MIN(QCReview!G26:G100)*G2","v":8.9},{"$f":"MIN(QCReview!H26:H100)*H2","v":180.71063948691099},{"$f":"MIN(QCReview!I26:I100)*I2","v":150.581516610791},{"$f":"MIN(QCReview!J26:J100)*J2","v":449.659182698216},{"$f":"MIN(QCReview!K26:K100)*K2","v":509.50512352205499},{"$f":"MIN(QCReview!L26:L100)*L2","v":378.75559869518997},{"$f":"MIN(QCReview!M26:M100)*M2","v":4.6648107431519001},{"$f":"MIN(QCReview!N26:N100)*N2","v":0.44575373577775101},{"$f":"MIN(QCReview!O26:O100)*O2","v":0.44711397777342698},{"$f":"MIN(QCReview!P26:P100)*P2","v":0.45988192477422901},{"$f":"MIN(QCReview!Q26:Q100)*Q2","v":17.709543709225699},{"$f":"MIN(QCReview!R26:R100)*R2","v":0.20697248682051},{"$f":"MIN(QCReview!S26:S100)*S2","v":0.88745615928769706},{"$f":"MIN(QCReview!T26:T100)*T2","v":4.5006272265695602},{"$f":"MIN(QCReview!U26:U100)*U2","v":42.7624616138785},{"$f":"MIN(QCReview!V26:V100)*V2","v":0.44235442630143601},{"$f":"MIN(QCReview!W26:W100)*W2","v":0.56798150017591098},{"$f":"MIN(QCReview!X26:X100)*X2","v":2.2798906623553399},{"$f":"MIN(QCReview!Y26:Y100)*Y2","v":7.9895749025331693},{"$f":"MIN(QCReview!Z26:Z100)*Z2","v":279.40181851843801},{"$f":"MIN(QCReview!AA26:AA100)*AA2","v":21.224836651076298},{"$f":"MIN(QCReview!AB26:AB100)*AB2","v":4.5993333721576004},{"$f":"MIN(QCReview!AC26:AC100)*AC2","v":2.4624686572192398},{"$f":"MIN(QCReview!AD26:AD100)*AD2","v":0.65732553977758001},{"$f":"MIN(QCReview!AE26:AE100)*AE2","v":0.544577383656354},{"$f":"MIN(QCReview!AF26:AF100)*AF2","v":0.81312160603667305},{"$f":"MIN(QCReview!AG26:AG100)*AG2","v":0.210932860609923},{"$f":"MIN(QCReview!AH26:AH100)*AH2","v":0.211447995002015},{"$f":"MIN(QCReview!AI26:AI100)*AI2","v":0.98118408712764604},{"$f":"MIN(QCReview!AJ26:AJ100)*AJ2","v":0.47512606879184399},{"$f":"MIN(QCReview!AK26:AK100)*AK2","v":0.36104432002066},{"$f":"MIN(QCReview!AL26:AL100)*AL2","v":22.033046859313501},{"$f":"MIN(QCReview!AM26:AM100)*AM2","v":0.45216308043174602},{"$f":"MIN(QCReview!AN26:AN100)*AN2","v":2.1326937795210901},{"$f":"MIN(QCReview!AO26:AO100)*AO2","v":0.40638711024189},{"$f":"MIN(QCReview!AP26:AP100)*AP2","v":0.39720220550103602},{"$f":"MIN(QCReview!AQ26:AQ100)*AQ2","v":9.04952796357958},{"$f":"MIN(QCReview!AR26:AR100)*AR2","v":0.50778669339083604},{"$f":"MIN(QCReview!AS26:AS100)*AS2","v":0.18995279692808301},{"$f":"MIN(QCReview!AT26:AT100)*AT2","v":1.9855488339795799},{"$f":"MIN(QCReview!AU26:AU100)*AU2","v":0.237213785892985},{"$f":"MIN(QCReview!AV26:AV100)*AV2","v":0.20103154901252401},{"$f":"MIN(QCReview!AW26:AW100)*AW2","v":4.0606714830781203E-2}]
["Container Name",null,null,"Lithium","Beryllium","Boron","Aluminium","Magnesium","Silicon","Phosphorus","Potassium","Calcium","Titanium","Vanadium","Chromium","Manganese","Iron","Cobalt","Nickel","Copper","Zinc","Gallium","Germanium","Arsenic","Selenium","Bromine","Rubidium","Strontium","Zirconium","Molybdenum","Palladium","Silver","Cadmium","Indium","Tin","Antimony","Tellurium","Iodine","Caesium","Barium","Tungsten","Platinum","Gold","Mercury","Thallium","Lead","Bismuth","Thorium","Uranium"]
["Green top"]
["Number of containers"]
[20,null,"Average",{"$f":"AVERAGE(OFFSET(D$12,0,0,$A$7,1))","v":3.2073684210526316},{"$f":"AVERAGE(OFFSET(E$12,0,0,$A$7,1))","v":1.578947368421053E-3},-0.764421052631579,21.634789473684208,62.195009894886333,48.015646237661699,8.3713377652508285,332.50412253534085,201.40623368589667,9.6042825196524476E-2,5.5575510547624808E-2,8.9425273890718734E-2,8.4558083079880886E-2,5.352522710816416,3.5110759571943019E-3,6.2807253327921919E-2,0.3571961550009955,349.03721390375193,0.57746446188485889,4.3781352908288046E-3,1.1947423351617584E-2,4.2700192630651399E-3,852.64831087747791,7.1579681427747344E-2,1.6129520746104635,0.11245642392191342,1.7710077462553884,-5.1034920359463443E-3,3.7332017245677652E-4,1.4320324413975164E-3,0,-6.9138397078227743E-4,8.9436717017536332E-2,0,-0.37596337503866761,1.2721139325932722E-3,96.250758635617572,9.194459601868021E-3,7.4398169759386928E-4,5.3057963312404699E-2,6.1084747234449774E-2,2.066219423528976E-3,5.0352104485931068E-2,4.9818609634038637E-4,-2.1468185126943031E-4,6.035888497458005E-3]
[null,null,"SD",{"$f":"STDEV(OFFSET(D$12,0,0,$A$7,1))","v":5.8522750112257653E-2},{"$f":"STDEV(OFFSET(E$12,0,0,$A$7,1))","v":9.01590537370498E-4},9.5157364279428794E-2,0.96478187396296211,4.3119988895741947,102.51681218633205,3.4416630223435964,30.051773811984205,15.492224572056918,3.772514147103706E-2,1.1107168176518157E-2,3.5562845208708102E-2,4.6860505645854575E-2,4.0302607938678774,3.6923727039200605E-3,0.18532074675635982,8.5700559447023775E-2,89.908944411773561,4.8036244629323196E-2,5.3604648747454372E-3,7.9401335627241783E-3,3.4752456995766512E-2,136.92125890453229,4.4722422836897632E-2,7.4368854774424606E-2,1.2896449124785912E-2,4.0711938129533998E-2,1.9107413454635055E-3,4.1251700287726011E-4,1.2239241872789442E-3,0,8.7100597204155267E-3,7.6166219815629764E-3,0,9.9788239986596514E-2,7.7508140711810166E-4,1.7729158583326601,1.1318741448071296E-3,3.9710093362704953E-4,0.74501939734070044,1.9837271876044588E-2,2.9645538618713771E-4,9.3984972009562885E-3,3.8480490334369107E-4,2.7966955941086493E-4,1.632584520991366E-3]
[null,null,"ave+2sd",{"$f":"D7+D8*2","v":3.3244139212771469},{"$f":"E7+E8*2","v":3.382128443162049E-3},-0.57410632407272144,23.564353221610133,70.819007674034722,253.04927061032578,15.254663809938021,392.60767015930924,232.39068283001049,0.17149310813859858,7.7789846900661125E-2,0.16055096430813492,0.17827909437159004,13.413044298552171,1.0895821365034423E-2,0.43344874684064155,0.52859727389504307,528.85510272729903,0.67353695114350531,1.509906504031968E-2,2.7827690477065941E-2,7.3774933254598163E-2,1126.4908286865425,0.16102452710154261,1.7616897841593127,0.13824932217148525,1.8524316225144564,-1.2820093450193332E-3,1.1983541782112968E-3,3.8798808159554049E-3,0,1.6728735470048776E-2,0.10466996098066228,0,-0.17638689506547459,2.8222767468294753E-3,99.796590352282891,1.145820789148228E-2,1.5381835648479684E-3,1.5430967579938055,0.10075929098653895,2.6591301959032515E-3,6.9149098887843652E-2,1.2677959030277684E-3,3.4465726755229955E-4,9.3010575394407366E-3]
[null,null,"Pass/Review",{"$f":"IF(D9>MIN(D2:D3),\"Review\",\"Pass\")","v":"Pass"},{"$f":"IF(E9>MIN(E2:E3),\"Review\",\"Pass\")","v":"Pass"},"Pass","Review","Pass","Review","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Review","Review","Pass","Pass","Pass","Review","Pass","Pass","Pass","Review","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Review","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass"]
[]
["Containers"]
["025SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A13,ConcTable!$A:$A,0)-1,0)","v":"19286427-001"},{"$f":"OFFSET(SQ!$D$1,MATCH($A13,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:00:10"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.33},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.65600000000000003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.213000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":63.4912044107095},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":45.0398522725541},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":4.6865062831858504},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":319.40276756943803},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":195.89632945595},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.154232974978737},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.65653229020398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":4.8473380143762998E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":7.8188186051413894E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.0380030934308202},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.87939985738687E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":3.3661305492135599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.36544691964916598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":361.19103426047297},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.62262261292667698},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.4991230801282399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.8864234753814099E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":832.71821376140394},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.9907812796367301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.63736512611301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.11544028701620999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7335742135113501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.21146292861823E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.4686955960942905E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.40072946775499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.5788569559229306E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.184888111688673},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.0911638958172E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":98.510793639238301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":9.7466056210821408E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":8.7244616198562796E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":2.4412296507629798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":8.9114090888553696E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.2337520332230401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.6201498112512902E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":7.6172898091019401E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-1.64200149031535E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":5.07605638477649E-3}]
["026SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A14,ConcTable!$A:$A,0)-1,0)","v":"19286427-002"},{"$f":"OFFSET(SQ!$D$1,MATCH($A14,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:01:38"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.1720000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.77900000000000003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.547999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":61.352194781947802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":51.872284134678502},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":7.5023627215957296},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":282.77490472845398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":212.27721886086201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":6.1196400137132997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":4.95877514295553E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.185779596184666},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.14874406432760201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.40878802228189},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.8151605621368201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":0.821558528975186},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.52711690611765605},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":309.15431109251602},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.62998222107972701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-3.5135060528698001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":3.5429017698576102E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.05641178321035E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":891.66042347063399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":8.7307212905958595E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.6018115877475001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.12600334690963499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.75218970037525},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-4.3391155980295401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-9.8832052049348892E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":2.0227364529245401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.5899942458092401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.6922809414604493E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.37021292173272302},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.0292830465577199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":95.949891551133305},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":9.6948085929422605E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.4722473922604701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":1.5439548041378801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":8.9074055350800305E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.2660612487443799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.4353034579836798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":3.4859787184530799E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-4.8190821936929998E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":5.4626945185020804E-3}]
["027SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A15,ConcTable!$A:$A,0)-1,0)","v":"19286427-003"},{"$f":"OFFSET(SQ!$D$1,MATCH($A15,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:03:06"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.1930000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.72599999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":20.721},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":63.9827978598424},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":126.05708626360899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":7.1607279564380999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":290.69830034052097},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":211.254457091951},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.119547924283051},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.6506435722617897E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.11269709702557799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.21849468517179399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":20.511420342716701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":2.2025354263837801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":6.7193216554276403E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.37287919700609201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":314.76402872650903},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.58996678583025997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":5.4843540621438303E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":5.45103830204954E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.27527417351602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":749.70062867053502},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":5.5203116884592299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.53182158684368},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":9.9419568241586306E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.8455785497480699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.9537891006211504E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":1.63955112920327E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.36216563378081E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.50778725906179E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.1966413715769904E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.350260984176753},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":6.0112616434124797E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":94.919960079291599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.3015429130394395E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":2.08558612491557E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.40212067730648599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":7.5531487883070098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.7268875591105899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.8294304641010297E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":9.7505511008169404E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":3.0720756399051801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":5.8610768704855697E-3}]
["028SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A16,ConcTable!$A:$A,0)-1,0)","v":"19286427-004"},{"$f":"OFFSET(SQ!$D$1,MATCH($A16,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:04:34"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":3.0000000000000001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.77700000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.195},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":64.144061974241296},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":213.408753966067},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":5.7814106197622097},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":342.32843028786601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":191.516783773463},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":2.32143605796203E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":3.7141705622680599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":5.9396726726476001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.1776965498117E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":4.0929856024818303},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":4.6335232445378401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.5869318894424098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.32136625203824598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":254.494300480916},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.58371725834287802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.0573172453844201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":9.6955966227617804E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.31792555391531E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":715.04456372796506},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":8.3783727154277607E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.57785065140402},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.107506402413062},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.8680357940846599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.9668461499897502E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.25363232612874E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.0505342492072901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.9088209875076801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.6711340623068395E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.36133835160071798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":8.5293754379636797E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":97.488844712174398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":9.2842130570216404E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":6.4498463354536995E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.420323017957279},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":8.0674411168897606E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.3190630998677401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.2641864444097302E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":4.3346956522400499E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.5228408767120002E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.1268826069966397E-3}]
["029SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A17,ConcTable!$A:$A,0)-1,0)","v":"19286427-005"},{"$f":"OFFSET(SQ!$D$1,MATCH($A17,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:06:03"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.258},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.754},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":19.699000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":63.957129155750799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-7.5361235531122501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":6.7252729930738999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":285.02617760894799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":169.68423873056901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.10476775793160401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.3831941854075797E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":7.5796661041408003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.101622225492196},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":8.6287165415228806},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":7.0360607956609803E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":-1.6262620476379501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.28470451374449501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":218.117592522671},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.51462182134938295},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.2793311187244E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":3.8208456554634E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-1.4994038273350599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":672.05543941269696},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":5.44315261297523E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5159689465991899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.107037425950034},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.77888224595105},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.30159097233432E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.39475788091429E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-6.0295728820621702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":7.8826819432781997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.27723039400175198},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.23428227756802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":95.965010989977401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":7.8923330673012805E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.3643032768986001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.20663556583241299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":7.12063582455357E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.1450936139706701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":3.9433202304981298E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":7.7093756402037399E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.5479158834783701E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.59675386036177E-3}]
["030SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A18,ConcTable!$A:$A,0)-1,0)","v":"19286427-006"},{"$f":"OFFSET(SQ!$D$1,MATCH($A18,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:07:32"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2410000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":3.0000000000000001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.79900000000000004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.763000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":54.875119249938301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":77.371123563995795},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":11.185961858426399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":362.71816338720799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":185.749422782268},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":9.9126608283260906E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":8.0092948920798407E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":5.65384947033368E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.88658971331354E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":2.1553622098460501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.4520371172484599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":9.6731062227653104E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.34319494723718602},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":417.39551057858199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.64077697895722596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.35113016536361E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":2.0434508053648499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":970.86536304905405},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.81704330515655E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.7801860889902299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.111629564911262},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.8017301900087099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.9012126817211503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-4.7806549619083499E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.6793361889954603E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-1.4005810604407101E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.8044289025989301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.246425035311678},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":2.3180481563130199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":97.237542130627105},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":7.2251323739871403E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":2.1289586725666199E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-3.52853721518194E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":5.9756788409716301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.9419429618965499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.5051042123841603E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":1.07145010943114E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":1.03963220988186E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":7.0580086175119697E-3}]
["031SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A19,ConcTable!$A:$A,0)-1,0)","v":"19286427-007"},{"$f":"OFFSET(SQ!$D$1,MATCH($A19,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:09:02"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.165},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.78400000000000003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.131},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":64.068500692308902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":67.114353244794103},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":11.233547058553},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":335.74419777819202},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":217.43296133464699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":8.7991227915685602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.2794896760962003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":4.3144485670675699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":8.0459430694323195E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.4601686468015198},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.0540022254501998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":-3.18862478866458E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.289058525586674},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":380.54014810031703},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.545177075176982},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.8891253281115E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.6954960953962501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.2346155436415601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":783.78577243793495},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":0.13206852657176399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.7183582434834399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.119835052259715},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.74796089906184},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.8525687509638803E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":8.6272614060890296E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.29112894788909E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-1.1312263804383201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.2396612973842695E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.45461256635144398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":8.6624672639661103E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":100.343381635691},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.10517983342937E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":6.5072828518225297E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.238008778753936},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":6.86768673902933E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.6407675021586701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.5916267592408397E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":5.8514638283932202E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":4.2067836209728397E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":7.1505663573191202E-3}]
["032SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A20,ConcTable!$A:$A,0)-1,0)","v":"19286427-008"},{"$f":"OFFSET(SQ!$D$1,MATCH($A20,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:10:36"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2170000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.51800000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":20.803000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":58.450414751316998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":159.34964988054901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":14.022929382925099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":347.557717019775},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":198.25841764348201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":8.7374447419496404E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.0229962977707501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":8.0314550165689905E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.10910580927215401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":5.0217383576821399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.8007360974317101E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":-1.0081240826155099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.35537300399028299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":290.51746641676198},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.61789212781553904},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-6.9804828803798598E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":9.6580410133061904E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.2049665034831599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":688.00083429181598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":7.4792445764409493E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.54159905220056},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.11202683708515999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7517996887885301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.0373084414770401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.161482170186E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":9.9614164262248005E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-5.9017828282609403E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.56465053201144E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.34199925347501398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":3.15953447056586E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":95.514555141757796},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.7085060905308792E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":3.1900383505443998E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.118207133402691},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":6.2405243328435098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.4103435176756199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.7002199043130098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-2.6629681897535602E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-4.1769731961250402E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":3.62774315845778E-3}]
["033SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A21,ConcTable!$A:$A,0)-1,0)","v":"19286427-009"},{"$f":"OFFSET(SQ!$D$1,MATCH($A21,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:12:06"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.1859999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.66500000000000004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":20.736000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":62.204869761249803},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":73.287912316501703},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.7924861579253699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":318.87257586417098},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":188.54451019598201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.2497684118345798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.7897070403559001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.15472884628380801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":8.8069963366828402E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":5.1435651787493804},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.9361517453107798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":4.9786422421463902E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.47095658725037998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":294.05408904202301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.49989215567300099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":3.9576982819256104E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.51048445595773E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.2327092173658299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":769.42392889263704},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":9.0546105713352795E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5788814372809601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.113893904508163},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.79439277057172},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.8748937410074299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.25651089814039E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":4.0005366507700301E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-4.5541751981371201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.0661706805041195E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.38522424675620898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.9211459428094501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":97.249711718179498},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":9.7637828260273007E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":8.6027093013494704E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.16359427796051901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":4.1722154506029403E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.31988938931252E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.4294641989416602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":8.6762082017945196E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.35092838049771E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.9623449608090101E-3}]
["034SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A22,ConcTable!$A:$A,0)-1,0)","v":"19286427-010"},{"$f":"OFFSET(SQ!$D$1,MATCH($A22,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:13:36"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.0880000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.80100000000000005},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.748000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":66.857371943972893},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":109.419063821162},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":12.7684664175284},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":317.27421279048201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":196.235862420177},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.3658333237025503E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.6150378606981503E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":8.8816971919023199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":6.88743103942666E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.1788354439817099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":2.6413969002538899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":2.64475759801178E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.270304115696543},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":281.62266798591003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.52403589416738905},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.7674693996054701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.49184532798671E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":8.7460197728880307E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":722.70763405308696},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":7.3417781863249999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5226497449109599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":9.4545216722237702E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7816201821519599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-4.9819493090849003E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":6.3287860971198296E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-4.41436427473757E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.7066295775205499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.25356818709506901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.44365454111381E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":96.074624975813407},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":7.5861460456890196E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.15494774077605E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.27869960927957399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":3.26446067358694E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.7976419479327199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.4790649250225002E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":4.18322911082519E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.1151554015604701E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.5989241887317601E-3}]
["035SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A23,ConcTable!$A:$A,0)-1,0)","v":"19286427-011"},{"$f":"OFFSET(SQ!$D$1,MATCH($A23,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:15:14"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.278},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.83699999999999997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":23.213999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":70.725195464700406},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":38.6047802452917},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":10.8006506799583},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":348.16706834220997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":210.49701375356901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":5.9792930872762703E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.2524253037442995E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.102286607388701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.5607166177020202E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.6089003542958502},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.1724604001193701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.4811106796350801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.35923699589276398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":424.157863222245},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.60474546244534},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.1532935156376E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.31232322964355E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-3.21396827130775E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":969.13071007413805},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":1.6017083348509001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5770580440351001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.116256765730642},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.76368772742391},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.89328920596112E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.1963777019966299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":8.7655497605373001E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.9947152878701901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.2598277004106194E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.48334064971665502},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.0549308348388399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":97.370994282005498},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.7461917511704092E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":3.2032473435831101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.27949523536918902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":5.2315048137355699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.9482707213708901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.4135033061711302E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-4.7312829970648301E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.3221734086962491E-3}]
["036SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A24,ConcTable!$A:$A,0)-1,0)","v":"19286427-012"},{"$f":"OFFSET(SQ!$D$1,MATCH($A24,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:16:47"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.12},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.79300000000000004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":20.529},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":54.299929407721997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":185.554196797641},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":5.00569377526491},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":312.38376308351002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":192.34309940337599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.0202705128986603E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":3.8413741102069E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":6.4031945181542399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":7.4148029673349006E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":5.3034807113766202},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.0769996253284301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":9.5516260313095493E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.27899731708871101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":215.75779274737201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.638715062466928},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-2.1059736828014002E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.12501990961438E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":738.12940090994596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":9.6960849786132203E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5431304910118699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":8.8772574720244801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7739615844199399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-7.5310543878721803E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.0220611135031499E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.1387176017906701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-5.6625894011899398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.0232684734673099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.37243173799194901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.26007266275869E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":92.344249850109406},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.0174881843363098E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":8.3660902245721403E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.32664853263091898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":4.7988932205609201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.8478411774795801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.9819220673739999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-7.6543755765266601E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":8.3603154494653906E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.9452290533102596E-3}]
["037SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A25,ConcTable!$A:$A,0)-1,0)","v":"19286427-013"},{"$f":"OFFSET(SQ!$D$1,MATCH($A25,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:18:17"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.169},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":3.0000000000000001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.88100000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.709},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":68.303715744688304},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":127.427414042857},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":15.1448666505791},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":347.038681279971},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":215.11492039474001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":9.8459285203896907E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.9246874663452499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":9.8862512620969498E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":6.0558565879990603E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":4.2667486148368301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.1918969342189002E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":2.9986796445696798E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.365964756763349},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":351.72421260578801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.59248316360884101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.2415855588389299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":5.7768338510524198E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.2134950484808399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":839.74705356869799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":0.13224469050398499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5920184733034299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.123802281000383},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7640420919038999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.6895220559379198E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":8.3582600656752703E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.28040510998874E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":3.45227526713165E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.0355184606448499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.29512846956957201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.4674203475463501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":96.490957886081304},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.0656136587113001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0619249190902201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.323449168421726},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":9.6475276569859894E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.99573022116217E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.0489486373854099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":1.3988823832937499E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-5.34045994941432E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":6.2171854376920303E-3}]
["038SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A26,ConcTable!$A:$A,0)-1,0)","v":"19286427-014"},{"$f":"OFFSET(SQ!$D$1,MATCH($A26,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:19:45"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.234},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.79200000000000004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.076000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":56.487214783800198},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-88.406565954217996},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":5.1951799936415801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":312.82109968654902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":187.92511499188001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":8.5863872171037497E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":4.0279608787009098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.10111801268904599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.105326295410994},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":5.7789780936964998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":2.676334462902E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.6198617224039299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.28083615315953703},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":274.81878578115698},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.55981444033149297},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.4408466203047999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.09633561954833E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.2800671514793602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":715.23075313335596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":5.5299253168971801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5885195077746399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.103803981879735},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7794838433956699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.9308665494295096E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.1454752353233599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-5.6003790120762002E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":0.102850066585568},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.39346996429560799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":4.03438595407296E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":94.660604108147396},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.8465021151036202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":7.3657513822569895E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.28136256364519002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":7.8922416050916905E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.5689934487376799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":3.8431232107017099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":7.0279056841255499E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-5.3877789958971896E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.8948255120185701E-3}]
["039SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A27,ConcTable!$A:$A,0)-1,0)","v":"19286427-015"},{"$f":"OFFSET(SQ!$D$1,MATCH($A27,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:21:14"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2639999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.86199999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.452000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":63.5778400523578},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":15.646297550764301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":6.7308575351604301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":327.93060391586698},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":193.82065469075599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.111537309019748},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":4.57099904898636E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.10495618810681701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":6.0805176776318498E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":4.0197539876833099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.8529172107351498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.3384645435586901E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.29330559697439701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":385.43817485738998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.54520784377707798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":4.3160124126236802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":9.6486328527090309E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.20064829124345E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":995.33055514510897},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":0.159642847147559},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.6460330835327499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":9.9955382088793704E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7585091948324501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-6.4914251842484701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.4098208371763599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":6.2247742713692697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.5141406535439601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.26192477809019E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.35384382076486498},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.7498730484774301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":97.538174278575298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.08355993388261E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.3136225033944899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.36032480678668299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":4.5865823774178298E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.6976327241228499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.2140886600366497E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":1.10801239200574E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.3942135465668002E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":5.0531338360874098E-3}]
["040SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A28,ConcTable!$A:$A,0)-1,0)","v":"19286427-016"},{"$f":"OFFSET(SQ!$D$1,MATCH($A28,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:22:47"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.17},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.60199999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.962},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":64.338055460304801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":76.868068087858802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":7.6335130268620004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":385.41584639639501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":224.33958250982201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":6.6777352045271895E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.8221634618199398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.106653330369191},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":6.9643892330284701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.9969849262681598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.49445672905838E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":4.9220869625525202E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.29727130601920398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":507.14327913143399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.61855460539460305},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.2216987226493399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.14923561221535E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":4.5063187438641801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":1040.1174486495199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":9.9096502541444295E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.6889125896286601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.11613449529183301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7146579957332699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.5866152125861601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.14731470018987E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.2099927773813299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.47545603927825E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":7.8897535798586493E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.53243144205895299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.4667040783514101E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":94.690574526026793},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.07307653647256E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":3.18467243124394E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.42315995222007902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":4.4953370632290698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.3473642426840502E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.9271551246526902E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":8.5372148481862502E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-4.76494339540468E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.8909719953279807E-3}]
["041SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A29,ConcTable!$A:$A,0)-1,0)","v":"19286427-017"},{"$f":"OFFSET(SQ!$D$1,MATCH($A29,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:24:17"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.1629999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.82},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":21.594000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":61.645970349747898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-138.958506138053},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":7.9992419686737097},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":349.28878290366401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":191.04958457604101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.124747112431222},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.0080777269534402E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":6.7016751322083595E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.2535187326002403E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.9052331585470599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.4739341326274601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.11850390017698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.39884147029231998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":367.60516944843903},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.51141523492097396},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":5.3443387071382504E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":9.11614179423249E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":8.6123784053575195E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":1000.43041909909},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":0.104269684675025},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.63832982038819},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.107188033235614},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.78157578636596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-6.8700973788234703E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.1642071568717502E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.3357529619286699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":2.9595116449265602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":0.10054947560118099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.484423917930277},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":2.1046339725258999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":96.037420250377707},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.5544363586212892E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":4.25477862286116E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.34392265348295897},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":3.4729826422427897E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.2346392628174402E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":7.87616563113364E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":7.1279779939564E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-1.8611252329355199E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":6.8464854892321602E-3}]
["042SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A30,ConcTable!$A:$A,0)-1,0)","v":"19286427-018"},{"$f":"OFFSET(SQ!$D$1,MATCH($A30,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:25:45"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2610000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.89},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.584},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":58.307881652263497},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-154.321951244579},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":8.4507575660661605},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":340.17320235418401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":213.42422983313699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.18630623365305601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.4954133874775499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":6.8509588528921306E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.0153140559928302E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.71752612484448},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.5918779315000401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":-4.8835732299261801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.57883737183634099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":449.088718298103},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.51786985303475397},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":7.0220104826644999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.4469937136235801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":1065.3321867919101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":1.46081451668502E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.6399898443665},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.13190147990241399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.76958534445178},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-8.1624313672053993E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.0850317004660798E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.64214614872212E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-5.5864554019876998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.9937919854554405E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.55772527777499803},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.24185547174332E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":96.454347173466999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":9.5593446140983702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":8.4296045693354105E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.424816300568153},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":4.5680154041588102E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.1542327683060202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":6.0346635598647398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":2.7885775596816603E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.23610752180716E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.3341770712614897E-3}]
["043SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A31,ConcTable!$A:$A,0)-1,0)","v":"19286427-019"},{"$f":"OFFSET(SQ!$D$1,MATCH($A31,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:27:15"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2309999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.78800000000000003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.384},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":60.6357205059766},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-65.500410782789501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":9.2349848941454695},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":391.96183283406998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":231.35403758936499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.137519159324024},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.5705271361547098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":7.99584578519595E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.14362458698201899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":8.4607420944681699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.7245714452641899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":-5.4538513284960602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.33303500867557001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":534.12191887267898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.61433417851324601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":5.1012363930814601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":5.30512975124235E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.3021616647548201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":1040.90657753255},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.6777943696420299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.72560509798412},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.14151945464963001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.68787937607236},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-4.3903096670689203E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.68959704335846E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-4.9345154445322196E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.72258687215232E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.44474879344177398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":3.8989287137818302E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":93.922775148059301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":9.4933991995828707E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":5.19303638827556E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.40918802838791202},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":4.28732857131182E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.6620216064773699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.5315579178030302E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":2.01901082697429E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.0657624650521799E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.6566481241237593E-3}]
["044SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A32,ConcTable!$A:$A,0)-1,0)","v":"19286427-020"},{"$f":"OFFSET(SQ!$D$1,MATCH($A32,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:28:50"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":3.2189999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":2E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.752},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":22.765999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":55.7427625042709},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-83.585249433385798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.2365670569463501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":361.90665717157998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":221.26724725292701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":8.6274076438011799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.5313150332931701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":9.1050640022922305E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.116993639367174},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":3.4309026291137998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":2.31005451538503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":-1.03397518036099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.34876619829265099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":524.35399851627596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":0.63201110299658103},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":9.6881314738865197E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":2.07791931214295E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-1.15190524097274E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":922.44362454497605},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":9.0236201921961801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.6232006097860601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":0.14109660346219599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.7567825869491001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-4.2747974372973402E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":8.4297242744752203E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":4.8128141803936399E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-4.6308679069931996E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.4271923935281504E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.69799903048675604},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.2672875747036299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":96.166682185520798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":8.5112433133011697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":7.4755243093332501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.39193519679530697},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":3.9783989675278297E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.9486232013948999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.4339657424279103E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":2.13733453336687E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":4.9642342770150103E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.8620057400208996E-3}]
[]
[]
[]
["024SMPL.d","Blank Control1",{"$f":"OFFSET(SQ!$D$1,MATCH($A36,SQ!$C:$C,0)-1,0)","v":"2019-10-22T11:58:42"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.3E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.2809999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.35299999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.15036268865307301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":2518.7331151937201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":4.8422006866449996},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":120.936205129892},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":31.208131511031802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.140345387072493},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.8279876019621202E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.18544622907031399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.18172762913964},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":1.5511288539275701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":4.6055492089780601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":9.9825361812703906E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.107151520405879},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.40350090951414402},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":3.9459117120023597E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":2.2107971129237201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":5.6820531075426799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":4.4557379315151302E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":18.552606316371101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.9532263687237503E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":7.4251240425254397E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":8.9832144026047601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":1.5300607729219701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":8.4817288355203101E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":4.26604573543061E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.5689378859270301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.73988722856479E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":0.76956566717627695},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":6.4134136960588401E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":9.9916376125462805E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":3.9736076352837798E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.64679774706176996},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":2.19848512955208E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.19108777671899E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.5593519418243899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":2.17209290279733E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":1.0537742244832299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.3494943377644405E-5}]
["048SMPL.d","Blank Control2",{"$f":"OFFSET(SQ!$D$1,MATCH($A37,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:34:53"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.524},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.39},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-358.70099849266001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":0.42131315942116598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-4.0835241863831397},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-12.320416542132699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":1.8689508567091999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.1304574497763301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.3072382505052997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":-5.5540944025919397E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.35692561536431},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.7291371634993799E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.0939054802140901E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-3.9899789419343898E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.218347688796206},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":2.3485399390022698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-7.6941632106343298E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-3.87260188280653E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":4.0500945146277501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":15.2680514187578},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.0240800869242801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-3.6346593792729501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-4.7190520436239903E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":9.3452316447226597E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":1.07729225926167E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.3362286689054201E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-9.8353686608873697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-3.6938344109737101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.47405699542373397},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":3.6902404067675398E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":2.77862403014527E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.01758829446069E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.45754624908168601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":1.43817138539697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":4.9770105088035699E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.14044240674632E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-8.8776360046398405E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":-4.5724610687465703E-6}]

{"sheet":"Summary1","index":5,"n_rows":50,"n_cols":6,"rows_kept":50,"cut":false,"hidden":false,"merged":[[0,0,0,5]]}
[{"$f":"INDIRECT(A50&\"!A5\")","v":"Green top"}]
["Element","Result","Average (ppb)","Standard Deviation","Average + 2SD","Reporting Limit"]
["Lithium",{"$f":"OFFSET(INDIRECT($A$50&\"!$A$10\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":"Pass"},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$7\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":3.2073684210526316},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$8\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":5.8522750112257653E-2},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$9\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":3.3244139212771469},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$2\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":10}]
["Beryllium",{"$f":"OFFSET(INDIRECT($A$50&\"!$A$10\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":"Pass"},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$7\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":1.578947368421053E-3},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$8\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":9.01590537370498E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$9\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":3.382128443162049E-3},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$2\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":0.25}]
["Boron","Pass",-0.764421052631579,9.5157364279428794E-2,-0.57410632407272144,25]
["Aluminium","Review",21.634789473684208,0.96478187396296211,23.564353221610133,10]
["Magnesium","Pass",62.195009894886333,4.3119988895741947,70.819007674034722,200]
["Silicon","Review",48.015646237661699,102.51681218633205,253.04927061032578,500]
["Phosphorus","Pass",8.3713377652508285,3.4416630223435964,15.254663809938021,500]
["Potassium","Pass",332.50412253534085,30.051773811984205,392.60767015930924,500]
["Calcium","Pass",201.40623368589667,15.492224572056918,232.39068283001049,500]
["Titanium","Pass",9.6042825196524476E-2,3.772514147103706E-2,0.17149310813859858,5]
["Vanadium","Pass",5.5575510547624808E-2,1.1107168176518157E-2,7.7789846900661125E-2,0.5]
["Chromium","Pass",8.9425273890718734E-2,3.5562845208708102E-2,0.16055096430813492,0.5]
["Manganese","Pass",8.4558083079880886E-2,4.6860505645854575E-2,0.17827909437159004,0.5]
["Iron","Pass",5.352522710816416,4.0302607938678774,13.413044298552171,20]
["Cobalt","Pass",3.5110759571943019E-3,3.6923727039200605E-3,1.0895821365034423E-2,0.25]
["Nickel","Pass",6.2807253327921919E-2,0.18532074675635982,0.43344874684064155,1]
["Copper","Pass",0.3571961550009955,8.5700559447023775E-2,0.52859727389504307,5]
["Zinc","Review",349.03721390375193,89.908944411773561,528.85510272729903,50]
["Gallium","Review",0.57746446188485889,4.8036244629323196E-2,0.67353695114350531,0.5]
["Germanium","Pass",4.3781352908288046E-3,5.3604648747454372E-3,1.509906504031968E-2,0.5]
["Arsenic","Pass",1.1947423351617584E-2,7.9401335627241783E-3,2.7827690477065941E-2,2.5]
["Selenium","Pass",4.2700192630651399E-3,3.4752456995766512E-2,7.3774933254598163E-2,10]
["Bromine","Review",852.64831087747791,136.92125890453229,1126.4908286865425,250]
["Rubidium","Pass",7.1579681427747344E-2,4.4722422836897632E-2,0.16102452710154261,25]
["Strontium","Pass",1.6129520746104635,7.4368854774424606E-2,1.7616897841593127,5]
["Zirconium","Pass",0.11245642392191342,1.2896449124785912E-2,0.13824932217148525,2]
["Molybdenum","Review",1.7710077462553884,4.0711938129533998E-2,1.8524316225144564,0.5]
["Palladium","Pass",-5.1034920359463443E-3,1.9107413454635055E-3,-1.2820093450193332E-3,0.5]
["Silver","Pass",3.7332017245677652E-4,4.1251700287726011E-4,1.1983541782112968E-3,1]
["Cadmium","Pass",1.4320324413975164E-3,1.2239241872789442E-3,3.8798808159554049E-3,0.25]
["Indium","Pass",0,0,0,0.25]
["Tin","Pass",-6.9138397078227743E-4,8.7100597204155267E-3,1.6728735470048776E-2,1]
["Antimony","Pass",8.9436717017536332E-2,7.6166219815629764E-3,0.10466996098066228,0.5]
["Tellurium","Pass",0,0,0,0.5]
["Iodine","Pass",-0.37596337503866761,9.9788239986596514E-2,-0.17638689506547459,10]
["Caesium","Pass",1.2721139325932722E-3,7.7508140711810166E-4,2.8222767468294753E-3,0.5]
["Barium","Review",96.250758635617572,1.7729158583326601,99.796590352282891,2.5]
["Tungsten","Pass",9.194459601868021E-3,1.1318741448071296E-3,1.145820789148228E-2,0.5]
["Platinum","Pass",7.4398169759386928E-4,3.9710093362704953E-4,1.5381835648479684E-3,0.5]
["Gold","Pass",5.3057963312404699E-2,0.74501939734070044,1.5430967579938055,5]
["Mercury","Pass",6.1084747234449774E-2,1.9837271876044588E-2,0.10075929098653895,0.5]
["Thallium","Pass",2.066219423528976E-3,2.9645538618713771E-4,2.6591301959032515E-3,0.25]
["Lead","Pass",5.0352104485931068E-2,9.3984972009562885E-3,6.9149098887843652E-2,2.5]
["Bismuth","Pass",4.9818609634038637E-4,3.8480490334369107E-4,1.2677959030277684E-3,0.25]
["Thorium","Pass",-2.1468185126943031E-4,2.7966955941086493E-4,3.4465726755229955E-4,0.25]
["Uranium","Pass",6.035888497458005E-3,1.632584520991366E-3,9.3010575394407366E-3,0.05]
[]
["Container1"]

{"sheet":"Instructions","index":6,"n_rows":39,"n_cols":3,"rows_kept":39,"cut":false,"hidden":false,"merged":[]}
[]
[null,"In Masshunter activate the SemiQuant Tab"]
[null,"Select View1 settings"]
[null,"Highlight all rows (click empty square in top left of results table)"]
["close this workbook before running the script in the next line"]
[null,"From the tool menu select 'Container' and wait 20 seconds"]
[]
[null,"This workbook will now be open"]
[null,"Import results using the button below"]
[]
[]
[]
[]
[]
[null,"From the 'SQ' tab, paste the data file names for QCs into the yellow cells of the 'QCReview' tab"]
[]
[null,"If more than 1 container type is being used then enter the number into the yellow cell below and click the duplicate button"]
[]
[null,3,"How many container types?"]
[]
[]
[]
[]
[]
[null,"Paste Data file names for containers into each 'Container#' tab"]
[null,"There should be 20 replicates of each, but in a special circumstance where less than 20 replicates were performed,"]
[null,"enter the number of replicates in the A7 cell"]
[null,"Enter name of container type in \"A5\" cell for each 'Container#' tab"]
[]
[null,"Containers are acceptable for use if the average + 2 standard deviations is lower than the reporting limit and all QC values"]
[]
[null,"Notes for interpretation:"]
[null,"The full quant results contain relevant results for light elements Li, Be, B, and Al"]
[null,"Those results are imported into the FQ tab in a long file format"]
[null,"The semi quant results contain relevant results for the remaining elements and these are in the SQ tab in a wide format"]
[null,"The 'ConcTable' tab combines all relevant results into a table (wide format)"]
[null,"The 'QCReview' tab references the 'Conc Table' tab for the purpose of reviewing QCs"]
[null,"Each 'Container#' tab references the 'ConcTable' and the QC values from the 'QCReview' tab"]
[null,"A simple summary of all results for each container can be viewed on each 'Summary#' tab"]

{"sheet":"Container2","index":7,"n_rows":37,"n_cols":49,"rows_kept":37,"cut":false,"hidden":false,"merged":[]}
[null,"Name","Acq. Date-Time","Li_7","Be_9","B_11","Al_27","24  Mg","28  Si","31  P","39  K","44  Ca","49  Ti","51  V","52  Cr","55  Mn","56  Fe","59  Co","60  Ni","65  Cu","66  Zn","71  Ga","74  Ge","75  As","78  Se","79  Br","85  Rb","88  Sr","90  Zr","98  Mo","105  Pd","109  Ag","114  Cd","115  In","118  Sn","121  Sb","125  Te","127  I","133  Cs","137  Ba","184  W","195  Pt","197  Au","202  Hg","205  Tl","208  Pb","209  Bi","232  Th","238  U"]
[null,null,"RL",10,0.25,25,10,200,500,500,500,500,5,0.5,0.5,0.5,20,0.25,1,5,50,0.5,0.5,2.5,10,250,25,5,2,0.5,0.5,1,0.25,0.25,1,0.5,0.5,10,0.5,2.5,0.5,0.5,5,0.5,0.25,2.5,0.25,0.25,0.05]
[null,null,"Low QC",{"$f":"MIN(QCReview!D26:D100)*D2","v":8.5139999999999993},{"$f":"MIN(QCReview!E26:E100)*E2","v":0.21099999999999999},{"$f":"MIN(QCReview!F26:F100)*F2","v":19.602},{"$f":"MIN(QCReview!G26:G100)*G2","v":8.9},{"$f":"MIN(QCReview!H26:H100)*H2","v":180.71063948691099},{"$f":"MIN(QCReview!I26:I100)*I2","v":150.581516610791},{"$f":"MIN(QCReview!J26:J100)*J2","v":449.659182698216},{"$f":"MIN(QCReview!K26:K100)*K2","v":509.50512352205499},{"$f":"MIN(QCReview!L26:L100)*L2","v":378.75559869518997},{"$f":"MIN(QCReview!M26:M100)*M2","v":4.6648107431519001},{"$f":"MIN(QCReview!N26:N100)*N2","v":0.44575373577775101},{"$f":"MIN(QCReview!O26:O100)*O2","v":0.44711397777342698},{"$f":"MIN(QCReview!P26:P100)*P2","v":0.45988192477422901},{"$f":"MIN(QCReview!Q26:Q100)*Q2","v":17.709543709225699},{"$f":"MIN(QCReview!R26:R100)*R2","v":0.20697248682051},{"$f":"MIN(QCReview!S26:S100)*S2","v":0.88745615928769706},{"$f":"MIN(QCReview!T26:T100)*T2","v":4.5006272265695602},{"$f":"MIN(QCReview!U26:U100)*U2","v":42.7624616138785},{"$f":"MIN(QCReview!V26:V100)*V2","v":0.44235442630143601},{"$f":"MIN(QCReview!W26:W100)*W2","v":0.56798150017591098},{"$f":"MIN(QCReview!X26:X100)*X2","v":2.2798906623553399},{"$f":"MIN(QCReview!Y26:Y100)*Y2","v":7.9895749025331693},{"$f":"MIN(QCReview!Z26:Z100)*Z2","v":279.40181851843801},{"$f":"MIN(QCReview!AA26:AA100)*AA2","v":21.224836651076298},{"$f":"MIN(QCReview!AB26:AB100)*AB2","v":4.5993333721576004},{"$f":"MIN(QCReview!AC26:AC100)*AC2","v":2.4624686572192398},{"$f":"MIN(QCReview!AD26:AD100)*AD2","v":0.65732553977758001},{"$f":"MIN(QCReview!AE26:AE100)*AE2","v":0.544577383656354},{"$f":"MIN(QCReview!AF26:AF100)*AF2","v":0.81312160603667305},{"$f":"MIN(QCReview!AG26:AG100)*AG2","v":0.210932860609923},{"$f":"MIN(QCReview!AH26:AH100)*AH2","v":0.211447995002015},{"$f":"MIN(QCReview!AI26:AI100)*AI2","v":0.98118408712764604},{"$f":"MIN(QCReview!AJ26:AJ100)*AJ2","v":0.47512606879184399},{"$f":"MIN(QCReview!AK26:AK100)*AK2","v":0.36104432002066},{"$f":"MIN(QCReview!AL26:AL100)*AL2","v":22.033046859313501},{"$f":"MIN(QCReview!AM26:AM100)*AM2","v":0.45216308043174602},{"$f":"MIN(QCReview!AN26:AN100)*AN2","v":2.1326937795210901},{"$f":"MIN(QCReview!AO26:AO100)*AO2","v":0.40638711024189},{"$f":"MIN(QCReview!AP26:AP100)*AP2","v":0.39720220550103602},{"$f":"MIN(QCReview!AQ26:AQ100)*AQ2","v":9.04952796357958},{"$f":"MIN(QCReview!AR26:AR100)*AR2","v":0.50778669339083604},{"$f":"MIN(QCReview!AS26:AS100)*AS2","v":0.18995279692808301},{"$f":"MIN(QCReview!AT26:AT100)*AT2","v":1.9855488339795799},{"$f":"MIN(QCReview!AU26:AU100)*AU2","v":0.237213785892985},{"$f":"MIN(QCReview!AV26:AV100)*AV2","v":0.20103154901252401},{"$f":"MIN(QCReview!AW26:AW100)*AW2","v":4.0606714830781203E-2}]
["Container Name",null,null,"Lithium","Beryllium","Boron","Aluminium","Magnesium","Silicon","Phosphorus","Potassium","Calcium","Titanium","Vanadium","Chromium","Manganese","Iron","Cobalt","Nickel","Copper","Zinc","Gallium","Germanium","Arsenic","Selenium","Bromine","Rubidium","Strontium","Zirconium","Molybdenum","Palladium","Silver","Cadmium","Indium","Tin","Antimony","Tellurium","Iodine","Caesium","Barium","Tungsten","Platinum","Gold","Mercury","Thallium","Lead","Bismuth","Thorium","Uranium"]
["acid washed 10 mL"]
["Number of containers"]
[20,null,"Average",{"$f":"AVERAGE(OFFSET(D$12,0,0,$A$7,1))","v":-1.8736842105263159E-2},{"$f":"AVERAGE(OFFSET(E$12,0,0,$A$7,1))","v":1.0526315789473685E-4},-1.3235263157894737,2.2755263157894738,0.57779839834678215,-137.41852289998971,2.9908347801601599E-2,-6.90393418339292,-0.35220547206186081,-1.3513087294231955E-2,8.1328230015633484E-5,1.8293695988449672E-2,4.2592207625348745E-2,0.18403887057730073,-5.8347893426083234E-4,5.1570795675138358E-2,7.3589783428011646E-3,-1.1567740230399753E-2,-4.9380002720002811E-4,-9.5751740064903713E-4,1.5681559479164414E-5,-2.7071023218607444E-2,-6.431128423651514,1.0195221297253193E-2,-3.7353252213692088E-3,8.2129587809565412E-4,-3.1578815867300859E-3,-2.4847997352979607E-3,-2.2847814088326777E-2,2.408563090253719E-4,0,3.327248830340183E-3,8.5952878241022388E-4,3.3016388798718526E-4,-0.25834990699794941,8.0511180854246508E-6,-2.5825592456599337E-3,-1.1173840531912408E-3,7.7905589309486105E-5,-9.9450607999082188E-2,-2.4149828338544409E-2,-1.286901895364591E-5,-2.2911361861198197E-3,-3.074091127379561E-3,-2.8317582585904177E-2,2.3815924467229106E-5]
[null,null,"SD",{"$f":"STDEV(OFFSET(D$12,0,0,$A$7,1))","v":2.7047716121114946E-3},{"$f":"STDEV(OFFSET(E$12,0,0,$A$7,1))","v":3.1530176764230577E-4},0.21760498575299617,3.6001885809931475,0.75283555543704506,115.55372806151925,2.3405394467337381,4.9556854720123553,3.6348246977342762,3.0862538704661317E-2,9.6765361465759574E-3,2.5769511475327824E-2,2.8434520979437248E-2,0.35061943087619968,9.6934768980146656E-4,3.9964528209102693E-2,1.7945127725241063E-2,0.12605053909025299,4.5555940157618541E-3,5.9441844705509861E-3,1.630727523235388E-3,2.9127635064189968E-2,2.1711172905202725,2.4546894296741898E-2,4.4049049514727238E-3,1.947695521300321E-3,1.0777931335002604E-3,1.2521701317617026E-3,6.0167364980755322E-3,3.4963051524095697E-4,0,5.4201805905998435E-3,1.927719385556622E-3,1.4391510225426232E-3,0.2086354023872096,5.6250628775741636E-4,6.5116018928334735E-3,3.1780389876500737E-4,1.7209981676633303E-4,4.1379439405970019E-2,5.598440752665022E-3,1.1474254730317303E-4,2.2025472543715828E-3,4.0650218208467389E-3,4.5962289816614014E-3,3.9572576969886147E-5]
[null,null,"ave+2sd",{"$f":"D7+D8*2","v":-1.3327298881040169E-2},{"$f":"E7+E8*2","v":7.358666931793484E-4},-0.88831634428348138,9.4759034777757698,2.0834695092208722,93.688933223048792,4.7109872412690779,3.0074367606317907,6.917443923406692,4.8211990115090683E-2,1.9434400523167548E-2,6.9832718939105323E-2,9.9461249584223241E-2,0.88527773232970008,1.3552164453421007E-3,0.13149985209334375,4.3249233793283291E-2,0.24053333795010623,8.6173880043236804E-3,1.0930851540452934E-2,3.2771366059499403E-3,3.1184246909772492E-2,-2.0888938426109691,5.9289009890736988E-2,5.0744846815762388E-3,4.7166869206962964E-3,-1.0022953197295651E-3,1.954052822544447E-5,-1.0814341092175713E-2,9.4011733950728586E-4,0,1.416761001153987E-2,4.714967553523468E-3,3.2084659330724316E-3,0.15892089777646978,1.1330636936002574E-3,1.0440644540007014E-2,-4.8177625566122608E-4,4.2210522284215216E-4,-1.6691729187142151E-2,-1.2952946833214365E-2,2.1661607565270016E-4,2.113958322623346E-3,5.0559525143139168E-3,-1.9125124622581376E-2,1.029610784070014E-4]
[null,null,"Pass/Review",{"$f":"IF(D9>MIN(D2:D3),\"Review\",\"Pass\")","v":"Pass"},{"$f":"IF(E9>MIN(E2:E3),\"Review\",\"Pass\")","v":"Pass"},"Pass","Review","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass"]
[]
["Containers"]
["056SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A13,ConcTable!$A:$A,0)-1,0)","v":"19286426-001"},{"$f":"OFFSET(SQ!$D$1,MATCH($A13,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:46:51"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.2E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.66500000000000004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.14699999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":40.931436531421198},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.62733854132069},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-6.5100114602574504},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.2050022898977599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":2.6894468618009398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":2.9931103964862299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.1923000960916102E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.38488840290565E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.23776311292553001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":6.5029074597298801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.7818699055409198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-6.6097789649473401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":4.4715525326841803E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":1.67770593172858E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-3.3973587764041498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-1.97613077750061E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-3.4601439211053302},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":2.57898369601091E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-5.69793584796643E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.3878076489210901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.4022953328407901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-6.2432482713328395E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.16430346228544E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":3.39957918179954E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.7167188677014198E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.1556990611150599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":0.36007186909311001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.0623725972995601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-9.3884738980283197E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-4.6373139868205801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":2.7415549145091402E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.2874697473503401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.36468803726893E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.8516222490281998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-2.69705155421897E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-1.20542352858952E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["057SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A14,ConcTable!$A:$A,0)-1,0)","v":"19286426-002"},{"$f":"OFFSET(SQ!$D$1,MATCH($A14,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:48:21"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.4E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.97199999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.34300000000000003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.156829179004355},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":83.501196534003199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-0.18798500070090099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-4.0807461533227203},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-5.2449367247137104},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.5074037146459403E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":3.6176120880817E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.7740591384102E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.5625476506253398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.24111093314125401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.8651842334056399E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.6911652126535101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.4481976750854301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-6.1852101256193101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":3.53146926644907E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-5.4310921048946297E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.79084168705515E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.5219657376365202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-2.2935067509470102},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":8.9986436088454105E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.80260514346303E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-4.2112953301473398E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.0577118017894298E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.9421906326296001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-7.3295263004024697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":7.5424643187473799E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.3761477310425995E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.169250135344166},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-6.3411865180481501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-5.9274594723134396E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-7.8764535778008899E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-4.4892819626929897E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.31317940817651E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.9444643140747902E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.1958049798510601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.2173876675242398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.23696536858316E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["058SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A15,ConcTable!$A:$A,0)-1,0)","v":"19286426-003"},{"$f":"OFFSET(SQ!$D$1,MATCH($A15,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:49:51"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.6E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.085},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.48399999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":1.01775032914504},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-69.898320304579102},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-2.9434270314094801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-6.8269236767893302},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":1.5843288584170501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":5.4908382569827403E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-8.8327992314239708E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":-2.3516849611446702E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.6086776345952901E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.55216726910575098},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.13925896188029E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.0377747991117899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":1.19171546311938E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-8.6247868384080301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-5.1551215216098503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":6.7286683389814602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-5.5488055443014899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":7.7643618598775598E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-2.9707286289431399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-1.3318108901493699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.2821267330364599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.5364656894546799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.0682520310518399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.48462207435896E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.1176054501460899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":6.2731138717565197E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.16982196515511},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.1738262489110399E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":6.7430702525778097E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-7.8855266375631096E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-5.3748918135840998E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.8519630807241401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-1.17695107384018E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.42483388297394E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":1.35259474592843E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.5238018807107999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["059SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A16,ConcTable!$A:$A,0)-1,0)","v":"19286426-004"},{"$f":"OFFSET(SQ!$D$1,MATCH($A16,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:51:21"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.4999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.248},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":1.2589999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":2.0424807238025098},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-110.617407709334},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-1.1716394990351899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-5.5059748262398802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.9840302647682599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.1235351654538107E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.7899057366056199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":-2.92974912495858E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.7587895346219303E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.13700399442049699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.24299602668996E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.8158198707270597E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.3068436505789901E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":1.8062742920763899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":1.32814042946602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.07580070605872E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-4.3329478170536201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-3.8784705711633398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.0956392946013602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-5.7818915642756201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":9.0807018952387204E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.5512096023245601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.5251726393568401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.6633085426632301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.8411683047797099E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":4.9881732412727503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.5838558985527801E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.13719418307745601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":4.3435525034208998E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":3.0678765687231898E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-8.5372363181579101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0691497317163099E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.13165732733739299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.6606183185292601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.7519012417714199E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.10275139350303E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-2.40067049658015E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.6561867273503799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["060SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A17,ConcTable!$A:$A,0)-1,0)","v":"19286426-005"},{"$f":"OFFSET(SQ!$D$1,MATCH($A17,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:52:55"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.9E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.226},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.94399999999999995},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-76.525089045549706},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-2.8516995484908598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-13.386157681103301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":5.1623464074139003},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-2.5581558814900302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":5.6751805315864304E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.12643629094331E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":9.51488627293187E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":5.7133004210687698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.16435230757484E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.54002521789946E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.5130845751356601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":3.8009951259526699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":7.9664295561296707E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.71851905972076E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-6.5081637156821595E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-6.2915723765259601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.5959757454041001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-5.7774209288806097E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-8.7203489119341403E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.1858623664500802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.5684105645734801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.80549792462441E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":2.37524238011501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":8.10905485824674E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-3.5059592278733502E-7},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.35506262231453101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":2.10107487125978E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.8830392025981202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-9.3856560740613302E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-5.0601460994735802E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.7851332277234899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.8627317281723999E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-4.1683565989633501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.2749444000563797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.7136355397538101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.2144582851602599E-5}]
["061SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A18,ConcTable!$A:$A,0)-1,0)","v":"19286426-006"},{"$f":"OFFSET(SQ!$D$1,MATCH($A18,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:54:24"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.7999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.302},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":1.526},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.760236281691402},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-68.3183416208756},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.47485627517565},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-7.3559174801228302},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-4.1220054569305598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-5.01463943678317E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":9.4549932567033999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.81266439490935E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.58915834169957E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.15043216189403799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.39321685610857E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":8.4452036297394006E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":1.5817713927441E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-0.117050101393971},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-5.0392172138836603E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.3482600960083E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-6.5584408337568298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":2.5208343102980101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.3227340966136601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":3.1410657256581102E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-1.2856177567077401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.5299417179133802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.13871824222495E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":2.5529175549324603E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.9369915405294403E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":7.4357405820243798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-9.3648646377815697E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":6.3416564096460595E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-8.8470543911524802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-5.4556899954364899E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.106143926992859},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.4320399125142701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.1774017139406599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.59950435967317E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.9884266637212002E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.8510512285248001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["062SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A19,ConcTable!$A:$A,0)-1,0)","v":"19286426-007"},{"$f":"OFFSET(SQ!$D$1,MATCH($A19,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:55:52"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.3129999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":2.7080000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.74899526561985297},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-103.20958654142601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":2.5723047731001101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-7.4634995119967602},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":7.1817849800448998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.3138679060120399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-1.2715552717453799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.10976497124102801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.129420405484486},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":1.28737284007},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":9.9505544323867895E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":0.16699504542369401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":5.1720099920434401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.20922609833151101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":1.3569124606410601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":9.6832743520774297E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.8689053982021901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-6.5319091741921104E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-7.1306696230444198},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.0084554613012503E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-7.1440550747193396E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":4.3381278296259797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-1.0253156790188201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.85814368163647E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.5765592634744101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.3683816885506901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":6.6400862750498196E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.0878698604444401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.47687390527718498},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":5.9952868413603796E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":1.62165267060361E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.4105891474400999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.105162127388673},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.7174251015931301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-1.1830577676428201E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":2.79112276931747E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.86473584403366E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.9441427131799501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["063SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A20,ConcTable!$A:$A,0)-1,0)","v":"19286426-008"},{"$f":"OFFSET(SQ!$D$1,MATCH($A20,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:57:21"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.7000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.3819999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":1.139},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.46082358635628301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-14.789031327780799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-4.0147260855427902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-7.1540635054894999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":0.64995578515212304},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-6.1804741050557201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.16085853002375E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.6288091986637299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.6898611475156997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.151741492396407},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.8262637327984899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":7.4400901612495698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":1.0559756196997901E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-8.0997837394790301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-7.66214461864702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":4.5799545059044197E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.1579627024928402E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-4.92646086507607},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":1.7220836707189401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-7.2052664729299601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-2.19815477903161E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.1718489612472698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.59854242777433E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.8111974740462502E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":2.0551110479311301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-8.5387902830085698E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.37793879258756102},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-8.4093129215789602E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.40725843031943E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0620346232599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.12179158781199401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.4855675516584502E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.17949161696387E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.9902052038234097E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.2021499017629504E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.9142940309322199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["064SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A21,ConcTable!$A:$A,0)-1,0)","v":"19286426-009"},{"$f":"OFFSET(SQ!$D$1,MATCH($A21,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:58:55"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.36},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":1.0680000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.46561168732029201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-65.699115471445594},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":2.25211443064756},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-2.6714055283828202},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.91285326688747},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-4.4559112967322801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-8.3420505805847697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.24808468170487E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.4925699463072399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.17542021075971501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.2130009859051701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":6.8352091007377394E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-5.25463044588108E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":2.8355914886362198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-9.3653715115524596E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":5.9246915214927996E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":1.5316953547642401E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-7.6194926854637002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.3193307897595699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-7.1358314078349E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.8901574316485001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.1326917579419401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.1905413780560698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.1582620097078301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":2.33883908944823E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.80668637162259E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-8.3866134913015001E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.22802483474938001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.93274985613688E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-7.8257008605326101E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.4842695346867199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0808663040592001E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.116404601861646},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.35885744185385E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":6.3364371585016105E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.9608767021449302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.4101000998318998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.9670300293798899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.3830221146173099E-5}]
["065SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A22,ConcTable!$A:$A,0)-1,0)","v":"19286426-010"},{"$f":"OFFSET(SQ!$D$1,MATCH($A22,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:00:23"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.371},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":16.254000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-140.24662886139001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":0.74340012927380394},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-5.8016203597834597},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-6.8903679597884293E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.04198032815289E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-6.7153493220310197E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":7.1431657746650898E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":1.9447536971464702E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.27038052447655803},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-2.3633262693556001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":3.1797252235001502E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":4.1719434085151802E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":9.0138647022067406E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":6.6320177471237696E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-5.7239164357813302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.8287588486283399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-4.3090787823456697E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-4.2896663011826703},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":3.2364998052594702E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.38360005612561E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":3.1147030601966601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.5048439061731197E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.24597146022117E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.2663037988181201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.7379552802921001E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.1386657491701899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.2850824995313602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.49689248921856799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-8.4134429848171596E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":4.9244806163176403E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.2510702324114E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.11327092825150301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-3.0826403366021399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.0047212819483201E-7},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":2.6966438386480801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.7022820859217397E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.01871570323912E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":1.23312473399421E-4}]
["066SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A23,ConcTable!$A:$A,0)-1,0)","v":"19286426-011"},{"$f":"OFFSET(SQ!$D$1,MATCH($A23,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:01:53"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.375},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":2.0619999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.60282224154328401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-119.914427325017},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":4.4433468856183298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":0.68488175681207497},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-2.58991605464356},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-1.4170938953389599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-1.1755985788824599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.85637197616196E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":7.4047949742315697E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-8.7941538459466208E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.36635885334298E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":6.5632803826223399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-9.0579966267413797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.161416852211101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":2.77752740731156E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-2.6815317711616298E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":5.6476526811472198E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-4.2964287688720099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-7.0775172754493099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":3.8482471663735902E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.5616487586259901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":2.6724956583619001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.3380951191996702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.2115294110476702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.68780736819757E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":9.0544610847182599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":3.9484577898933398E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":3.4165884962744199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.47360378818073001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":2.0849629006687901E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-2.95981062606192E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.4859738243281999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":7.40446754546393E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.11698127106152},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.7052773463793499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.8776991812948103E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.91590295206914E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.1332983406930602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0134483342374899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.1019762150678202E-5}]
["067SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A24,ConcTable!$A:$A,0)-1,0)","v":"19286426-012"},{"$f":"OFFSET(SQ!$D$1,MATCH($A24,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:03:23"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.3939999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":1.1850000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.30650675582356002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-137.994185808997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-1.55136221961427},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-7.3655836208841698},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":1.6387204889060301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.7082617181093999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.0091096720182401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.51817196016197E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.60884524371753E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-5.9650714707744697E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-6.1246509982186396E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":9.5511707858797895E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-4.1036977825444597E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.238011531201237},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.0276717413783901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":4.6371624734710204E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-6.4735340818618106E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-8.2410636601880896},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-7.3361568266781E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-2.7226367767162899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":3.1900915077375299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.7263704282824799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.2538626892640301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.5520384489605499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":3.7114857279018298E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.3867715003163701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.4595377479449201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.30396242801794998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.0738059391357899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-4.8646725375306602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.1689610818715801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.142118678402917},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.5329406558501699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-1.16594259886363E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.8365181767424901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.2696324923113997E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0862355259846801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["068SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A25,ConcTable!$A:$A,0)-1,0)","v":"19286426-013"},{"$f":"OFFSET(SQ!$D$1,MATCH($A25,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:04:56"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.444},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.749},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.43469831968887801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-293.66738732201901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-1.7867934039626301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-13.872003444314},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-5.1164152459869596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-2.4363384793711099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":3.2986134611166398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.01833432745887E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.7036675935696498E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-3.0181144145649399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-7.5028166897307897E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":2.1962189110827399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":1.3630372598170201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-0.171461784153807},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.3369141989615699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-3.05349037226412E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-2.6013694010272898E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-5.5692928610012604},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":1.5761346297465399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-4.4195631586534702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":8.5527346599234197E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-5.2649764935650597E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.53605362439056E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.77555238602476E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.9594963688157399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-2.9633207534335602E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.192993489986514},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-4.2559022808995903E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.9051402782163603E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.33125799534978E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0511949351100099E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.105539196124048},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.76331461374042E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.9505333970677597E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-4.0932527085190702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.0665853382798097E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0892666194180601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.1528958343013002E-5}]
["069SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A26,ConcTable!$A:$A,0)-1,0)","v":"19286426-014"},{"$f":"OFFSET(SQ!$D$1,MATCH($A26,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:06:26"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.4790000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.59},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.59091761551120903},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-197.24843853780001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":7.6889489317300899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-9.5696283910300703},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-0.67969646774262704},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.5711078451230797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-7.9393383566158694E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.6167176141092501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.32647515543349E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-1.5655002084477E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.1219012410213802E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.5221185135233199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":2.61875165560186E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":3.5378648125250399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":3.8968579773675802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-3.8863521380381998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":5.5683980392107297E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":7.3399429065673595E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-10.6308025157174},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.6216628963355701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-8.7405197784197908E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":2.1821726454720699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.2306397667641001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.5819853207174298E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.2351531975500501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":9.4186509979829494E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.08113155306303E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-1.48286242197377E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.24037063448790899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.17654678946264E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.9079856633922597E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.02138221994895E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.131585624671298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.8711733310004699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-1.17714075146749E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.2009585021870999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.2791621734742302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0494511517217101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["070SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A27,ConcTable!$A:$A,0)-1,0)","v":"19286426-015"},{"$f":"OFFSET(SQ!$D$1,MATCH($A27,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:07:55"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.476},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.36699999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.30080649938221699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-204.93133536206099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-1.6317325486530201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-5.6208351338937499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":4.2251566134132101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":9.7936771384070608E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-9.8048729080406306E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.65081460275637E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.4565700201501599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-2.62044895718423E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-7.4062404098164901E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":7.8441641903820802E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.30750133674061E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-0.20524360227247401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-7.6991196980278499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":9.7236601682690002E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.85673434102852E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-6.5479132611223997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-10.6730429892161},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.6515108930045201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-7.35058928865306E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-2.21336983983548E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.9003628553673998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.5535889047850302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.9142811851389401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.3602698777993901E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-5.9941475554245899E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.50278588958024995},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-4.4282720601333998E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-5.0881331097740396E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.33766578567205E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.10546181462802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.88542361832979E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.8007687558300699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.4985596809511003E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.1383025500431101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["071SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A28,ConcTable!$A:$A,0)-1,0)","v":"19286426-016"},{"$f":"OFFSET(SQ!$D$1,MATCH($A28,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:09:24"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.9E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.508},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.91700000000000004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-281.47544405334799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-1.80552343609646},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-14.877706407946301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-6.5247564934269304},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-4.2541666244788903E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-1.10616304029982E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.4064323829496701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":1.7855277804864601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-0.233899305234994},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-2.2198022850435399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.9960150254132601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":7.60773507220698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-2.1925525194049299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.4746745869969799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-4.8729080631059298E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-6.5975713857224505E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-7.9120929694571203},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-3.2609038406376599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-3.20748706830699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":3.2812152043393402E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.0188229122295697E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.5162721550133198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":4.5402357792331903E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-1.0900436350563499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-6.1941935431882402E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.39078558534653601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-8.8908141951543809E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.41433109146785E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.13311886267510201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.6306951503706499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-7.5044590203584304E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-4.3141937560472396E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.3643356385051303E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0895967843570599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["072SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A29,ConcTable!$A:$A,0)-1,0)","v":"19286426-017"},{"$f":"OFFSET(SQ!$D$1,MATCH($A29,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:10:59"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.4890000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":4.343},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.15100878618063801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-216.149891022719},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":3.8562853550188301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":2.3176845352248798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-0.82201875963400906},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-1.39906348965352E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.38810936746625E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":9.0615640025609508E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":6.7551678483896593E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.57605679374248497},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.4299502882570699E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":2.4147871379164899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":1.9872115481993999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-1.35098821531237E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-7.6820959344048902E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-7.7031870752383696E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.8289706628246299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-6.4558656718680698},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":5.7018372520199201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-7.2665642285394396E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-3.1708037328257798E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.6456653660070699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.94328134863951E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":-1.82692996761087E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.1328669675700299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.6521633631717602E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.24096525023062201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.06018213960042E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-3.0189859111316802E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.3318170073350899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":2.0973468914928101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.114458656104776},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-3.03255737462029E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-1.17797121357317E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.0654718219375E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.7861658646661102E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0953670825216099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.0666566986465102E-5}]
["073SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A30,ConcTable!$A:$A,0)-1,0)","v":"19286426-018"},{"$f":"OFFSET(SQ!$D$1,MATCH($A30,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:12:28"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.526},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":4.1139999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":2.9386822975193398},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-322.54507994164101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":0.63618521778449699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-15.263643987612101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":2.8735437542697202},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-1.7027802606029398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-9.9127491224868195E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.7332276609676203E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.2118738175292401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.345349552024107},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":2.7661583879939498E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":2.0817397414501199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.5238660115122499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-0.102878566737324},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":2.5351845452156502E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-2.0258479395680899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.7680636049779201E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-6.4570358989267298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.6382874220407399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":8.5364821169174009E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-9.0176074144511997E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.8584952436674599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.9098141699986902E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.90477505429944E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":3.0388719823007702E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":1.3073252113581599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.5311342833586901E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-1.0723303480283299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.2394598320507001E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-2.09334904399994E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.2562974508641201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.123712070589236},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.8752161059000299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":2.27792455824306E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":1.2758544103096501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.4230844225651798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.1133480396335001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["074SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A31,ConcTable!$A:$A,0)-1,0)","v":"19286426-019"},{"$f":"OFFSET(SQ!$D$1,MATCH($A31,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:13:57"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.532},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":3.036},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-312.154857909246},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":0.83042628447925904},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-0.85159460733400705},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":0.26279384743744},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.8337158296575399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.17295631586209E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.2591362076641602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.8777390310254997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-0.310808538607662},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":3.0771056752912899E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":6.5828673146180905E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.6093259283981899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-0.18772675058887001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":1.07373166636126E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-7.7485136326430103E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-4.3365991379207998E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-7.1774967349878702},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":1.6024298238263401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.0195360742298499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-4.62033521352003E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.4167980652586201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.54465583170854E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.4963302450600099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.2237867821963101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.33597688357486E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-3.1544742563929503E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.407822158641582},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.2188874848016301E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-8.8695871295327595E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-9.5163554995427802E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0370019377001999E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.10032722846916101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.6131815203176201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-1.1844269829123599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.7741865112580798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.3551062143987497E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.0971440750570602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["075SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A32,ConcTable!$A:$A,0)-1,0)","v":"19286426-020"},{"$f":"OFFSET(SQ!$D$1,MATCH($A32,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:15:26"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.5249999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":2.077},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.59830338771091396},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-337.378546130887},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-0.45084237246483},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-10.4874675869941},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":1.6747822344444501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":2.7460097336352501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.98312818676614E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.93549446665294E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":1.1783136551051799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.57598143015000403},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.26486757429352E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":3.7542097904223699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":2.7265135865975301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":-3.14218014998872E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":1.4241210033365799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-6.7220779212731402E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.1115159003561101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-9.9755234699290405},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":5.7266408923935301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-4.3246741370461398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":3.54755475143148E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-1.51783359453523E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.98137948572527E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.9021004375504501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.25775722058702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.4209661679568299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.426431472874088},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-8.4185209528912005E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":2.2554012260785E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.4860020240075201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.13268316324768101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-3.08342364472192E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":5.8171436180499003E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":1.6294584125523201E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.2046449701532299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.1745413799778903E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.1004948608940702E-5}]
[]
[]
[]
["055SMPL.d","Blank Control1",{"$f":"OFFSET(SQ!$D$1,MATCH($A36,SQ!$C:$C,0)-1,0)","v":"2019-10-22T12:45:17"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-5.0000000000000001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-8.2000000000000003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.20300000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":2113.8075826409399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":5.8697407802854897},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":122.018874927563},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":25.275822426967299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":0.13026997098266299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":7.0066080000694003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.22574452766891301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.12965835835675099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":1.3597532227772999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":5.6493818431130804E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.5085837279854999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":7.5960260606831698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.22244254279942799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":8.9858031289419306E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.86456929372627E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":3.7015002829832201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":8.7131655300839397E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":26.975524187702799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.0430893553099902E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.1643545156889999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":2.64435675611399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":5.82437757862373E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":4.1731315588822001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":3.4003987120378297E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":1.4716630472016E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":7.3131058499069299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.7120978106384701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":0.66263577333549195},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.05183047319981E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":9.8320505104675607E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.56406076390156E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.22491686541486899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":3.4073310040638197E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":1.7580955767890601E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":5.6442292293266199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":4.7735695160255201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":3.2550438073356801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["079SMPL.d","Blank Control2",{"$f":"OFFSET(SQ!$D$1,MATCH($A37,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:21:28"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.5569999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.85799999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.15882922732186999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-174.89496021349501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":8.8354840455750896},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":5.6156938629900699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-3.1782012929093399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":2.3736835251514099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":8.9651817322323902E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":7.8579740419963601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.5860825352282501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.70163607809190598},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.2910611994337001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":8.6146350249957704E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":5.5142158234920399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.50775711061140105},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-3.6682738333033899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":5.6773969821747499E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-7.6322849844029399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.0855988207369E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-2.9080569540014E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.7213432370428001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":7.3345983860963901E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":2.4068185286491899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-3.27864587514336E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":8.9960281973894103E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":4.4703944655938299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.1897444219477299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.182196509753662},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":1.6842413458975199E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.0299171879105804E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.41285244142556E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":2.04829506032006E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.38365706977439901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-3.0413218427988602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-4.3010562961238903E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.6715004363840498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-3.2216276840783903E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]

{"sheet":"Summary2","index":8,"n_rows":50,"n_cols":6,"rows_kept":50,"cut":false,"hidden":false,"merged":[[0,0,0,5]]}
[{"$f":"INDIRECT(A50&\"!A5\")","v":"acid washed 10 mL"}]
["Element","Result","Average (ppb)","Standard Deviation","Average + 2SD","Reporting Limit"]
["Lithium",{"$f":"OFFSET(INDIRECT($A$50&\"!$A$10\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":"Pass"},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$7\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":-1.8736842105263159E-2},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$8\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":2.7047716121114946E-3},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$9\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":-1.3327298881040169E-2},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$2\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":10}]
["Beryllium",{"$f":"OFFSET(INDIRECT($A$50&\"!$A$10\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":"Pass"},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$7\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":1.0526315789473685E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$8\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":3.1530176764230577E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$9\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":7.358666931793484E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$2\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":0.25}]
["Boron","Pass",-1.3235263157894737,0.21760498575299617,-0.88831634428348138,25]
["Aluminium","Review",2.2755263157894738,3.6001885809931475,9.4759034777757698,10]
["Magnesium","Pass",0.57779839834678215,0.75283555543704506,2.0834695092208722,200]
["Silicon","Pass",-137.41852289998971,115.55372806151925,93.688933223048792,500]
["Phosphorus","Pass",2.9908347801601599E-2,2.3405394467337381,4.7109872412690779,500]
["Potassium","Pass",-6.90393418339292,4.9556854720123553,3.0074367606317907,500]
["Calcium","Pass",-0.35220547206186081,3.6348246977342762,6.917443923406692,500]
["Titanium","Pass",-1.3513087294231955E-2,3.0862538704661317E-2,4.8211990115090683E-2,5]
["Vanadium","Pass",8.1328230015633484E-5,9.6765361465759574E-3,1.9434400523167548E-2,0.5]
["Chromium","Pass",1.8293695988449672E-2,2.5769511475327824E-2,6.9832718939105323E-2,0.5]
["Manganese","Pass",4.2592207625348745E-2,2.8434520979437248E-2,9.9461249584223241E-2,0.5]
["Iron","Pass",0.18403887057730073,0.35061943087619968,0.88527773232970008,20]
["Cobalt","Pass",-5.8347893426083234E-4,9.6934768980146656E-4,1.3552164453421007E-3,0.25]
["Nickel","Pass",5.1570795675138358E-2,3.9964528209102693E-2,0.13149985209334375,1]
["Copper","Pass",7.3589783428011646E-3,1.7945127725241063E-2,4.3249233793283291E-2,5]
["Zinc","Pass",-1.1567740230399753E-2,0.12605053909025299,0.24053333795010623,50]
["Gallium","Pass",-4.9380002720002811E-4,4.5555940157618541E-3,8.6173880043236804E-3,0.5]
["Germanium","Pass",-9.5751740064903713E-4,5.9441844705509861E-3,1.0930851540452934E-2,0.5]
["Arsenic","Pass",1.5681559479164414E-5,1.630727523235388E-3,3.2771366059499403E-3,2.5]
["Selenium","Pass",-2.7071023218607444E-2,2.9127635064189968E-2,3.1184246909772492E-2,10]
["Bromine","Pass",-6.431128423651514,2.1711172905202725,-2.0888938426109691,250]
["Rubidium","Pass",1.0195221297253193E-2,2.4546894296741898E-2,5.9289009890736988E-2,25]
["Strontium","Pass",-3.7353252213692088E-3,4.4049049514727238E-3,5.0744846815762388E-3,5]
["Zirconium","Pass",8.2129587809565412E-4,1.947695521300321E-3,4.7166869206962964E-3,2]
["Molybdenum","Pass",-3.1578815867300859E-3,1.0777931335002604E-3,-1.0022953197295651E-3,0.5]
["Palladium","Pass",-2.4847997352979607E-3,1.2521701317617026E-3,1.954052822544447E-5,0.5]
["Silver","Pass",-2.2847814088326777E-2,6.0167364980755322E-3,-1.0814341092175713E-2,1]
["Cadmium","Pass",2.408563090253719E-4,3.4963051524095697E-4,9.4011733950728586E-4,0.25]
["Indium","Pass",0,0,0,0.25]
["Tin","Pass",3.327248830340183E-3,5.4201805905998435E-3,1.416761001153987E-2,1]
["Antimony","Pass",8.5952878241022388E-4,1.927719385556622E-3,4.714967553523468E-3,0.5]
["Tellurium","Pass",3.3016388798718526E-4,1.4391510225426232E-3,3.2084659330724316E-3,0.5]
["Iodine","Pass",-0.25834990699794941,0.2086354023872096,0.15892089777646978,10]
["Caesium","Pass",8.0511180854246508E-6,5.6250628775741636E-4,1.1330636936002574E-3,0.5]
["Barium","Pass",-2.5825592456599337E-3,6.5116018928334735E-3,1.0440644540007014E-2,2.5]
["Tungsten","Pass",-1.1173840531912408E-3,3.1780389876500737E-4,-4.8177625566122608E-4,0.5]
["Platinum","Pass",7.7905589309486105E-5,1.7209981676633303E-4,4.2210522284215216E-4,0.5]
["Gold","Pass",-9.9450607999082188E-2,4.1379439405970019E-2,-1.6691729187142151E-2,5]
["Mercury","Pass",-2.4149828338544409E-2,5.598440752665022E-3,-1.2952946833214365E-2,0.5]
["Thallium","Pass",-1.286901895364591E-5,1.1474254730317303E-4,2.1661607565270016E-4,0.25]
["Lead","Pass",-2.2911361861198197E-3,2.2025472543715828E-3,2.113958322623346E-3,2.5]
["Bismuth","Pass",-3.074091127379561E-3,4.0650218208467389E-3,5.0559525143139168E-3,0.25]
["Thorium","Pass",-2.8317582585904177E-2,4.5962289816614014E-3,-1.9125124622581376E-2,0.25]
["Uranium","Pass",2.3815924467229106E-5,3.9572576969886147E-5,1.029610784070014E-4,0.05]
[]
["Container2"]

{"sheet":"Container3","index":9,"n_rows":37,"n_cols":49,"rows_kept":37,"cut":false,"hidden":false,"merged":[]}
[null,"Name","Acq. Date-Time","Li_7","Be_9","B_11","Al_27","24  Mg","28  Si","31  P","39  K","44  Ca","49  Ti","51  V","52  Cr","55  Mn","56  Fe","59  Co","60  Ni","65  Cu","66  Zn","71  Ga","74  Ge","75  As","78  Se","79  Br","85  Rb","88  Sr","90  Zr","98  Mo","105  Pd","109  Ag","114  Cd","115  In","118  Sn","121  Sb","125  Te","127  I","133  Cs","137  Ba","184  W","195  Pt","197  Au","202  Hg","205  Tl","208  Pb","209  Bi","232  Th","238  U"]
[null,null,"RL",10,0.25,25,10,200,500,500,500,500,5,0.5,0.5,0.5,20,0.25,1,5,50,0.5,0.5,2.5,10,250,25,5,2,0.5,0.5,1,0.25,0.25,1,0.5,0.5,10,0.5,2.5,0.5,0.5,5,0.5,0.25,2.5,0.25,0.25,0.05]
[null,null,"Low QC",{"$f":"MIN(QCReview!D26:D100)*D2","v":8.5139999999999993},{"$f":"MIN(QCReview!E26:E100)*E2","v":0.21099999999999999},{"$f":"MIN(QCReview!F26:F100)*F2","v":19.602},{"$f":"MIN(QCReview!G26:G100)*G2","v":8.9},{"$f":"MIN(QCReview!H26:H100)*H2","v":180.71063948691099},{"$f":"MIN(QCReview!I26:I100)*I2","v":150.581516610791},{"$f":"MIN(QCReview!J26:J100)*J2","v":449.659182698216},{"$f":"MIN(QCReview!K26:K100)*K2","v":509.50512352205499},{"$f":"MIN(QCReview!L26:L100)*L2","v":378.75559869518997},{"$f":"MIN(QCReview!M26:M100)*M2","v":4.6648107431519001},{"$f":"MIN(QCReview!N26:N100)*N2","v":0.44575373577775101},{"$f":"MIN(QCReview!O26:O100)*O2","v":0.44711397777342698},{"$f":"MIN(QCReview!P26:P100)*P2","v":0.45988192477422901},{"$f":"MIN(QCReview!Q26:Q100)*Q2","v":17.709543709225699},{"$f":"MIN(QCReview!R26:R100)*R2","v":0.20697248682051},{"$f":"MIN(QCReview!S26:S100)*S2","v":0.88745615928769706},{"$f":"MIN(QCReview!T26:T100)*T2","v":4.5006272265695602},{"$f":"MIN(QCReview!U26:U100)*U2","v":42.7624616138785},{"$f":"MIN(QCReview!V26:V100)*V2","v":0.44235442630143601},{"$f":"MIN(QCReview!W26:W100)*W2","v":0.56798150017591098},{"$f":"MIN(QCReview!X26:X100)*X2","v":2.2798906623553399},{"$f":"MIN(QCReview!Y26:Y100)*Y2","v":7.9895749025331693},{"$f":"MIN(QCReview!Z26:Z100)*Z2","v":279.40181851843801},{"$f":"MIN(QCReview!AA26:AA100)*AA2","v":21.224836651076298},{"$f":"MIN(QCReview!AB26:AB100)*AB2","v":4.5993333721576004},{"$f":"MIN(QCReview!AC26:AC100)*AC2","v":2.4624686572192398},{"$f":"MIN(QCReview!AD26:AD100)*AD2","v":0.65732553977758001},{"$f":"MIN(QCReview!AE26:AE100)*AE2","v":0.544577383656354},{"$f":"MIN(QCReview!AF26:AF100)*AF2","v":0.81312160603667305},{"$f":"MIN(QCReview!AG26:AG100)*AG2","v":0.210932860609923},{"$f":"MIN(QCReview!AH26:AH100)*AH2","v":0.211447995002015},{"$f":"MIN(QCReview!AI26:AI100)*AI2","v":0.98118408712764604},{"$f":"MIN(QCReview!AJ26:AJ100)*AJ2","v":0.47512606879184399},{"$f":"MIN(QCReview!AK26:AK100)*AK2","v":0.36104432002066},{"$f":"MIN(QCReview!AL26:AL100)*AL2","v":22.033046859313501},{"$f":"MIN(QCReview!AM26:AM100)*AM2","v":0.45216308043174602},{"$f":"MIN(QCReview!AN26:AN100)*AN2","v":2.1326937795210901},{"$f":"MIN(QCReview!AO26:AO100)*AO2","v":0.40638711024189},{"$f":"MIN(QCReview!AP26:AP100)*AP2","v":0.39720220550103602},{"$f":"MIN(QCReview!AQ26:AQ100)*AQ2","v":9.04952796357958},{"$f":"MIN(QCReview!AR26:AR100)*AR2","v":0.50778669339083604},{"$f":"MIN(QCReview!AS26:AS100)*AS2","v":0.18995279692808301},{"$f":"MIN(QCReview!AT26:AT100)*AT2","v":1.9855488339795799},{"$f":"MIN(QCReview!AU26:AU100)*AU2","v":0.237213785892985},{"$f":"MIN(QCReview!AV26:AV100)*AV2","v":0.20103154901252401},{"$f":"MIN(QCReview!AW26:AW100)*AW2","v":4.0606714830781203E-2}]
["Container Name",null,null,"Lithium","Beryllium","Boron","Aluminium","Magnesium","Silicon","Phosphorus","Potassium","Calcium","Titanium","Vanadium","Chromium","Manganese","Iron","Cobalt","Nickel","Copper","Zinc","Gallium","Germanium","Arsenic","Selenium","Bromine","Rubidium","Strontium","Zirconium","Molybdenum","Palladium","Silver","Cadmium","Indium","Tin","Antimony","Tellurium","Iodine","Caesium","Barium","Tungsten","Platinum","Gold","Mercury","Thallium","Lead","Bismuth","Thorium","Uranium"]
["50 mL acid washed"]
["Number of containers"]
[20,null,"Average",{"$f":"AVERAGE(OFFSET(D$12,0,0,$A$7,1))","v":-1.8315789473684212E-2},{"$f":"AVERAGE(OFFSET(E$12,0,0,$A$7,1))","v":1.0526315789473685E-4},-1.0217368421052633,0.26452631578947372,-0.23937399537704823,61.498573352765142,-2.0984003652770014,13.855357016978015,2.4217043604347812,3.0607059693709856E-2,1.0135291784705174E-2,1.8631254442402264E-2,3.3166009174656064E-2,6.8803584603456258E-2,-7.7363073864977164E-4,5.2396048035437869E-2,3.0134953890754702E-2,55.298010406535731,-6.9671118726126425E-3,-2.9923741184568055E-3,4.993252632980148E-4,6.6493017544048015E-3,-2.577818671796392,-3.6801715364541202E-2,-6.2263461856045593E-3,4.7525736752840922E-4,-4.1385303518445644E-3,-1.8984765528776725E-3,-2.055794694006529E-2,3.8174908764113347E-4,0,-2.711920979409739E-2,1.2287891592981403E-3,0,-0.67247404019236179,-1.570220179316621E-3,-4.2092183067321412E-3,-1.6458432890333836E-4,5.1936664085896786E-5,-0.17626135906752857,-2.2118587363171544E-2,-3.3094019716774876E-4,-4.2623593794651886E-4,-4.5381258635102716E-3,-2.4455585830378008E-2,1.9680714594182863E-5]
[null,null,"SD",{"$f":"STDEV(OFFSET(D$12,0,0,$A$7,1))","v":2.6885070532618278E-3},{"$f":"STDEV(OFFSET(E$12,0,0,$A$7,1))","v":3.1530176764230571E-4},0.2125372757124063,0.37506790028548348,0.25217668809230359,82.8549529112762,2.818894508279385,6.3625910336330822,2.9388545389867593,3.1184410465421933E-2,7.4396480185479107E-3,1.3654643294718412E-2,1.2454164706893353E-2,0.16584020769731189,9.756822268868631E-4,7.2296409120795144E-2,9.2270765809901789E-2,5.3040588351120919,4.9195854821148461E-3,5.2860191868483435E-3,2.2714159949779253E-3,2.4296997213841598E-2,2.3847520688702621,2.9298135604016284E-2,1.8688844618340767E-2,1.5679287237252203E-3,1.2937880965186632E-3,1.2858415080153898E-3,6.0927383537682546E-3,3.4947622272601313E-4,0,7.486699341197548E-3,1.005469241236458E-3,0,0.12737118997157426,4.8583245863993615E-4,3.0826211601820324E-3,2.0667553112884397E-4,1.102180531137353E-4,4.1740136742664705E-2,4.6693506190975502E-3,2.117066303786603E-4,3.9538714425378452E-3,6.2724315127319686E-4,3.7632563243504844E-3,3.4623157953879255E-5]
[null,null,"ave+2sd",{"$f":"D7+D8*2","v":-1.2938775367160557E-2},{"$f":"E7+E8*2","v":7.3586669317934829E-4},-0.59666229068045062,1.0146621163604408,0.26497938080755895,227.20847917531754,3.5393886512817687,26.58053908424418,8.2994134384082994,9.2975880624553725E-2,2.5014587821800993E-2,4.5940541031839088E-2,5.807433858844277E-2,0.40048399999808004,1.1777337151239547E-3,0.19698886627702816,0.21467648551055829,65.906128076759913,2.8720590916170496E-3,7.5796642552398812E-3,5.0421572532538652E-3,5.5243296182087995E-2,2.1916854659441323,2.1794555843491367E-2,3.1151343051076974E-2,3.6111148149788499E-3,-1.550954158807238E-3,6.7320646315310699E-4,-8.3724702325287809E-3,1.0807015330931598E-3,0,-1.2145811111702294E-2,3.2397276417710564E-3,0,-0.41773166024921327,-5.9855526203674871E-4,1.9560240136319235E-3,2.4876673335434956E-4,2.723727703133674E-4,-9.2781085582199163E-2,-1.2779886124976443E-2,9.2473063589571848E-5,7.4815069471291717E-3,-3.2836395609638779E-3,-1.6929073181677037E-2,8.892703050194137E-5]
[null,null,"Pass/Review",{"$f":"IF(D9>MIN(D2:D3),\"Review\",\"Pass\")","v":"Pass"},{"$f":"IF(E9>MIN(E2:E3),\"Review\",\"Pass\")","v":"Pass"},"Pass","Pass","Pass","Review","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Review","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass","Pass"]
[]
["Containers"]
["087SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A13,ConcTable!$A:$A,0)-1,0)","v":"19286424-001"},{"$f":"OFFSET(SQ!$D$1,MATCH($A13,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:33:29"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.0999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.48799999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.316},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.11499585145772},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":112.964426194016},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-3.3270832943264801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":13.8350140730922},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-0.35525727598773599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":4.7145065258997303E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.1568189757301599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.5022185815487001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.2860995256321798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-0.15297831166977099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-4.23786899548978E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.5209851519857299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-9.4952334245034103E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":52.689281508778798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.26937237736702E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":6.57298137187271E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":7.7179344272655503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":6.8324193900649394E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-0.46771231647584},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-3.0569403158917E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":6.4148297545301904E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-3.9857981997355702E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.6237945974462302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":2.0142846147415401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.8380753999235299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-7.9413556540595693E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.19195628735153E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.48436263555099301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.2323668903842799E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-4.8176489012563897E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.6829441643680099E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.91899075670585E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-9.8363136411364002E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.3123206798936599E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.0477332341054101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":2.560771708229E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.7242949835092401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-1.19094769355762E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A13,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":8.3196097545209995E-5}]
["088SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A14,ConcTable!$A:$A,0)-1,0)","v":"19286424-002"},{"$f":"OFFSET(SQ!$D$1,MATCH($A14,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:35:04"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.2999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.61099999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.193},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.101378451448013},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":143.63425941518099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-2.6156514883370399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":15.354846649750399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":4.3143575534070902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":4.9618809968932601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-3.69698319240292E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.6736670336472501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":7.6732308623999601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":7.65512189554753E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.56065459277985E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.70259932179154E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-3.0426793833779298E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":57.256171544946298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.39808047438395E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":8.9199798906808708E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.0831157470965998E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":0.48949526818019601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":4.5365160599544603E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":6.8250897079884507E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.43164357802159E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.4095302089607099E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.8741769510122301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.2478222716295499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":4.4239741256570401E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.1339529947328498E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":5.1333269501246601E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.431948090248576},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.4374024477927E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-5.7908590152390797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-2.2527421358099501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.11449777587047499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.57466155197401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.47237026465035E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-5.0629326967945704E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.2025287954616198E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-1.9132887439822999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A14,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["089SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A15,ConcTable!$A:$A,0)-1,0)","v":"19286424-003"},{"$f":"OFFSET(SQ!$D$1,MATCH($A15,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:36:33"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.6E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.746},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.13500000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":141.532248890402},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-4.4694514026023899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":2.9528492404728399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.5880458954619801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":6.0678459678491104E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":2.4016353619010899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":-6.1426611467152404E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.4948446479090199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":6.3023229494429497E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":6.1842335630566899E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.1973285855846499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.12033503441906E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":59.723492183071798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.26425594596237E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-8.4769452287546294E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":2.07501641696592E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-5.5788707604231798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.34327965899304E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.28696102895618E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-1.7434703082958301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.6171279247009499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.1865693017547E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.51660139503286E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":4.4900512280168699E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-1.9872721292815299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.48086612462051E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.55625816857838295},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.2323775624149299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-5.8132670354615404E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-2.27649184036425E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":4.3099073291900101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.12661445175602101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.31233726550345E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.0477531215263101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.67837241481194E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.3736083503913298E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.0776643615659798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A15,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.1597630354344902E-5}]
["090SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A16,ConcTable!$A:$A,0)-1,0)","v":"19286424-004"},{"$f":"OFFSET(SQ!$D$1,MATCH($A16,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:38:03"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.7999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.85799999999999998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.11899999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.41607272702869802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":106.29691087435999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-3.2906232317461201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":16.148937880267699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":3.6708252872743099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":5.4606417523409403E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.06205697703569E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":5.32857425041017E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.1198339391764002E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-0.262987133703456},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-2.0333412841291401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.7917081321837399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.7981017833861801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":56.761338036477298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-8.5698198295035104E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":2.9346658963756399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":4.0820861896633696E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.2946406202412201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":1.0650504340212701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.18192112259068E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-9.6855596247497795E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-3.4427413040941801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.3459254074252E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-8.1781758106908996E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.6362469691052699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.3583517005669605E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-8.12668144680527E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.5033624836059199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.63834468199051697},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-5.2447935034213801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-1.5806331276084001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.93883919546846E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.114214714876695},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.6445509106595299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.2132515089746098E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.0358924234250599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.3347360591524404E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.2921438346455499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A16,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.3091757218143797E-5}]
["091SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A17,ConcTable!$A:$A,0)-1,0)","v":"19286424-005"},{"$f":"OFFSET(SQ!$D$1,MATCH($A17,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:39:32"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.6E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-0.89400000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.38200000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.398188018513979},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":141.21508749780801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-2.45710615819773},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":20.746761604381501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":6.5764943349017697},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":6.0072000509229298E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.20112208204914E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.87402217306861E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.3891593622340002E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.175307495437032},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.0569152283892199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.6602432452371898E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":5.0212068560227703E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":51.851422039889698},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-5.6960836267530403E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-7.2421242179551001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":4.8954622396369399E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-0.25573375557355099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-6.2633337853126803E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-9.6716220358399503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.12276152712819E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.7767722612142901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-5.1935570527720804E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.82851144000909E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.7024319902994101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.5600541651643401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.76590685553673099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.8568252094687901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-4.7267347225851703E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.3942043373982299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":4.3678602464260196E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.14208371237432299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.0552858069553401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-7.1676406524622904E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.2013080638443299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.44220825581518E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A17,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["092SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A18,ConcTable!$A:$A,0)-1,0)","v":"19286424-006"},{"$f":"OFFSET(SQ!$D$1,MATCH($A18,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:41:05"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.7999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.0069999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.125},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":6.9500092405132802E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":89.6049469134596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":3.1907994904518602},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":12.899783748150201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":0.13142791518696501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":1.02403723754362E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.42408669001236E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.6855012797463499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.6421586631579198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.41701996706279798},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-6.9853064187544499E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":0.32994505853800099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.169192505652733},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":56.884917027458201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.1185369272718E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":2.1622382384070202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.8226688169287598E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-4.46099007527389},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-2.8253836671989E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-8.0842609813342806E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.2886740427191801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-5.8509167542042395E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.4932972679640899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.9896335250851702E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.1999569465367098E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.60179524190696E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.68834090895709599},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.84913289832703E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-5.5758822111163196E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.9314184217407199E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.169234295893369},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.1561099322557401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-2.7574541570272201E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.21932784986812E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.6345028818968103E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.3953160921326801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A18,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["093SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A19,ConcTable!$A:$A,0)-1,0)","v":"19286424-007"},{"$f":"OFFSET(SQ!$D$1,MATCH($A19,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:42:34"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.7999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.1220000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":9.2999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.55503872355013495},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":109.183683192537},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-4.9977369538224004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":18.912750230140201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":4.6195573221071697},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":2.2278548292052398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":7.3275567674456303E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.00951137369176E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.78437835909212E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":2.24643394801218E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.2026799659497699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.32819925733242E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-5.28608097294585E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":50.211688930238203},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-4.1953841227122396E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-6.0835738251059698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":3.2727730714858902E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":0.72639477148988596},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-3.6348773869319798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.42050575637515E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-1.2598123122222801E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.3425638245980701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.1627440077983001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.0310152656825099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":5.6196975854247097E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.6862187583302201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.5099471730034599E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.76205838729831299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.8474965920112899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.7343014745678701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-2.1579191023369699E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.16153652254162701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.4378031791754198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.21144711590118E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.7821614041324897E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.49710569139064E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A19,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["094SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A20,ConcTable!$A:$A,0)-1,0)","v":"19286424-008"},{"$f":"OFFSET(SQ!$D$1,MATCH($A20,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:44:03"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.7999999999999999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.0720000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.218},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.38962642121885499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":116.142624602943},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":2.2164200210383602},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":14.5706016190589},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":1.21071396462454},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.5716469920202797E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":2.5620870397969099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":-3.88972175921326E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":1.8813243297412598E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.110842927754086},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-2.6698266983526198E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":7.7523662083417003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.6981772893254601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":47.549788172596998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.8054094716993599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-6.1920747202801303E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.7283537394026901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.04170466968248E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-2.03057816837743},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-4.5968908790201997E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-6.7198566672288899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-8.2693222800870204E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.0195436958682098E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.5336698154769998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.9810157703569999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.6833558268181099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":6.2150810442962803E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.55488506170901997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.0832111356764399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-8.1250122558604299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-3.8556989695787602E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":2.1737841292329501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.17740812935547601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.4605195728048999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.23286207288658E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.7587234810265901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.0216566002337397E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.4615965736673798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A20,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["095SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A21,ConcTable!$A:$A,0)-1,0)","v":"19286424-009"},{"$f":"OFFSET(SQ!$D$1,MATCH($A21,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:45:33"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.9E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.0069999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.14000000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":66.747631660152805},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-4.4188765958207297},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":15.0327143877544},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":4.8354488772082096},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-1.17121756409924E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":8.8376274236588704E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.3491264962729699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.3536996617385503E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-0.15748513197682401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":4.4289701223735201E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.077018044692E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-6.7075486019014304E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":58.330366461816702},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-7.1137490365825303E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.03851437682934E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-2.0435180693809899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-2.3254886262471799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-6.2703450030859295E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-8.1583676521545904E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":4.7213489916372196E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-5.7346968460542797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.84645119871276E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.17480532325003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.2085949551079105E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.2484734268198499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-2.1020078115338599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.44628789341641001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.4192346485162701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.7599402363653603E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-3.82134778906494E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.15987452376101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-1.9412723085362502E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.6120080999644903E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.5021246582514498E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.01226790616987E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.5252185938097398E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A21,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.2619310782601998E-5}]
["096SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A22,ConcTable!$A:$A,0)-1,0)","v":"19286424-010"},{"$f":"OFFSET(SQ!$D$1,MATCH($A22,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:47:06"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.9E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.0980000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.124},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":120.023029508836},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-1.64552927024242},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":17.805310490846299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":3.2923611313080201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-1.5363517245603199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":7.9627781403368292E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.55280407484226E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.3631905040949699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.164475764021427},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.7219768163612499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":7.7551960513912804E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.1369383247547601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":60.6177812647473},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-8.2498648970415998E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.6862526895532299E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-1.89160247599678},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-2.71492189469556E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-7.9245087568573895E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.0980688329104099E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-5.3539923364308798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-8.0459325412828801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-1.89462308035716E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":6.6229106769906897E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.76790990986228E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.02145983873297E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.78606326453042796},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.28871331871296E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.7272016648107398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-2.1466247038513101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.17937442343872601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.3784484158616202E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.2102338516694599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.5734591803161502E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.8549457812457897E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.5870285224132698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A22,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["097SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A23,ConcTable!$A:$A,0)-1,0)","v":"19286424-011"},{"$f":"OFFSET(SQ!$D$1,MATCH($A23,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:48:36"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.133},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.108},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.54443931149138702},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":101.30666937318099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.1184995165073299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":6.4562709685143496},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-0.10054277274019099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":1.9256056715541099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.19084796945957E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.5857898431718903E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.6942268081687699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.15179187201196001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-4.8133388313523998E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":7.65513356581418E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-9.6761120671011395E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":51.319098943439997},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-5.3787900891108902E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.1279047323278999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.6738978893274701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":3.9203780544130899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-6.9493934185892998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-1.8360693714192699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.5729027064768902E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":2.0337589421650501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.7183334707881202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.1749755366610398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.1648895588655199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":6.8407494037985201E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.9438260270990401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.60453729145942E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.78734265864626096},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.0692389190488502E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-5.70673709253807E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.34341298617075E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.16838388589894E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.207516667245323},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.3255286645130699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.2134003805127797E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-3.4934710250776602E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.6338752396586202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.5762222081979699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A23,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["098SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A24,ConcTable!$A:$A,0)-1,0)","v":"19286424-012"},{"$f":"OFFSET(SQ!$D$1,MATCH($A24,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:50:05"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.077},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.20100000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":70.425805054992196},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-3.0515919578806399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":18.174079045692999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-0.109139369938706},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":1.5823129150755198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-1.5639639334433699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.2746160769816299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.02328932043352E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.23799537597873899},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":2.3506349199791299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":4.3863566464275701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":4.4376861718892898E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":68.195409545047696},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-7.0363225267994003E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-7.2178270557066403E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.5949238272744202E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-3.8505330830877802},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-7.1162224719527004E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.57950354888515E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":2.8698954468399499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.5922484634989399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.8499391421690501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.5702151149098001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.0429979418488601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.54450569815945E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.74283970483998896},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.4264217147102901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":2.34644167040297E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-2.22008244450563E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.1300784046316101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.200455369764658},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.3392194528117499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-3.4120620873807997E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.2718390365180901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.7214746044793596E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.6163240925194699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A24,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["099SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A25,ConcTable!$A:$A,0)-1,0)","v":"19286424-013"},{"$f":"OFFSET(SQ!$D$1,MATCH($A25,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:51:47"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-1.9E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.0980000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.252},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.37212943367554702},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":103.39914517572301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-6.8769921385627004},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":20.872363626989902},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":7.4936719118016297},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":8.1687245021134003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.5763151828133801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.3704874277666698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.10961894394613E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":5.4356954680691003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":6.5105728042999403E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":7.1797467257777595E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.0538490871706101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":54.988360971931002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-2.5646820793255101E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-6.9984574790696004E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":-1.9874796774363102E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-5.5128168554714101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-4.4919961410820899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.26667032581077E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-8.0261283076962699E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-2.5988765823299498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.8511910320128901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.0961808096018698E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":6.02762406262417E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.0445021816893798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":9.4889031865033602E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.62167976998439101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-7.7927840985483399E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.7772115816638898E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-3.8350838398908197E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.18991735718693301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.5075994453163501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.22513785697986E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":1.0446575492652E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.5762839621025899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.6937773555841302E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A25,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["100SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A26,ConcTable!$A:$A,0)-1,0)","v":"19286424-014"},{"$f":"OFFSET(SQ!$D$1,MATCH($A26,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:53:21"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.2250000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.14799999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.71431680789830898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":47.578054212232701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-4.3198457299978603},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":11.7073112093908},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-0.31457743384321502},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":3.6185461267816903E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.5392186482708097E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.9901453496161102E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.0289660723890601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.14185459962062399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":7.5233250274761305E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.50592893234277E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.3479179752680199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":49.851859640658397},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-5.6283354445825296E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-8.8739771701807496E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":2.5789461132830801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.9003839222464002E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-3.1878032200884898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-2.00008112973563E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.120572359478E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.06735615225877E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.40103182769854E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-1.8610131831064401E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.83036844981089E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.0029450315591798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":3.5890986638624201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.67046852863061601},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.64511426513162E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-3.0313127902258802E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.22222871448558801},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.7302502489528899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.62264447474721E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-1.9037423816392399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.1626841687148596E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.7056662672150001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A26,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["101SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A27,ConcTable!$A:$A,0)-1,0)","v":"19286424-015"},{"$f":"OFFSET(SQ!$D$1,MATCH($A27,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:54:51"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.1439999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.129},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":8.0347275415264499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-14.458752971989},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-2.4540135163921399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":20.478960252670898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":6.9565591521732397},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":7.2930195109554605E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":9.5814069993207501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.8149310286089401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":1.5194970641849799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-7.7783337895785995E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-9.8421088215416711E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":2.2872941596224401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-1.78683177838318E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":48.563367548096998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":5.0978960059561698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-7.3233403798387102E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":2.1526839438452201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.2737273872795402E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-2.7845785702508898},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-7.1478579038141296E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-9.7956951954687108E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":9.8614568815502896E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-6.1431398446613701E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.85937610531572E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.2790837054536699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":8.9006542782140504E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.72414123325866E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":-2.4706491026977402E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.79509800934793096},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.6555980136658201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-7.8011429138877498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":1.70634694973802E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.19151927186604001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.6800213884530799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-5.2326096359550504E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.1104650175964601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.3833785467669298E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.6662776282784E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A27,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["102SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A28,ConcTable!$A:$A,0)-1,0)","v":"19286424-016"},{"$f":"OFFSET(SQ!$D$1,MATCH($A28,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:56:21"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.17},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":1.7809999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":11.559543971016399},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-4.9679471186996196},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":20.4249737873347},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":3.4077600187094701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.1112240702259502E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.24235885108651E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.28202212937124E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":5.59802943086204E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.193879550381796},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.31953844374611E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":6.6155624777741798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-3.9870743050956199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":53.996588955522803},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.2557845312457501E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-2.0723911416143698E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":2.1825913746506301E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.3926261630628398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-3.17835615993378},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-8.8306545104179499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.4331577721114701E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-8.3647436335355203E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-5.2126358271890197E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-9.2619930259847801E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.6667654082294899E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.5352815344251797E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.5931586239873602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.4508841312516499E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.82109044910278095},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-8.1616894402550299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-3.8693686904423501E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.23655285296093101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.0850423149454601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.6479957200834299E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-9.07282150233233E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.1688485659711798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.65041033053534E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A28,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":4.1358020505684701E-5}]
["103SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A29,ConcTable!$A:$A,0)-1,0)","v":"19286424-017"},{"$f":"OFFSET(SQ!$D$1,MATCH($A29,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:57:50"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.2509999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.13},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.26081497225153299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-110.883822174061},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.05964358407422},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":8.5964629156523493},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.85711555427518},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-7.6696782285873601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":1.4343152420381E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":-1.66892042140238E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.1539259577270601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":-0.11800280861099401},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-7.0566637499629896E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":4.5669644246482096E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.15324825596189701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":52.512200361949397},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.1412603516288301E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-2.6508641739782001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.8132281650537899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":4.4789662364234303E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-4.8112674317786297},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-4.80201278539448E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.15709613520593E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":-4.4102684387111103E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.8767224696526596E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.5677953661096199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.4722623532134499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":7.5044371009892698E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.08203155880632E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.0898195638691301E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.82677904027365301},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.6857790149047501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-3.9820237566726901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.6118254635912599E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.02294391733556E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.20798298688489},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.5895292205992598E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-2.95625423359402E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.53185458030118E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.8273264521709498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.70049915842831E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A29,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["104SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A30,ConcTable!$A:$A,0)-1,0)","v":"19286424-018"},{"$f":"OFFSET(SQ!$D$1,MATCH($A30,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:59:23"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.234},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.17599999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.57141011334633196},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-91.200517992871298},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":0.35864636306768799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":-2.1896108456938901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":0.36434898771517199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":5.3225420254285602E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":3.55715497448549E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.6778039640963799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.9169158983041597E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":4.0289267729745303E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.6718014997922899E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":1.6061867531475201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":2.2420990202892E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":63.622516002870697},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-6.2411348389531203E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-5.7297430642964499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":0.21900016453075499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-3.21879752735321E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.4509523628892801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.29302356777118E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.7792509787917001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-3.19843258065637E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.5979195769019001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-3.24400175857677E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":7.9752856153037302E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.64302178308112101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-2.0999829317185702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.8815160564610501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-3.9180475446004098E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.22379417228617199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.7502222031401201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-4.2235778554032002E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-4.5521227847469302E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.7292132451121699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A30,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["105SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A31,ConcTable!$A:$A,0)-1,0)","v":"19286424-019"},{"$f":"OFFSET(SQ!$D$1,MATCH($A31,SQ!$C:$C,0)-1,0)","v":"2019-10-22T14:00:52"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-2.1000000000000001E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":1E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.1779999999999999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.25600000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.259542448103806},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-96.598079695381998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":1.07883294122578},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":10.471402438115501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":3.46353469409026},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":2.37540271685922E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":9.4631690437900796E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":1.74600623921944E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":2.88893585681438E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.22665226871357499},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-8.6849446685382603E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.5013953104655199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":0.34099004733845101},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":55.736548584640602},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-2.1454080684648498E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-2.5664202547361001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":-1.8012149300871999E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-4.1927704847854299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-3.1281896976926203E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.0043213485379499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":9.0130860023281004E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.5007984423171797E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":-2.55774178759405E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.79833162863647E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.8325185591920401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":2.2531432192711199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.75423087193166405},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.88536114030873E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.8522824724979699E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-7.9640858511798696E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.22579674332342201},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.7445934276740299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":1.52922867019341E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-3.0563862560460901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.7447044288670799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A31,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":1.22070760883489E-4}]
["106SMPL.d",{"$f":"OFFSET(ConcTable!$B$1,MATCH($A32,ConcTable!$A:$A,0)-1,0)","v":"19286424-020"},{"$f":"OFFSET(SQ!$D$1,MATCH($A32,SQ!$C:$C,0)-1,0)","v":"2019-10-22T14:02:22"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.149},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.222},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-9.9373967863818996E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-95.651381558545594},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":6.1546208845900701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":12.710624130235299},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.63532184488276},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":-3.1994126943177E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":-2.0591685997233601E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":2.2888968381674198E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":4.8313883882211403E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.13807078350568699},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":1.01527136627371E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":5.6410906442159299E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-2.0154735320495099E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":48.596385576886199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-1.13863860899058E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-1.58316460801418E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":8.2124514457458705E-5},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-4.1551177621226199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-7.2315807314655298E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.0100557469619201E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":3.0692612879042501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-3.3917442650977501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.8019028296204101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":9.6685531943848497E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":8.0527461805448295E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.66507916236173903},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.27559958664691E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-1.6185239993008E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":-0.220116514417839},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.69628151305405E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.6812910340045401E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.8558947121834202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.0375370771726397E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.6612837342144799E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A32,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
[]
[]
[]
["086SMPL.d","Blank Control1",{"$f":"OFFSET(SQ!$D$1,MATCH($A36,SQ!$C:$C,0)-1,0)","v":"2019-10-22T13:32:00"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-4.0000000000000001E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-8.2000000000000003E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.19700000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":0.87619866229130405},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":1637.4178708966799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":8.2297101347621098},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":107.85756440185099},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":24.313789336799001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":9.4850906347703101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":6.4191173547548996E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":0.27780916360322999},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":0.139857891637583},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":1.5409540970059701},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":6.2746054985686897E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":6.9090364242712199E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":9.5804681126961597E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.27342486087787998},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":1.6661352361936999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":1.66515168885807E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":3.6934386431786499E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":4.3458313119352603E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":20.100763185252799},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":9.6713139170295101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":1.8852552854408601E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":2.19446218935923E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":7.1059286784700398E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":3.5116188924587399E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":3.0893933932825999E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":3.8385747039204401E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.9864251613338798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":1.0355091333684501},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":2.5104610780678199E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":8.8000029288708894E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":5.4439859377283697E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0535021038907101E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.31904884281034901},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":3.11958508161315E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":5.82797215220839E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":4.4040968648245997E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":5.4557597140578798E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":2.8360009986096901E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A36,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]
["110SMPL.d","Blank Control2",{"$f":"OFFSET(SQ!$D$1,MATCH($A37,SQ!$C:$C,0)-1,0)","v":"2019-10-22T14:08:29"},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(D$1,ConcTable!$1:$1,0)-1),0)","v":-0.02},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(E$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(F$1,ConcTable!$1:$1,0)-1),0)","v":-1.5980000000000001},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(G$1,ConcTable!$1:$1,0)-1),0)","v":0.27100000000000002},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(H$1,ConcTable!$1:$1,0)-1),0)","v":-0.71223130746462404},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(I$1,ConcTable!$1:$1,0)-1),0)","v":-41.8212412720275},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(J$1,ConcTable!$1:$1,0)-1),0)","v":-0.30922673109827098},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(K$1,ConcTable!$1:$1,0)-1),0)","v":19.5033760134787},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(L$1,ConcTable!$1:$1,0)-1),0)","v":-1.87529016895731},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(M$1,ConcTable!$1:$1,0)-1),0)","v":4.5441030106041103E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(N$1,ConcTable!$1:$1,0)-1),0)","v":4.8360186908304702E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(O$1,ConcTable!$1:$1,0)-1),0)","v":3.6616422601480798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(P$1,ConcTable!$1:$1,0)-1),0)","v":3.6810740201091902E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Q$1,ConcTable!$1:$1,0)-1),0)","v":0.584150931303558},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(R$1,ConcTable!$1:$1,0)-1),0)","v":-1.08212336288338E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(S$1,ConcTable!$1:$1,0)-1),0)","v":4.3983716691303798E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(T$1,ConcTable!$1:$1,0)-1),0)","v":-5.3781362800300898E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(U$1,ConcTable!$1:$1,0)-1),0)","v":0.445390029633631},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(V$1,ConcTable!$1:$1,0)-1),0)","v":-5.6520987092785503E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(W$1,ConcTable!$1:$1,0)-1),0)","v":-6.1804811515542204E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(X$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Y$1,ConcTable!$1:$1,0)-1),0)","v":2.5926220786887202E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(Z$1,ConcTable!$1:$1,0)-1),0)","v":-3.6418454877821702},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AA$1,ConcTable!$1:$1,0)-1),0)","v":-2.9282634140443499E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AB$1,ConcTable!$1:$1,0)-1),0)","v":-1.2812173995027E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AC$1,ConcTable!$1:$1,0)-1),0)","v":1.8831894543406E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AD$1,ConcTable!$1:$1,0)-1),0)","v":-4.0438626685527501E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AE$1,ConcTable!$1:$1,0)-1),0)","v":3.0677149098487101E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AF$1,ConcTable!$1:$1,0)-1),0)","v":-2.92067947888613E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AG$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AH$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AI$1,ConcTable!$1:$1,0)-1),0)","v":-2.6955795281847699E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AJ$1,ConcTable!$1:$1,0)-1),0)","v":1.1673246294258901E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AK$1,ConcTable!$1:$1,0)-1),0)","v":0},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AL$1,ConcTable!$1:$1,0)-1),0)","v":-0.75019098027478504},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AM$1,ConcTable!$1:$1,0)-1),0)","v":-1.45363761520707E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AN$1,ConcTable!$1:$1,0)-1),0)","v":-6.8242577112667497E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AO$1,ConcTable!$1:$1,0)-1),0)","v":-2.3010217901129701E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AP$1,ConcTable!$1:$1,0)-1),0)","v":1.0876006461429201E-6},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AQ$1,ConcTable!$1:$1,0)-1),0)","v":0.37826917102246199},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AR$1,ConcTable!$1:$1,0)-1),0)","v":-2.9565623544152E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AS$1,ConcTable!$1:$1,0)-1),0)","v":-4.06153294398568E-4},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AT$1,ConcTable!$1:$1,0)-1),0)","v":-2.9209124953830201E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AU$1,ConcTable!$1:$1,0)-1),0)","v":-5.3125898768027896E-3},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AV$1,ConcTable!$1:$1,0)-1),0)","v":-2.7376025538340801E-2},{"$f":"IFERROR(--OFFSET(ConcTable!$A$1,MATCH($A37,ConcTable!$A:$A,0)-1,MATCH(AW$1,ConcTable!$1:$1,0)-1),0)","v":0}]

{"sheet":"Summary3","index":10,"n_rows":50,"n_cols":6,"rows_kept":50,"cut":false,"hidden":false,"merged":[[0,0,0,5]]}
[{"$f":"INDIRECT(A50&\"!A5\")","v":"50 mL acid washed"}]
["Element","Result","Average (ppb)","Standard Deviation","Average + 2SD","Reporting Limit"]
["Lithium",{"$f":"OFFSET(INDIRECT($A$50&\"!$A$10\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":"Pass"},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$7\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":-1.8315789473684212E-2},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$8\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":2.6885070532618278E-3},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$9\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":-1.2938775367160557E-2},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$2\"),0,MATCH($A3,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":10}]
["Beryllium",{"$f":"OFFSET(INDIRECT($A$50&\"!$A$10\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":"Pass"},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$7\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":1.0526315789473685E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$8\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":3.1530176764230571E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$9\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":7.3586669317934829E-4},{"$f":"OFFSET(INDIRECT($A$50&\"!$A$2\"),0,MATCH($A4,INDIRECT($A$50&\"!$4:$4\"),0)-1)","v":0.25}]
["Boron","Pass",-1.0217368421052633,0.2125372757124063,-0.59666229068045062,25]
["Aluminium","Pass",0.26452631578947372,0.37506790028548348,1.0146621163604408,10]
["Magnesium","Pass",-0.23937399537704823,0.25217668809230359,0.26497938080755895,200]
["Silicon","Review",61.498573352765142,82.8549529112762,227.20847917531754,500]
["Phosphorus","Pass",-2.0984003652770014,2.818894508279385,3.5393886512817687,500]
["Potassium","Pass",13.855357016978015,6.3625910336330822,26.58053908424418,500]
["Calcium","Pass",2.4217043604347812,2.9388545389867593,8.2994134384082994,500]
["Titanium","Pass",3.0607059693709856E-2,3.1184410465421933E-2,9.2975880624553725E-2,5]
["Vanadium","Pass",1.0135291784705174E-2,7.4396480185479107E-3,2.5014587821800993E-2,0.5]
["Chromium","Pass",1.8631254442402264E-2,1.3654643294718412E-2,4.5940541031839088E-2,0.5]
["Manganese","Pass",3.3166009174656064E-2,1.2454164706893353E-2,5.807433858844277E-2,0.5]
["Iron","Pass",6.8803584603456258E-2,0.16584020769731189,0.40048399999808004,20]
["Cobalt","Pass",-7.7363073864977164E-4,9.756822268868631E-4,1.1777337151239547E-3,0.25]
["Nickel","Pass",5.2396048035437869E-2,7.2296409120795144E-2,0.19698886627702816,1]
["Copper","Pass",3.0134953890754702E-2,9.2270765809901789E-2,0.21467648551055829,5]
["Zinc","Review",55.298010406535731,5.3040588351120919,65.906128076759913,50]
["Gallium","Pass",-6.9671118726126425E-3,4.9195854821148461E-3,2.8720590916170496E-3,0.5]
["Germanium","Pass",-2.9923741184568055E-3,5.2860191868483435E-3,7.5796642552398812E-3,0.5]
["Arsenic","Pass",4.993252632980148E-4,2.2714159949779253E-3,5.0421572532538652E-3,2.5]
["Selenium","Pass",6.6493017544048015E-3,2.4296997213841598E-2,5.5243296182087995E-2,10]
["Bromine","Pass",-2.577818671796392,2.3847520688702621,2.1916854659441323,250]
["Rubidium","Pass",-3.6801715364541202E-2,2.9298135604016284E-2,2.1794555843491367E-2,25]
["Strontium","Pass",-6.2263461856045593E-3,1.8688844618340767E-2,3.1151343051076974E-2,5]
["Zirconium","Pass",4.7525736752840922E-4,1.5679287237252203E-3,3.6111148149788499E-3,2]
["Molybdenum","Pass",-4.1385303518445644E-3,1.2937880965186632E-3,-1.550954158807238E-3,0.5]
["Palladium","Pass",-1.8984765528776725E-3,1.2858415080153898E-3,6.7320646315310699E-4,0.5]
["Silver","Pass",-2.055794694006529E-2,6.0927383537682546E-3,-8.3724702325287809E-3,1]
["Cadmium","Pass",3.8174908764113347E-4,3.4947622272601313E-4,1.0807015330931598E-3,0.25]
["Indium","Pass",0,0,0,0.25]
["Tin","Pass",-2.711920979409739E-2,7.486699341197548E-3,-1.2145811111702294E-2,1]
["Antimony","Pass",1.2287891592981403E-3,1.005469241236458E-3,3.2397276417710564E-3,0.5]
["Tellurium","Pass",0,0,0,0.5]
["Iodine","Pass",-0.67247404019236179,0.12737118997157426,-0.41773166024921327,10]
["Caesium","Pass",-1.570220179316621E-3,4.8583245863993615E-4,-5.9855526203674871E-4,0.5]
["Barium","Pass",-4.2092183067321412E-3,3.0826211601820324E-3,1.9560240136319235E-3,2.5]
["Tungsten","Pass",-1.6458432890333836E-4,2.0667553112884397E-4,2.4876673335434956E-4,0.5]
["Platinum","Pass",5.1936664085896786E-5,1.102180531137353E-4,2.723727703133674E-4,0.5]
["Gold","Pass",-0.17626135906752857,4.1740136742664705E-2,-9.2781085582199163E-2,5]
["Mercury","Pass",-2.2118587363171544E-2,4.6693506190975502E-3,-1.2779886124976443E-2,0.5]
["Thallium","Pass",-3.3094019716774876E-4,2.117066303786603E-4,9.2473063589571848E-5,0.25]
["Lead","Pass",-4.2623593794651886E-4,3.9538714425378452E-3,7.4815069471291717E-3,2.5]
["Bismuth","Pass",-4.5381258635102716E-3,6.2724315127319686E-4,-3.2836395609638779E-3,0.25]
["Thorium","Pass",-2.4455585830378008E-2,3.7632563243504844E-3,-1.6929073181677037E-2,0.25]
["Uranium","Pass",1.9680714594182863E-5,3.4623157953879255E-5,8.892703050194137E-5,0.05]
[]
["Container3"]